## Périmètre connecté
Quant | DESIGNATION | MARQUE | TYPE MODELE | CONNECTIVITE CAPSULE | Waveform Capable | Waveform Supported | DRIVER | COMMENTAIRES/PRE-REQUIS
3 | GENER. DIALYSE | GAMBRO | ARTIS 230V | Oui | Non | Non | BaxterC 5.3.0
2 | MONITEUR SURVEILL. | DELTEX MEDICAL | CARDIO Q ODM | Oui | Non | Non | DeltexA 5.0.1
20 | RESPIRATEUR REANIMATION | DRAGER | EVITA INFINITY V500 | Oui | Oui | Oui | DragerMedibus 5.3.21
6 | RESPIRATEUR REANIMATION | DRAGER | EVITA V800 | Oui | Oui | Oui | DragerMedibus 5.3.21
20 | STATION ARRIMAGE | FRESENIUS VIAL | EXELIA COMBOX | Non | Non | Non | FreseniusB 5.3.0 | Connectivité via la Gateway VSS de Fresenius. Driver en cours de développement, attendu pour fin Q1 2022
200 | POUSSE-SERINGUE | FRESENIUS VIAL | EXELIA SP | Non | Non | Non | FreseniusB 5.3.0 | Connexion possible seulement via la Gateway VSS (Vigilant Software Suite). Le Driver pour ce dispositif est en cours de développement et attendu pour end of Q1 2021
40 | POMPE A PERFUSION | FRESENIUS VIAL | EXELIA VP | Non | Non | Non | FreseniusB 5.3.0 | Connexion possible seulement via la Gateway VSS (Vigilant Software Suite). Le Driver pour ce dispositif est en cours de développement et attendu pour end of Q1 2021
15 | POMPE DE NUTRITION | FRESENIUS VIAL | APPLIX SMART | Oui | Non | Non | FreseniusD 5.0.2
2 | MONITEUR HEMOFILTR. | FRESENIUS MEDICAL CARE | MULTIFILTRATE PRO | Oui | Non | Non | FreseniusG 5.2.4 | Nport 5210 (MOXA) à acquérir par l'hôpital pour pouvoir connecter ce dispositif par câble série
1 | GENER. DIALYSE | FRESENIUS MEDICAL CARE | 5008 | Oui | Non | Non | FreseniusI 5.2.0 | L'accès à ce driver est soumis à l'autorisation de Fresenius pour le site
8 | LIT SPECIALISE PESEE | HILL ROM | TOTALCARE SCALE | Oui | Non | Non | HillRomA 5.0.0 | Bed Manager Gateway est nécessaire pour pouvoir connecter ces lits
1 | MONITEUR SURVEILL. IRM | PHILIPS | Expression IP5 | Oui | Non | Non | InvivoB 5.1.6 | Si le dispositif possède un port USB, alors il est nécessaire d'acquérir le câble USB-Serial Adapter for IP5 CP-US-03 auprès de Philips, s'il est pas déjà inclus avec l'appareil
1 | MONITEUR SURVEILL. IRM | PHILIPS | Expression MR400 | Oui | Non | Non | InvivoB 5.1.6 | Les données de cet appareil peuvent être collectées via le moniteur IP5
3 | GENERATEUR HYPOTHERMIE | MTRE | CRITICOOL | Oui | Non | Non | MTREA 5.1.0
21 | MODULE MULTIPARAMETR | PHILIPS | INTELLIVUE X3 | Oui | Oui | Oui | PhilipsDataExport 5.3.18
4 | MONITEUR SURVEILL. | PHILIPS | MP20-NELLCOR SPO2 | Oui | Oui | Oui | PhilipsDataExport 5.3.18 | Carte MIB/RS232 Interface / (Ref Philips : Option J13) est nécessaire pour pouvoir connecter ce dispositif par cable série
3 | MONITEUR SURVEILL. | PHILIPS | MP70 - NELLCOR SPO2 | Oui | Oui | Oui | PhilipsDataExport 5.3.18 | Carte MIB/RS232 Interface / (Ref Philips : Option J13) est nécessaire pour pouvoir connecter ce dispositif par cable série
21 | MONITEUR SURVEILL. | PHILIPS | MX800 | Oui | Oui | Oui | PhilipsDataExport 5.3.18 | Carte MIB/RS232 Interface / (Ref Philips : Option J13) est nécessaire pour pouvoir connecter ce dispositif par cable série
5 | UNITE CENTRALE | HP | PIIC IX HARDWARE | Oui | Non | Non | PhilipsIIC 5.3.14
1 | MONITEUR SURVEILL. | PULSION | PICCO2 | Oui | Oui | Oui | PulsionA 5.0.0
7 | RESPIRATEUR VNI | RESPIRONICS | V60 | Oui | Oui | Non | RespironicsEsprit 5.1.6 | Ce dispositif est capable d'envoyer des courbes mais le driver actuel ne les prend pas en charge. Un update du driver peut être fait si participation financière au développement
5 | RESPIRATEUR | AIR LIQUIDE | MONNAL T60 | Oui | Oui | Non | TaemaA 5.3.3 | Nécessite acquisition du câble YR135600 USB - RS232 auprès d'Air Liquide Medical Systems
3 | ILA ACTIVE SENSORBOX | NOVALUNG | Xenios Console | Oui | Non | Non | XeniosA 5.3.0

## Artis
ID | Label | Unit | Description | Type | Group | Device | Source
1 | Heart Rate | 1/min | The MeasurementTime of BP variables are only updated when a new measurement is performed. In between measurements, the last valid values are sent repeatedly. | Long | VitalSigns | Common
This variable is associated with SourceID identifier.
2 | NIBP Systolic | mmHg | The MeasurementTime of BP variables are only updated when a new measurement is performed. In between measurements, the last valid values are sent repeatedly. | Long | VitalSigns | Common
3 | NIBP Diastolic | mmHg | The MeasurementTime of BP variables are only updated when a new measurement is performed. In between measurements, the last valid values are sent repeatedly. | Long | VitalSigns | Common
456 | Dialysate Inlet Pressure | mmHg |  | Long | VitalSigns | Common
457 | Dialysate Outlet Pressure | mmHg |  | Long | VitalSigns | Common
650 | Stroke Volume | mL |  | Long | VitalSigns | Common
757 | Patient Weight | kg |  | Double | PatientInfo | Common
1011 | Therapy Running Status | unitless | Possible values: | Long | DeviceInfo | Common
0: No treatment running
32: Treatment running
3003 | Blood Flow Rate Setting | mL/min |  | Long | Settings | Common
3006 | Fluid Removal Volume Setting | mL |  | Long | Settings | Common
3029 | Transmembrane Pressure | mmHg |  | Long | VitalSigns | Common
3168 | Fluid Removal Actual Rate | mL/h |  | Long | VitalSigns | Common
3236 | Relative Blood Volume Change | % |  | Double | VitalSigns | Common
3248 | Generic Venous Line Pressure | mmHg |  | Long | VitalSigns | Common
3479 | Equivalent Dialysate Sodium (Na) | mmol/L |  | Long | VitalSigns | Common
3493 | Desired Equivalent Dialysate Sodium (Na) | mmol/L |  | Long | Settings | Common
3627 | Venous Pump Flow | mL/min |  | Long | VitalSigns | Common
3704 | Arterial/Main Pump Flow | mL/min |  | Long | VitalSigns | Common
4226 | Replacement Cumulated Volume | L |  | Double | VitalSigns | Common
4229 | Fluid Removal Cumulated Volume | mL |  | Long | VitalSigns | Common
4236 | Transmembrane Pressure Setting | mmHg |  | Long | Settings | Common
4237 | Desired Blood Volume Change Percentage | % |  | Double | Settings | Common
4240 | Plasma Sodium (Na) | mmol/L |  | Long | VitalSigns | Common
4241 | Dilution Ratio (Qf/Qb) | unitless |  | Double | VitalSigns | Common
4347 | Therapy Remaining Time | s |  | Long | OtherMeasuredData | Common
5850 | Anticoagulant Cumulated Volume | mL |  | Double | VitalSigns | Common
5853 | Delivered Treatment Effectiveness (Kt/V) | unitless |  | Double | VitalSigns | Common
5855 | Blood Flow Rate | mL/min |  | Long | VitalSigns | Common
5856 | Dialysate Flow Rate | mL/min |  | Long | VitalSigns | Common
5861 | Effective Blood Volume Processed (Kt) | L |  | Double | VitalSigns | Common
5863 | Dialysate Temperature | °C |  | Double | VitalSigns | Common
5884 | Replacement Cumulated Bolus Volume | mL |  | Long | VitalSigns | Common
5888 | Dialysate pH | pH |  | Double | VitalSigns | Common
5889 | Ionic Clearance | mL/min | Ionic Dialysance | Long | VitalSigns | Common
5890 | Ionic Mass Balance | mmol/L |  | Long | VitalSigns | Common
5892 | Bicarbonate Conductivity (8.4%) | mS/cm |  | Double | VitalSigns | Common
5899 | Actual Anticoagulant Rate | mL/h |  | Double | VitalSigns | Common
6643 | Substitution Fluid And Blood Ratio (Qi/Qb) | unitless |  | Long | Measurements | Common
7438 | Generic Arterial Line Pressure | mmHg |  | Long | VitalSigns | Common
7979 | Dialysate Conductivity | mS/cm |  | Double | VitalSigns | Common
7988 | Plasma Conductivity | mS/cm |  | Double | VitalSigns | Common
8811 | Therapy Elapsed Time | min |  | Long | OtherMeasuredData | Common
8813 | Processed Cumulated Blood Volume | L |  | Double | VitalSigns | Common
8856 | Therapy Operation Status | unitless | Consult detailed help for more information on this variable | Long | DeviceInfo | Common
9642 | Replacement Flow Rate | mL/h |  | Long | VitalSigns | Common

## CardioQ
ID | Label | Unit | Description | Type | Group | Device | Source
1 | Heart Rate | 1/min |  | Double | VitalSigns | Common
414 | Cardiac Output | L/min | This variable is available only if patient category 'adult' or 'paediatric' is selected. | Double | VitalSigns | Common
585 | Patient Category | unitless | A: adult C: paediatric L: linear | String | PatientInfo | Common
649 | Cardiac Index | L/min/m2 | This variable is available only if patient category 'adult' or 'paediatric' is selected. | Double | VitalSigns | Common
650 | Stroke Volume | mL | This variable is available only if patient category 'adult' or 'paediatric' is selected. | Double | VitalSigns | Common
651 | Stroke Volume Index | mL/m2 | This variable is available only if patient category 'adult' or 'paediatric' is selected. | Double | VitalSigns | Common
2255 | Device Date/Time | unitless |  | String | DeviceInfo | Common
6412 | Blood Minute Distance | cm |  | Double | VitalSigns | Common
6413 | Blood Stroke Distance | cm |  | Double | VitalSigns | Common
6414 | Flow Time To Peak | ms |  | Double | VitalSigns | Common
6415 | Blood Mean Acceleration | m/s2 |  | Double | VitalSigns | Common
8117 | Corrected Flow Time | ms |  | Double | VitalSigns | Common
8118 | Peak Velocity | cm/s |  | Double | VitalSigns | Common

## Evita V500
ID | Label | Unit | Description | Type | Group | Device | Source
20 | End Tidal CO2 (EtCO2) | - | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Evita V300, Fabius Family. This variable is associated with UpperAlarmLimit attribute, which is available for Perseus A500, Atlan, Evita Infinity V500, Babylog VN500, Savina 300, Primus, Primus IE, Apollo, Zeus IE, Oxylog 3000 plus, Evita V300. This variable is associated with LowerAlarmLimit attribute, which is available for Perseus A500, Atlan, Evita Infinity V500, Babylog VN500, Savina 300, Primus, Primus IE, Apollo, Zeus IE, Oxylog 3000 plus, Evita V300. Possible units: mmHg, %, kPa. This variable is in Multiset. See variables 2355 and 6551 for concentration and partial pressure respectively. For Evita Infinity V500, Babylog VN500: This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. end tidal CO2 etCO2 Note: etCO2 values and alarm limits are transmitted according to the configuration on the device screen (mmHg or Vol%). If kPa is configured be aware that the Medibus alarm limits will be transmitted with mmHg. | Double (Multiset) | VitalSigns | Evita Infinity V500 [MEDIBUS.X]
60 | Tidal Volume | mL | For Evita 4, Evita 2 dura, Evita XL: Expiratory tidal volume. This variable is associated with UpperAlarmLimit attribute, which is also sent in separate variable 2763. Values below 1 L will be transmitted in units of mL, otherwise in units of L. Device firmware version 04.00 and higher: Limit is not transmitted, if alarm is switched off. For Savina: This variable is associated with UpperAlarmLimit attribute.Device SW version 02.00 and higher: Limit is not transmitted, if alarm is switched off. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300, Fabius Family. This variable is associated with UpperAlarmLimit attribute, which is available for Perseus A500, Evita Infinity V500, Savina 300, Evita V300. This variable is associated with LowerAlarmLimit attribute, which is available for Evita Infinity V500, Evita V300. For Evita Infinity V500, Babylog VN500: This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. For Zeus, Zeus IE, Primus, Primus IE, Apollo, Pallas: Exp. Tidal Volume. | Double | VitalSigns | Evita Infinity V500 [MEDIBUS.X]
61 | Dynamic Compliance | L/bar | For Evita 4, Evita 2 dura, Evita XL: Values below 100 will be transmitted with one decimal place, otherwise without a decimal place. For Savina: Available since device SW version 02.00. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Evita V300. | Double | VitalSigns | Evita Infinity V500 [MEDIBUS.X]
114 | CO2 Concentration Waveform | - | For Evita 4, Evita 2 dura, Evita XL: Possible units: mmHg, kPa, and % For Evita 2: Possible units: mmHg, kPa, and % Firmware version 2.00 or higher. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Evita V300. Possible units: mmHg, %, kPa. Unit "mmHg" is available for Perseus A500, Atlan, Evita Infinity V500, Babylog VN500, Savina 300, Primus, Primus IE, Apollo, Zeus IE, Oxylog 3000 plus, Evita V300. Unit "kPa" and "%" are available for Perseus A500, Atlan, Savina 300, Primus, Primus IE, Apollo, Zeus IE, Oxylog 3000 plus. See variables 8866 and 7872 for concentration and partial pressure respectively. For Evita Infinity V500, Babylog VN500: Possible units: mmHg, %. Expiratory CO2 concentration. For Zeus, Zeus IE, Primus, Primus IE, Apollo, Pallas, Vamos, Vamos plus: Possible units: mmHg, %, kPa. | Double | Waveforms | Evita Infinity V500 [MEDIBUS.X]
118 | Airway Flow Wave | L/min | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300, Fabius Family. For Evita Infinity V500, Babylog VN500, Carina, Zeus, Zeus IE, Primus, Primus IE, Apollo, Pallas: Flow (insp./exp.) | Double | Waveforms | Evita Infinity V500 [MEDIBUS.X]
177 | Alarm Data | unitless | This variable is same as 3878 but sent as a string, which is an array containing all alarm messages and separated by delimiter "¦¦". This variable, along with variables 588 and 4131, are marked as obsolete because they use either"¦¦" or "," in the array string as delimiter, which are often not compatible with customers' applications. Variables 6164, 3878, and 2596 should be used instead. The HL7 output can contain an '&' character to indicate a LOW alarm. In order to avoid issues, the '&' character will be replaced by a '?' character in the DDI output. | String | Alarms | Evita Infinity V500 [MEDIBUS.X] | Alarm And Text Message
582 | Resistance | mbar/L/s | For Evita 4, Evita 2 dura, Evita XL: Values below 100 will be transmitted with one decimal place, otherwise without a decimal place. For Savina: Available since device SW version 02.00. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Zeus IE, Evita V300. | Double | VitalSigns | Evita Infinity V500 [MEDIBUS.X]
584 | Ventilation Mode | unitless | The ventilation mode is reported in variable 584 and 2756. Variable 2756 contains mode ID numeric value, while variable 584 contains string representation which is the complete ventilation mode information concatenated from basic ventilation mode together with available options such as "Mode VC-CMV /AF". See variable 2756 for full list of supported ventilation mode. | String | DeviceInfo | Evita Infinity V500 [MEDIBUS.X]
588 | Alarm Message | unitless | This variable is same as 6164 but sent as a string, which is an array containing all alarm codes and separated by delimiter ",". This variable, along with variables 177 and 4131, are marked as obsolete because they use either"¦¦" or "," in the array string as delimiter, which are often not compatible with customers' applications. Variables 6164, 3878, and 2596 should be used instead. | String | Alarms | Evita Infinity V500 [MEDIBUS.X] | Alarm And Text Message
635 | Inspired O2 (FiO2) | % | For Savina: This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. Available since device SW version 03.00. Limits are transmitted in LPO-Mode (Option) only. For Fabius GS, Fabius GS premium, Fabius Tiro, Fabius plus, Fabius plus XL, Fabius OS, Fabius MRI: This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Evita V300, Fabius Family. This variable is associated with UpperAlarmLimit attribute, which is available for Perseus A500, Atlan, Primus, Primus IE, Apollo, Zeus IE, Fabius Family. This variable is associated with LowerAlarmLimit attribute, which is available for Perseus A500, Atlan, Primus, Primus IE, Apollo, Zeus IE, Fabius Family. Inspiratory oxygen fraction, FiO2. For Zeus, Zeus IE: This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. For Primus, Primus IE, Apollo, Pallas: This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. | Double | VitalSigns | Evita Infinity V500 [MEDIBUS.X]
754 | Device Name | unitless | This variable contains the device name defined by the manufacture. See variable 3591 for device model identification number. | String | DeviceInfo | Common | Device Information
775 | Inspiratory Component of I:E Ratio | unitless | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: Obsolete variable that may be removed in future release. Partial information of I:E ratio. For Evita Infinity V500, Savina: Partial information. Please use variable 1190 instead. Obsolete variable that may be removed in future release. | Double | VitalSigns | Evita Infinity V500 [MEDIBUS.X]
776 | Expiratory Component of I:E Ratio | unitless | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: Obsolete variable that may be removed in future release. Partial information of I:E ratio. For Evita Infinity V500, Savina: Partial information. Please use variable 1190 instead. Obsolete variable that may be removed in future release. | Double | VitalSigns | Evita Infinity V500 [MEDIBUS.X]
1189 | Positive End Expiratory Pressure (PEEP) | mbar | For Fabius GS, Fabius GS premium, Fabius Tiro, Fabius plus, Fabius plus XL, Fabius OS, Fabius MRI: 1mbar = 1cmH2O. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300, Fabius Family. | Double | VitalSigns | Evita Infinity V500 [MEDIBUS.X]
1190 | I:E Ratio | unitless | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Evita V300. Format: 1:x if E > I, x:1 if I > E, or 1:1 if I = E. The useless leading or ending “0” or “.” is not transmitted, and the maximum digits after the floating point is 2. For Evita Infinity V500, Babylog VN500: Format: 1:x if E > I, x:1 if I > E, or 1:1 if I = E. The useless leading or ending “0” or “.” is not transmitted, and the maximum digits after the floating point is 2. | String | VitalSigns | Evita Infinity V500 [MEDIBUS.X]
1197 | Message Text | unitless | This variable, along with variable 3909 are marked as obsolete because they use either"¦¦" or "," in the array string as delimiter, which are often not compatible with customers' applications. Variables 8366 and 8696 should be used instead. | String | Messages | Evita Infinity V500 [MEDIBUS.X] | Alarm And Text Message
1306 | Tidal Volume | mL | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: Obsolete variable that may be removed in future release. This variable is marked as deprecated in order to synchronize with other models. Please use variable 60 instead. For Evita Infinity V500, Babylog VN500: Deprecated in order to synchronize with other models. Please use variable 60 instead. | Double | VitalSigns | Evita Infinity V500 [MEDIBUS.X]
1307 | Minute Volume | L/min | For Evita 4, Evita 2 dura, Evita XL: This variable is associated with UpperAlarmLimit attribute, which is also sent in separate variable 1400. This variable is associated with LowerAlarmLimit attribute, which is also sent in separate variable 1399. For device firmware version 04.00 and higher: Lower alarm limit is not transmitted, if alarm is switched off. Values may be transmitted with one or two decimal place, depended on firmware version. Note that the value sent by the device is less precise than what is displayed on device screen. For example, a displayed value of 5.12 is sent by the device and thus by DDI as 5.1. For Evita 2, Evita, Fabius GS, Fabius GS premium, Fabius Tiro, Fabius plus, Fabius plus XL, Fabius OS, Fabius MRI, Evita Infinity V500, Babylog VN500, Carina: This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. This variable is available only with a precision to the tenth. For example: for a minute volume value of 7.43, the device displays 7.43 L/min but the DDI will output it as 7.4 L/min. For Savina: This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300, Fabius Family. This variable is associated with UpperAlarmLimit attribute, which is available for Perseus A500, Atlan, Savina 300, Primus, Primus IE, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Fabius Family. This variable is associated with LowerAlarmLimit attribute, which is available for Perseus A500, Atlan, Savina 300, Primus, Primus IE, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Fabius Family. This variable is available only with a precision to the tenth. For example: for a minute volume value of 7.43, the device displays 7.43 L/min but the DDI will output it as 7.4 L/min. For Zeus, Zeus IE, Primus, Primus IE, Apollo, Pallas: This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. Respiratory Minute Volume This variable is available only with a precision to the tenth. For example: for a minute volume value of 7.43, the device displays 7.43 L/min but the DDI will output it as 7.4 L/min. For Evita Infinity V500 and Babylog VN500: This variable is only output if the device version is less than 2.30, otherwise variable 2326 is exported. | Double | VitalSigns | Evita Infinity V500 [MEDIBUS.X]
1314 | Spontaneous Respiration Rate | 1/min | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Zeus IE, Oxylog 3000 plus, Evita V300. This variable is associated with UpperAlarmLimit attribute, which is available for Oxylog 3000 plus. | Double | VitalSigns | Evita Infinity V500 [MEDIBUS.X]
1318 | Static Compliance | mL/mbar | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500 Evita Infinity V600, Evita Infinity V800. For Evita Infinity V500: Cstat (Low Flow) Cstat. | Double | VitalSigns | Evita Infinity V500 [MEDIBUS.X]
1320 | Mandatory Respiration Rate Setting | 1/min | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300, Fabius Family. Respiratory rate, RR. For Evita Infinity V500, Babylog VN500: Respiratory Rate RR. | Double | Settings | Evita Infinity V500 [MEDIBUS.X]
1321 | Tidal Volume Setting | - | For Evita 4, Evita 2 dura, Evita XL, Evita 2, Evita, Savina, Fabius GS, Fabius GS premium, Fabius Tiro, Fabius plus, Fabius plus XL, Fabius OS, Fabius MRI, Carina, Zeus, Zeus IE, Primus, Primus IE, Apollo, Pallas: Inspired tidal volume setting. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: Obsolete variable that may be removed in future release. Variable is in Multiset format. Please use separate variable 2771 and 5044 instead. For Evita Infinity V500, Babylog VN500: Variable 2771 and 5044 contains same information but in different units. This variable combines these two variables and sends whichever is significant. For Zeus, Zeus IE: Obsolete variable that may be removed in future release. Multiset variable. Please use variables 2771 and 5044 instead. | Double (Multiset) | Settings | Evita Infinity V500 [MEDIBUS.X]
1324 | Sigh Interval Setting | s | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. Time interval between sighs, Interval sigh. | Double | Settings | Evita Infinity V500 [MEDIBUS.X]
1326 | Sigh Number Setting | unitless | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. Number of sigh cycles, Cycles sigh. | Double | Settings | Evita Infinity V500 [MEDIBUS.X]
1328 | Apnea Tidal Volume Setting | - | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Oxylog 3000 plus, Carina, Evita V300. Tidal volume for apnea ventilation, Vtapn. Possible units: L, mL. Unit "L" is available for Evita Infinity V500, Babylog VN500, Savina 300, Oxylog 3000 plus, Carina, Evita V300. Unit "mL" is available for Evita Infinity V500, Babylog VN500, Evita V300. For Evita Infinity V500, Babylog VN500: Possible units: L, mL. VT Apnea VTapn. | Double (Multiset) | Settings | Evita Infinity V500 [MEDIBUS.X]
1329 | Apnea Respiration Rate Setting | 1/min | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Oxylog 3000 plus, Carina, Evita V300. Respiratory rate for apnea ventilation, Rrapn. | Double | Settings | Evita Infinity V500 [MEDIBUS.X]
1332 | Pressure Support Level Above PEEP Setting | mbar | For Evita 4, Evita 2 dura, Evita XL: This variable is calculated from variables 1392 and 2104. Available only with Evita XL. For Medibus.X compatible device, Perseus A500, Atlan, Evita V300, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300, Fabius Family. Pressure amplitude above PEEP in Pressure Support, DPsupp. For Evita Infinity V500, Babylog VN500: ΔPsupp. For Carina: Assisted spontaneous breathing. | Double | Settings | Evita Infinity V500 [MEDIBUS.X]
1381 | Inspiratory Flow Setting | L/min | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300, Fabius Family. | Double | Settings | Evita Infinity V500 [MEDIBUS.X]
1383 | Inspiratory Flow Trigger Setting | L/min | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Evita V300, Fabius Family. Flow trigger threshold, Flow trigger. For Primus, Primus IE, Apollo, Pallas: Flow trigger. | Double | Settings | Evita Infinity V500 [MEDIBUS.X]
1386 | Intermittent PEEP Setting | mbar | For Fabius GS, Fabius GS premium, Fabius Tiro, Fabius plus, Fabius plus XL, Fabius OS, Fabius MRI: 1mbar = 1cmH2O. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Evita V300. Additional intermittent PEEP for sighs, PEEPsigh. For Evita Infinity V500, Babylog VN500: Interm. PEEP ΔintPEEP. For Zeus, Zeus IE: PEEP for all vent mode different from Man/Spont. | Double | Settings | Evita Infinity V500 [MEDIBUS.X]
1387 | APRV Low Pressure Setting | mbar | For Evita: BIPAP low pressure For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Evita V300. Lower pressure level during APRV, Plow. For Evita Infinity V500, Babylog VN500: APRV Low Pressure Plow. | Double | Settings | Evita Infinity V500 [MEDIBUS.X]
1388 | APRV Low Time Setting | s | For Evita: BIPAP low time For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Evita V300. Time of lower pressure level during APRV, Tlow. For Evita Infinity V500, Babylog VN500: APRV Low Time Tlow. | Double | Settings | Evita Infinity V500 [MEDIBUS.X]
1389 | APRV High Pressure Setting | mbar | For Evita: BIPAP high pressure For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Evita V300. Upper pressure level during APRV, Phigh. For Evita Infinity V500, Babylog VN500: APRV High Pressure Phigh. | Double | Settings | Evita Infinity V500 [MEDIBUS.X]
1390 | APRV High Time Setting | s | For Evita: BIPAP high time For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Evita V300. Time of upper pressure level during APRV, Thigh. For Evita Infinity V500, Babylog VN500: APRV High Time Thigh. | Double | Settings | Evita Infinity V500 [MEDIBUS.X]
1391 | Apnea Time Setting | s | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Oxylog 3000 plus, Carina, Evita V300. For Evita Infinity V500, Babylog VN500: Apnea Time Tapn. For Carina: Apnea alarm time. | Double | Settings | Evita Infinity V500 [MEDIBUS.X]
1393 | Maximum Inspired Pressure Setting | mbar | For Fabius GS, Fabius GS premium, Fabius Tiro, Fabius plus, Fabius plus XL, Fabius OS, Fabius MRI: 1mbar = 1cmH2O. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300, Fabius Family. Pressure limitation, Pmax. For Evita Infinity V500, Babylog VN500: Max Pressure Pmax. For Carina: Pressure limitation. For Primus, Primus IE, Apollo, Pallas: Max. insp. Airway pressure. | Double | Settings | Evita Infinity V500 [MEDIBUS.X]
1415 | Mean Airway Pressure | mbar | For Fabius GS, Fabius GS premium, Fabius Tiro, Fabius plus, Fabius plus XL, Fabius OS, Fabius MRI: This variable is associated with LowerAlarmLimit attribute. 1mbar = 1cmH2O. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300, Fabius Family. For Zeus, Zeus IE, Primus, Primus IE, Apollo, Pallas: Mean Breathing Pressure. | Double | VitalSigns | Evita Infinity V500 [MEDIBUS.X]
1422 | Rapid Shallow Breathing Index | - | For Evita 4, Evita 2 dura, Evita XL: Firmware version 04.00 or higher For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Atlan, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Evita V300. Possible units: 1/L * min, 1/mL*min. Unit "1/L * min" is available for Evita Infinity V500, Babylog VN500, Savina 300, Evita V300. Unit "1/mL*min" is available for Evita Infinity V500, Babylog VN500, Evita V300 (added with SW 2.30 for Babylog VN500, added with SW 2.30 for Evita Infinity V500). | Double (Multiset) | VitalSigns | Evita Infinity V500 [MEDIBUS.X]
1426 | Intrinsic PEEP | mbar | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. | Double | VitalSigns | Evita Infinity V500 [MEDIBUS.X]
1569 | Minimum Airway Pressure | mbar | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. | Double | VitalSigns | Evita Infinity V500 [MEDIBUS.X]
1570 | P0.1 Occlusion Pressure | mbar | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. Occlusion pressure in P0.1 maneuver. For Evita Infinity V500, Babylog VN500: Only available for Evita Infinity V500 | Double | VitalSigns | Evita Infinity V500 [MEDIBUS.X]
1571 | Lung Volume Trapped (TrpVol) | mL | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. Volume trapped in the lung by intrinsic PEEP, Vtrap. For Evita Infinity V500: Trapped Volume Vtrap. | Double | VitalSigns | Evita Infinity V500 [MEDIBUS.X]
1575 | Respired Volume Wave | mL | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Evita V300. Respiratory volume since start of inspiration, Volume. For Evita Infinity V500, Babylog VN500: Volume since inspiration began. | Double | Waveforms | Evita Infinity V500 [MEDIBUS.X]
1637 | Ventilation Pressure Variability Setting | % | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800. For Evita Infinity V500, Babylog VN500: Supported through Medibus.X protocol. | Double | Settings | Evita Infinity V500 [MEDIBUS.X]
1638 | APRV Low Time Limitation | s | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. Maximum time of low pressure level during APRV/AutoRelease, Tlow max. For Evita Infinity V500, Babylog VN500: Tlow max. | Double | Settings | Evita Infinity V500 [MEDIBUS.X]
1640 | Spontaneous Mean Tidal Volume | mL | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500 Evita Infinity V600, Evita Infinity V800. Spontaneous mean tidal volume, VTspon mean. For Evita Infinity V500, Babylog VN500: VTspon mean (Not available on Babylog VN500). Only available for Evita Infinity V500 | Double | VitalSigns | Evita Infinity V500 [MEDIBUS.X]
1641 | Spontaneous Inspired Mean Tidal Volume | mL | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500 Evita Infinity V600, Evita Infinity V800. For Evita Infinity V500, Babylog VN500: VTispon mean (Not available on Babylog VN500). Only available for Evita Infinity V500 | Double | VitalSigns | Evita Infinity V500 [MEDIBUS.X]
1642 | Spontaneous Expired Mean Tidal Volume | mL | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800. For Evita Infinity V500, Babylog VN500: VTespon mean (Not available on Babylog VN500). Only available for Evita Infinity V500 | Double | VitalSigns | Evita Infinity V500 [MEDIBUS.X]
1650 | SmartCare FiO2 Limitation | % | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. FiO2 in SmartCare - upper limit, SC-FiO2high For Evita Infinity V500, Babylog VN500: SC-Upper limit inspiratory O2 FiO2 max. | Double | Settings | Evita Infinity V500 [MEDIBUS.X]
1651 | SmartCare PEEP Limitation | mbar | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. PEEP in SmartCare - upper limit, SC-PEEPmax. For Evita Infinity V500, Babylog VN500: SC-Upper limit PEEP, PEEP max. | Double | Settings | Evita Infinity V500 [MEDIBUS.X]
1652 | SmartCare Spontaneous Respiration Rate Lower Limitation | 1/min | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. Spontaneous respiratory rate in SmartCare - lower limit, SC-RRspon low. For Evita Infinity V500, Babylog VN500: SC-Lower limit respiratory rate in the guideline, RRspon low. | Double | Settings | Evita Infinity V500 [MEDIBUS.X]
1653 | SmartCare Spontaneous Respiration Rate Upper Limitation | 1/min | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. Spontaneous respiratory rate in SmartCare - upper limit, SC-RRspon high. For Evita Infinity V500, Babylog VN500: SC-Upper limit respiratory rate in the guideline, RRspon high. | Double | Settings | Evita Infinity V500 [MEDIBUS.X]
1654 | SmartCare Tidal Volume Limitation | mL/kg | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. Tidal volume in SmartCare - lower limit, SC-VTlow. For Evita Infinity V500, Babylog VN500: SC-Lower limit tidal volume, VT low. | Double | Settings | Evita Infinity V500 [MEDIBUS.X]
1655 | SmartCare EtCO2 Limitation | mmHg | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. End-tidal CO2 in SmartCare - upper limit, SC-etCO2high. For Evita Infinity V500, Babylog VN500: SC-Upper limit endtidal CO2 in the guideline etCO2 high. | Double | Settings | Evita Infinity V500 [MEDIBUS.X]
1656 | SmartCare Patient Height | cm | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. Patient height in SmartCare, SC-Ht. For Evita Infinity V500: SC-Body height. | Double | Settings | Evita Infinity V500 [MEDIBUS.X]
1761 | Airway Resistance | mbar/L/s | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. | Double | VitalSigns | Evita Infinity V500 [MEDIBUS.X]
2044 | P01 Repetition Interval Setting | min | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. P0.1 repetition interval. | Double | Settings | Evita Infinity V500 [MEDIBUS.X]
2092 | Inspired Minute Volume | L/min | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Carina, Evita V300. This variable is associated with UpperAlarmLimit attribute, which is available for Carina. This variable is associated with LowerAlarmLimit attribute, which is available for Carina. | Double | VitalSigns | Evita Infinity V500 [MEDIBUS.X]
2102 | Inspiratory Time Setting | s | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300, Fabius Family. | Double | Settings | Evita Infinity V500 [MEDIBUS.X]
2104 | Positive End Expiratory Pressure (PEEP) Setting | mbar | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300, Fabius Family. For Evita Infinity V500, Babylog VN500: PEEP / CPAP PEEP. For Zeus, Zeus IE: PEEP (CPAP). PEEP for mode Man/Spont. | Double | Settings | Evita Infinity V500 [MEDIBUS.X]
2110 | Inspired O2 (FiO2) Setting | % | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Zeus IE, Oxylog 3000 plus, Carina, Evita V300. For Primus, Primus IE, Apollo, Pallas: Not in device Apollo and Pallas | Double | Settings | Evita Infinity V500 [MEDIBUS.X]
2217 | Manual Breath Ventilation Pressure Setting | mbar | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300(added with SW 2.30 for Babylog VN500, added with SW 2.30 for Evita Infinity V500). Manual inspiratory pressure for Neonates during NIV - Non-Invasive Ventilation, PmanInsp. | Double | Settings | Evita Infinity V500 [MEDIBUS.X]
2245 | Manual Breath Maximum Time Limit | s | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300(added with SW 2.30 for Babylog VN500, added with SW 2.30 for Evita Infinity V500). Maximum time of pressure for manual inspiration pressure for Neonates during NIV - Non-Invasive Ventilation, TmanInsp. | Double | Settings | Evita Infinity V500 [MEDIBUS.X]
2250 | CO2 Tidal Production | mL | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. CO2 production volume per breath, VTCO2. For Evita Infinity V500, Babylog VN500: Supported through Medibus.X protocol. | Double | VitalSigns | Evita Infinity V500 [MEDIBUS.X]
2251 | CO2 Minute Production | mL/min | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Evita V300. CO2 production, V'CO2. | Double | VitalSigns | Evita Infinity V500 [MEDIBUS.X]
2255 | Device Date/Time | unitless | Format: YYYYMMDDhhmmss, where ss is always present but might be 00. | String | DeviceInfo | Evita Infinity V500 [MEDIBUS.X] | Device Information
2324 | Inspired Tidal Volume | mL | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Carina, Evita V300. This variable is associated with UpperAlarmLimit attribute, which is available for Evita Infinity V500, Zeus IE, Evita V300. This variable is associated with LowerAlarmLimit attribute, which is available for Evita Infinity V500, Evita V300. For Primus, Primus IE, Apollo, Pallas: Insp. Tidal Volume. | Double | VitalSigns | Evita Infinity V500 [MEDIBUS.X]
2325 | Expired Tidal Volume | mL | For Evita XL: Neo expired tidal volume. Available only with firmware version 7.01 or higher. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Evita V300. | Double | VitalSigns | Evita Infinity V500 [MEDIBUS.X]
2326 | Expired Minute Volume | L/min | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. This variable is associated with UpperAlarmLimit attribute, which is available for Evita Infinity V500, Babylog VN500, Evita V300. This variable is associated with LowerAlarmLimit attribute, which is available for Evita Infinity V500, Babylog VN500, Evita V300. For Evita Infinity V500, Babylog VN500: This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. | Double | VitalSigns | Evita Infinity V500 [MEDIBUS.X]
2355 | End Tidal CO2 Percentage | % | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Evita V300, Fabius Family. This variable is associated with UpperAlarmLimit attribute, which is available for Perseus A500, Atlan, Evita Infinity V500, Babylog VN500, Savina 300, Primus, Primus IE, Apollo, Zeus IE, Oxylog 3000 plus, Evita V300. This variable is associated with LowerAlarmLimit attribute, which is available for Perseus A500, Atlan, Evita Infinity V500, Babylog VN500, Savina 300, Primus, Primus IE, Apollo, Zeus IE, Oxylog 3000 plus, Evita V300. end tidal CO2 etCO2 | Double | VitalSigns | Evita Infinity V500 [MEDIBUS.X]
2377 | Fresh Gas O2 Flow Rate | L/min | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Evita V300. O2 therapy flow, Flow. For Evita Infinity V500, Babylog VN500: Flow O2 Therapy Flow. | Double | DeviceInfo | Evita Infinity V500 [MEDIBUS.X]
2382 | Ventilator Operation Mode | unitless | This variable contains ventilation operation mode in string format, which is also send as a numerical value in variable 2758. See detailed description of variable 2758 for the list of supported values. For Evita 4, Evita 2 dura, Evita XL: Firmware version 04.00 or higher For Evita 2: Variable label is not correct. Please use variable 3853 instead. Obsolete variable that may be removed in future release. | String | DeviceInfo | Evita Infinity V500 [MEDIBUS.X]
2416 | Spontaneous Inspired Tidal Volume | mL | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. For Zeus, Zeus IE: This variable is associated with UpperAlarmLimit attribute. Insp. Tidal volume. | Double | VitalSigns | Evita Infinity V500 [MEDIBUS.X]
2421 | Airway Dead Space | mL | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. Serial dead space volume, Vds. For Evita Infinity V500, Babylog VN500: Serial dead space Vds. | Double | VitalSigns | Evita Infinity V500 [MEDIBUS.X]
2423 | Dynamic Compliance Over Last 20% Breath | unitless | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. Ratio of compliance during the last 20 % of inspiration over dynamic compliance, C20/Cdyn. For Evita Infinity V500, Babylog VN500: C20/Cdyn. | Double | VitalSigns | Evita Infinity V500 [MEDIBUS.X]
2432 | Expired Spontaneous Tidal Volume | mL | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. | Double | VitalSigns | Evita Infinity V500 [MEDIBUS.X]
2433 | Mechanical Expired Minute Volume | L/min | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Primus, Primus Infinity Empowered, Apollo, Evita V300. Mandatory expiratory minute volume, MVemand | Double | VitalSigns | Evita Infinity V500 [MEDIBUS.X]
2434 | Spontaneous Expired Minute Volume | L/min | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Primus, Primus Infinity Empowered, Apollo, Evita V300. | Double | VitalSigns | Evita Infinity V500 [MEDIBUS.X]
2467 | Patient Type ID | unitless | Consult detailed help for more information on this variable | UByte | PatientInfo | Evita Infinity V500 [MEDIBUS.X]
2542 | Airway Pressure Wave | mbar | For Fabius GS, Fabius GS premium, Fabius Tiro, Fabius plus, Fabius plus XL, Fabius OS, Fabius MRI: 1mbar = 1cmH2O. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300, Fabius Family. For Evita Infinity V500, Babylog VN500: Airway pressure Paw. | Double | Waveforms | Evita Infinity V500 [MEDIBUS.X]
2564 | Peak Airway Pressure | mbar | For Evita 4, Evita 2 dura, Evita XL: This variable is associated with UpperAlarmLimit attribute, which is also sent in separate variable 7040. For Evita 2, Evita, Savina: This variable is associated with UpperAlarmLimit attribute. For Fabius GS, Fabius GS premium, Fabius Tiro, Fabius plus, Fabius plus XL, Fabius OS, Fabius MRI: 1mbar = 1cmH2O. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300, Fabius Family. This variable is associated with UpperAlarmLimit attribute, which is available for Perseus A500, Atlan, Evita Infinity V500, Babylog VN500, Savina 300, Zeus IE, Oxylog 3000 plus, Carina, Evita V300, Fabius Family. This variable is associated with LowerAlarmLimit attribute, which is available for Zeus IE, Fabius Family. Peak inspiratory pressure, PIP. For Evita Infinity V500, Babylog VN500, Carina: This variable is associated with UpperAlarmLimit attribute. For Zeus, Zeus IE: This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. Peak Breathing Press. For Primus, Primus IE, Apollo, Pallas: Peak Breathing Pressure. | Double | VitalSigns | Evita Infinity V500 [MEDIBUS.X]
2565 | Plateau Airway Pressure | mbar | For Fabius GS, Fabius GS premium, Fabius Tiro, Fabius plus, Fabius plus XL, Fabius OS, Fabius MRI: 1mbar = 1cmH2O. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Evita V300, Fabius Family. For Evita Infinity V500, Babylog VN500: Only available for Evita Infinity V500 | Double | VitalSigns | Evita Infinity V500 [MEDIBUS.X]
2570 | Spontaneous Tidal Volume | - | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Zeus IE, Evita V300. Possible units: L, mL. For Evita Infinity V500, Babylog VN500: Possible units: L, mL. | Double (Multiset) | VitalSigns | Evita Infinity V500 [MEDIBUS.X]
2583 | I:E Ratio Setting | unitless | String format: 1:x if E > I, x:1 if I > E, or 1:1 if I = E. The useless leading or ending “0” or “.” is not transmitted, and the maximum digits after the floating point is 2. For Carina, Primus, Primus IE, Apollo, Pallas, Perseus A500: This variable is calculated by the DDI from respiration rate and inspiratory time settings. For Medibus.X compatible device: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300, Fabius Family. For Primus, Primus IE, Apollo, Perseus A500, Carina and Oxylog 3000 plus, it is calculated by the DDI from respiration rate and inspiratory time settings. For Atlan, Evita Infinity V500, Babylog VN500, Savina 300, Zeus IE, Evita V300, and Fabius Family, it is calculated by the DDI from inspiratory and expiratory components. For Evita, Evita 2, Evita 4, Evita 2 dura, and Evita XL, Fabius GS, Fabius GS premium, Fabius Tiro, Fabius plus, Fabius plus XL, Fabius OS, Fabius MRI, Evita Infinity V500, Babylog VN500, Zeus IE, Evita V300: This variable is calculated by the DDI from inspiratory and expiratory components of I:E ratio sent by the devices. | String | Settings | Evita Infinity V500 [MEDIBUS.X]
2596 | Alarm Severity Level | unitless | This variable is in multiset format. This variable contains alarm priority level, variable 6164 contains alarm codes, and variable 3878 contains alarm messages. They need to be interpreted together. One alarm code may have one or more priority values and one or more alarm messages associated with it. The possible alarms are depended on device firmware version. Please see device user manuals for details. See variables 177 and 588 and 4131 for string format. | UByte (Multiset) | Alarms | Evita Infinity V500 [MEDIBUS.X] | Alarm And Text Message
2754 | Ventilator General Status | unitless | Ventilator Standby mode status. | UByte | DeviceInfo | Evita Infinity V500 [MEDIBUS.X]
Consult detailed help for more information on this variable
2756 | Ventilation Mode ID | unitless | This variable contains ID of basic ventilation mode ID without options. See also variable 584. | UByte | DeviceInfo | Evita Infinity V500 [MEDIBUS.X]
Consult detailed help for more information on this variable
2758 | Ventilator Operation Mode ID | unitless | For Evita 4, Evita 2 dura, Evita XL: Firmware version 04.00 or higher For Evita 2: Variable label is not correct. Please use variable 2467 instead. Obsolete variable that may be removed in future release. | UByte | DeviceInfo | Evita Infinity V500 [MEDIBUS.X]
Consult detailed help for more information on this variable
2771 | Tidal Volume Setting | L | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300, Fabius Family. Inspiratory tidal volume, Vti. For Evita Infinity V500, Babylog VN500: See also variable 1321 and 5044. For Zeus, Zeus IE: This variable contains same information as variable 1321, except that the value of different units (L or mL) are sent in separate variables. See also variable 5044. | Double | Settings | Evita Infinity V500 [MEDIBUS.X]
2784 | ASB Ramp Time Setting | s | For Evita 2: Firmware version 2.00 or higher. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300. Pressure rise time, Slope. For Carina: ASB ramp. For Primus, Primus IE, Apollo, Pallas: Slope time (Ramp). | Double | Settings | Evita Infinity V500 [MEDIBUS.X]
2952 | PS Activation State | unitless | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300, Fabius Family. A value of 1 indicates that Pressure Support (PS) is added to the current ventilation mode. A value of 0 indicates that Pressure Support (PS) is not added to the current ventilation mode. Text message displayed on device screen: "/PS" | Boolean | DeviceInfo | Evita Infinity V500 [MEDIBUS.X]
2953 | VS Activation State | unitless | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. A value of 1 indicates that Volume support (VS) is added to the current ventilation mode. A value of 0 indicates that Volume support (VS) is not added to the current ventilation mode. Text messages displayed on device screen: "/VS" | Boolean | DeviceInfo | Evita Infinity V500 [MEDIBUS.X]
2954 | APRV High Pressure (Phigh) | mbar | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. Upper pressure level during APRV, Phigh. | Double | VitalSigns | Evita Infinity V500 [MEDIBUS.X]
2959 | APRV Low Pressure (Plow) | mbar | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. Lower pressure level during APRV, Plow. | Double | VitalSigns | Evita Infinity V500 [MEDIBUS.X]
2987 | Inspired Pressure Setting | mbar | For Evita 4, Evita 2 dura, Evita XL: Available only with EvitaXL firmware version 7.00 or higher. For Fabius GS, Fabius GS premium, Fabius Tiro, Fabius plus, Fabius plus XL, Fabius OS, Fabius MRI: 1mbar = 1cmH2O. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300, Fabius Family. | Double | Settings | Evita Infinity V500 [MEDIBUS.X]
3326 | Device Average Gas Flow | L/min | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. Average device flow, FlowDev. | Double | OtherMeasuredData | Evita Infinity V500 [MEDIBUS.X]
3503 | Auto Flow Active Status | unitless | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300. A value of 1 indicates that AutoFlow (AF) is added to the current ventilation mode. A value of 0 indicates that AutoFlow (AF) is not added to the current ventilation mode. Text message displayed on device screen: "/AF" | Boolean | DeviceInfo | Evita Infinity V500 [MEDIBUS.X]
3555 | ATC Activation Status | unitless | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Evita V300. A value of 1 indicates that automatic tube compensation (ATC) is added to the current ventilation mode. A value of 0 indicates that automatic tube compensation (ATC) is not added to the current ventilation mode. Text message displayed on device screen: "/ATC" For Savina 300 (MEDIBUS.X): added with SW 5.00 | Boolean | DeviceInfo | Evita Infinity V500 [MEDIBUS.X]
3591 | Device Number | unitless | Consult detailed help for more information on this variable | String | DeviceInfo | Common | Device Information
3593 | Device Communications Status | unitless | Value "stop" sent when a request to stop the communication is received from the device | String | DeviceInfo | Common | Device Information
3853 | Patient Type | unitless | This variable contains patient type in string format, which is also send as a numerical value in variable 2467. See detailed description of variable 2467 for the list of supported values. | String | PatientInfo | Evita Infinity V500 [MEDIBUS.X]
3878 | Generic Alarm Message | unitless | This variable is in multiset format. This variable contains alarm messages, variable 2596 contains alarm priority level, variable 6164 contains alarm codes. They need to be interpreted together. One alarm code may have one or more priority values and one or more alarm messages associated with it. The possible alarms are dependent on device firmware version. Please see device user manuals for details. See variables 177 and 588 and 4131 for string format. The HL7 output can contain an '&' character to indicate a LOW alarm. In order to avoid issues, the '&' character will be replaced by a '?' character in the DDI output. | String (Multiset) | Alarms | Evita Infinity V500 [MEDIBUS.X] | Alarm And Text Message
3909 | Message Code | unitless | This variable, along with variable 1197 are marked as obsolete because they use either"¦¦" or "," in the array string as delimiter, which are often not compatible with customers' applications. Variables 8366 and 8696 should be used instead. | String | Messages | Evita Infinity V500 [MEDIBUS.X] | Alarm And Text Message
3966 | Inspired Mandatory Tidal Volume | mL | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. | Double | VitalSigns | Evita Infinity V500 [MEDIBUS.X]
3968 | Spontaneous Fraction of Minute Volume | % | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. Spontaneous breathing portion of minute volume, % Mvspon. For Evita Infinity V500, Babylog VN500: Spontaneous Fraction Min. Vol. %MVspon (MVspon : MVtotal). | Double | VitalSigns | Evita Infinity V500 [MEDIBUS.X]
3971 | Negative Inspiratory Force | mbar | For Evita 4, Evita 2 dura, Evita XL: The value is not transmitted, if it is displayed on Evita 4 / Evita 2 dura as "[less than] –45 mbar". Firmware version 04.00 or higher For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. | Double | VitalSigns | Evita Infinity V500 [MEDIBUS.X]
3998 | Auto Release Activation Status | unitless | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. A value of 1 indicates that AutoRelease is added to the current ventilation mode. A value of 0 indicates that AutoRelease is not added to the current ventilation mode. Text messages displayed on device screen: "/AutoRelease" | Boolean | DeviceInfo | Evita Infinity V500 [MEDIBUS.X]
4003 | Generic Respiration Rate | 1/min | For Evita 4, Evita 2 dura, Evita XL: This variable is associated with UpperAlarmLimit attribute, which is also sent in separate variable 1339. Total Respiratory Rate (RRtotal). For Evita 2, Evita, Savina, Carina: This variable is associated with UpperAlarmLimit attribute. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300, Fabius Family. This variable is associated with UpperAlarmLimit attribute, which is available for Evita Infinity V500, Babylog VN500, Savina 300, Carina, Evita V300. Respiratory rate, RR. For Evita Infinity V500, Babylog VN500: This variable is associated with UpperAlarmLimit attribute. Respiratory Rate (RRtotal) RR. | Double | VitalSigns | Evita Infinity V500 [MEDIBUS.X]
4010 | Minute Volume Alarm Delay | s | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. Delay for minute volume high alarm, MVhigh delay. This variable is associated with a ChannelID identifier. | Double | Settings | Evita Infinity V500 [MEDIBUS.X]
4011 | Flow Assist Level Setting | 0.1 mbar/L/s | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. Flow assist during PPS ventilation, Flow Assist. For Evita Infinity V500, Babylog VN500: Flow Assist. | Double | Settings | Evita Infinity V500 [MEDIBUS.X]
4012 | Volume Assist Level Setting | 0.1 mbar/L | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. Volume assist during PPS ventilation, Vol. Assist. For Evita Infinity V500, Babylog VN500: Vol. Assist. | Double | Settings | Evita Infinity V500 [MEDIBUS.X]
4015 | Disconnection Alarm Delay Time | s | For Evita 4, Evita 2 dura, Evita XL: Firmware version 4.00 or higher. For Savina: Tdeconnect. Available since Device SW version 02.00. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Carina, Evita V300. Disconnection alarm time, Tdiscon. For Evita Infinity V500, Babylog VN500: Tdisconnect. For Carina: Disconnection detection timeout. | Double | Settings | Evita Infinity V500 [MEDIBUS.X]
4044 | Medibus Revision | unitless |  | String | DeviceInfo | Common | Device Information
4045 | Device Revision | unitless | Format: mm.nn | String | DeviceInfo | Common | Device Information
4131 | Alarm Priority | unitless | This variable is same as 2596 but sent as a string, which is an array containing all alarm priorities and separated by "," delimiter. This variable, along with variables 177 and 588, are marked as obsolete because they use either"¦¦" or "," in the array string as delimiter, which are often not compatible with customers' applications. Variables 6164, 3878, and 2596 should be used instead. | String | Alarms | Evita Infinity V500 [MEDIBUS.X] | Alarm And Text Message
4150 | VG Activation Status | unitless | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Carina, Evita V300. A value of 1 indicates that Volume Guarantee (VG) is added to the current ventilation mode. A value of 0 indicates that Volume Guarantee (VG) is not added to the current ventilation mode. Text message displayed on device screen: "/VG" | Boolean | DeviceInfo | Evita Infinity V500 [MEDIBUS.X]
4732 | Inspiratory Flow Waveform | L/min | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. | Double | Waveforms | Evita Infinity V500 [MEDIBUS.X]
4889 | Spontaneous Inspiration Time | s | For Evita 4, Evita 2 dura, Evita XL: Available only with EvitaXL firmware version 7.02 or higher. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. | Double | VitalSigns | Evita Infinity V500 [MEDIBUS.X]
4891 | Lung Time Constant | s | For Evita 4, Evita 2 dura, Evita XL: Available only with EvitaXL firmware version 7.00 or higher. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. | Double | VitalSigns | Evita Infinity V500 [MEDIBUS.X]
4892 | Airway Correlation Coefficient | unitless | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. Correlation factor, r2. For Evita Infinity V500, Babylog VN500: r². | Double | VitalSigns | Evita Infinity V500 [MEDIBUS.X]
4966 | Sigh Activation Status | unitless | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. A value of 1 indicates that the sigh function has been added to the current ventilation mode. A value of 0 indicates that the sigh function has not been added to the current ventilation mode. Text message displayed on device screen: "/Sigh" | Boolean | DeviceInfo | Evita Infinity V500 [MEDIBUS.X]
5044 | Inspired Tidal Volume Setting | mL | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Zeus IE, Evita V300. Inspiratory tidal volume, Vti. For Evita Infinity V500, Babylog VN500: See also variable 1321 and 2771. For Zeus, Zeus IE: This variable contains same information as variable 1321, except that the value of different units (L or mL) are sent in separate variables. See also variable 2771. | Double | Settings | Evita Infinity V500 [MEDIBUS.X]
5087 | Dead Space to Tidal Volume Ratio | % | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. Ratio of serial dead space volume to expiratory tidal volume, Vds/VTe. For Evita Infinity V500, Babylog VN500: Dead space ventilation Vds/VTe. | Double | VitalSigns | Evita Infinity V500 [MEDIBUS.X]
5340 | Expired Mandatory Tidal Volume | mL | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. | Double | VitalSigns | Evita Infinity V500 [MEDIBUS.X]
5374 | Expiratory Time Constant (RCexp) | s | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. | Double | VitalSigns | Evita Infinity V500 [MEDIBUS.X]
5728 | EtCO2 Phase III Increment | - | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. Possible units: mmHg/L, kPa/L, Vol%/L. Increase of measured CO2 value in phase III of the capnogram, CO2 slope. | Double (Multiset) | VitalSigns | Evita Infinity V500 [MEDIBUS.X]
5815 | Spontaneous I:E Ratio | unitless | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. Format: 1:x if E > I, x:1 if I > E, or 1:1 if I = E. The useless leading or ending “0” or “.” is not transmitted, and the maximum digits after the floating point is 2. For Evita Infinity V500, Babylog VN500: Format: 1:x if E > I, x:1 if I > E, or 1:1 if I = E. The useless leading or ending “0” or “.” is not transmitted, and the maximum digits after the floating point is 2. | String | VitalSigns | Evita Infinity V500 [MEDIBUS.X]
6164 | General Alarm Code | unitless | This variable is in multiset format. This variable contains alarm codes, variable 3878, alarm messages, variable 2596 contains alarm priority level. They need to be interpreted together. One alarm code may have one or more priority values and one or more alarm messages associated with it. The possible alarms are depended on device firmware version. Please see device user manuals for details. See variables 177 and 588 and 4131 for string format. | UShort (Multiset) | Alarms | Evita Infinity V500 [MEDIBUS.X] | Alarm And Text Message
Consult detailed help for more information on this variable
6389 | Endotracheal Tube Size | mm | For Evita 4, Evita 2 dura, Evita XL: Possible units: "0.1mm" in firmware versions 7.00 and 7.01, "mm" otherwise. Available only with EvitaXL firmware version 7.00 or higher. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Evita V300. Inner diameter of the endotracheal/tracheostoma tube. For Evita Infinity V500, Babylog VN500: ATC Tube Diameter. | Double | DeviceInfo | Evita Infinity V500 [MEDIBUS.X]
6551 | End Tidal CO2 Pressure | - | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Evita V300, Fabius Family. This variable is associated with UpperAlarmLimit attribute, which is available for Perseus A500, Atlan, Evita Infinity V500, Babylog VN500, Savina 300, Primus, Primus IE, Apollo, Zeus IE, Oxylog 3000 plus, Evita V300. This variable is associated with LowerAlarmLimit attribute, which is available for Perseus A500, Atlan, Evita Infinity V500, Babylog VN500, Savina 300, Primus, Primus IE, Apollo, Zeus IE, Oxylog 3000 plus, Evita V300. Possible units: mmHg, kPa. end tidal CO2 etCO2 | Double | VitalSigns | Evita Infinity V500 [MEDIBUS.X]
6701 | Expiratory Time Setting | s | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Zeus IE, Evita V300. | Double | Settings | Evita Infinity V500 [MEDIBUS.X]
6702 | Inspiratory Component of I:E Ratio Setting | unitless | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Zeus IE, Evita V300, Fabius Family, Atlan For Evita Infinity V500: Partial information. Please use variable 2583 instead. Obsolete variable that may be removed in future release. | Double | Settings | Evita Infinity V500 [MEDIBUS.X]
6703 | Expiratory Component of I:E Ratio Setting | unitless | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800,Savina 300, Zeus IE, Evita V300, Fabius Family, Atlan For Evita Infinity V500: Partial information. Please use variable 2583 instead. Obsolete variable that may be removed in future release. | Double | Settings | Evita Infinity V500 [MEDIBUS.X]
6735 | Lung Elastance | mbar/L | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. | Double | VitalSigns | Evita Infinity V500 [MEDIBUS.X]
6736 | Rapid Shallow Breathing Index (SBI) | - | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: Obsolete variable that may be removed in future release. Variable is in Multiset format. This variable is same as 1422 but DDI buffer and send all values from different data packages together. | Double (Multiset) | VitalSigns | Evita Infinity V500 [MEDIBUS.X]
7076 | Generic Ventilation Mode | unitless | Consult detailed help for more information on this variable | String | Settings | Evita Infinity V500 [MEDIBUS.X]
7077 | Ventilation Mode Primary | unitless | Consult detailed help for more information on this variable | String | Settings | Evita Infinity V500 [MEDIBUS.X]
7078 | Ventilation Mode Options | unitless | Consult detailed help for more information on this variable | String | Settings | Evita Infinity V500 [MEDIBUS.X]
7318 | Variable PS Activation Status | unitless | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800. A value of 1 indicates that Variable PS (VPS) has been added to the current ventilation mode. A value of 0 indicates that Variable PS (VPS) has not been added to the current ventilation mode. Text message displayed on device screen: "VariablePS" | Boolean | Settings | Evita Infinity V500 [MEDIBUS.X]
7364 | Patient Operation Mode | unitless | Possible values: 1 for Invasive, 2 for Noninvasive | UByte | PatientInfo | Evita Infinity V500 [MEDIBUS.X]
7637 | Artificial Airway Compensation Status | % | For Evita 4, Evita 2 dura, Evita XL: Tube Compensation. Available only with EvitaXL firmware version 7.00 or higher. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. Compensation rate during ATC, % comp. For Evita Infinity V500, Babylog VN500: ATC Compensation Compens. | Double | DeviceInfo | Evita Infinity V500 [MEDIBUS.X]
7640 | Mandatory Respiration Rate | 1/min | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Zeus IE, Evita V300. | Double | VitalSigns | Evita Infinity V500 [MEDIBUS.X]
7642 | Mandatory Tidal Volume | - | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Babylog VN500, Babylog VN600, Babylog VN800, Zeus IE, Evita V300(added with SW 2.30 for Babylog VN500, added with SW 2.30 for Evita Infinity V500). Possible units: L, mL. Unit "L" available for Evita Infinity V500, Babylog VN500, Evita V300 (added with SW 2.30 for Babylog VN500, added with SW 2.30 for Evita Infinity V500). Unit "mL" available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Zeus IE, Evita V300. | Double (Multiset) | VitalSigns | Evita Infinity V500 [MEDIBUS.X]
7872 | CO2 Partial Pressure Waveform | - | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Evita V300. Possible units: mmHg, kPa. Unit "mmHg" is available for Perseus A500, Atlan, Evita Infinity V500, Babylog VN500, Savina 300, Primus, Primus IE, Apollo, Zeus IE, Oxylog 3000 plus, Evita V300. Unit "kPa" is available for Perseus A500, Atlan, Savina 300, Primus, Primus IE, Apollo, Zeus IE, Oxylog 3000 plus. | Double | Waveforms | Evita Infinity V500 [MEDIBUS.X]
7879 | Time Duration at Low Pressure (Tlow) | s | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. Effective expiratory time during APRV/AutoRelease, Tlow. For Evita Infinity V500, Babylog VN500: Tlow (resulting Tlow from expiratory flow cycling). | Double | VitalSigns | Evita Infinity V500 [MEDIBUS.X]
7904 | Spontaneous Tidal Volume | - |  | Double (Multiset) | VitalSigns | Evita Infinity V500 [MEDIBUS.X]
8366 | Generic Message Code | unitless | This variable is in multiset format. This variable, along with variable 8696, contains text message code and alarm text respectively. They need to be interpreted together. | UShort (Multiset) | Messages | Evita Infinity V500 [MEDIBUS.X] | Alarm And Text Message
8696 | Generic Message Text | unitless | This variable is in multiset format. This variable, along with variable 8366, contains text message code and alarm text respectively. They need to be interpreted together. | String (Multiset) | Messages | Evita Infinity V500 [MEDIBUS.X] | Alarm And Text Message
9146 | End Inspiratory Pressure | mbar | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. End-inspiratory pressure, EIP. | Double | VitalSigns | Evita Infinity V500 [MEDIBUS.X]
9163 | Tracheal Pressure Waveform | mbar | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Evita V300. For Evita Infinity V500, Babylog VN500: Ptrach. | Double | Waveforms | Evita Infinity V500 [MEDIBUS.X]
9288 | Tidal Volume Indexed By Body Weight | mL/kg | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. Tidal volume per kg body weight, VT/Wt. The leakage is determined by the difference between the delivered flow on the inspiratory side and the measured flow on the expiratory side, which is displayed as the leakage minute volume MVleak and relative as % leak (MVleak to MV). For Evita Infinity V500, Babylog VN500: VT/kg BW. | Double | VitalSigns | Evita Infinity V500 [MEDIBUS.X]
9466 | Mandatory Triggered Respiration Rate | 1/min | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300(added with SW 2.42 for Evita Infinity V500, Evita V300 and Babylog VN500). Respiratory rate of triggered, mandatory breaths, RRtrig For Evita 4, Evita 2 dura, Evita XL: Available only with Evita XL firmware version 7.00 or higher. | Double | VitalSigns | Evita Infinity V500 [MEDIBUS.X]
9475 | Volume of Leakage (MVLeak) | L/min | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Carina, Evita V300. Leakage minute volume, Mvleak. The leakage is determined by the difference between the delivered flow on the inspiratory side and the measured flow on the expiratory side, which is displayed as the leakage minute volume MVleak and relative as % leak (MVleak to MV). For Evita Infinity V500, Babylog VN500: Mvleak. The leakage is determined by the difference between the delivered flow on the inspiratory side and the measured flow on the expiratory side, which is displayed as the leakage minute volume MVleak and relative as % leak (MVleak to MV). For Carina: Leakage minute volume. For Evita XL: Leakage minute volume. Available since device SW version 7.02. | Double | OtherMeasuredData | Evita Infinity V500 [MEDIBUS.X]
9476 | Percent of Leakage | % | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Evita V300. Leakage minute volume in % of inspiratory minute volume, % leak Leakage rel. % leak - Infinity V500 determines the difference between the delivered flow on the inspiratory side and the measured flow on the expiratory side. This difference provides a measure of the amount of leakage and is displayed by Infinity V500 as the leakage minute volume MVleak and relative as % leak (MVleak to MV). For Evita Infinity V500, Babylog VN500: Leakage rel. % leak. | Double | OtherMeasuredData | Evita Infinity V500 [MEDIBUS.X]
9477 | Inspiratory Component of Spontaneous I:E Ratio | unitless | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: Obsolete variable that may be removed in future release. Partial information of I:E ratio. For Evita Infinity V500: Partial information. Please use variable 5815 instead. Obsolete variable that may be removed in future release. | Double | VitalSigns | Evita Infinity V500 [MEDIBUS.X]
9478 | Expiratory Component of Spontaneous I:E Ratio | unitless | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: Obsolete variable that may be removed in future release. Partial information of I:E ratio. For Evita Infinity V500: Partial information. Please use variable 5815 instead. Obsolete variable that may be removed in future release. | Double | VitalSigns | Evita Infinity V500 [MEDIBUS.X]
9479 | SmartCare Target Delta Support Pressure | mbar | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. Dpressure support goal in SmartCare, SC-DPsupp goal. For Evita Infinity V500: SC-Target ΔPsupp ΔPsupp goal. | Double | VitalSigns | Evita Infinity V500 [MEDIBUS.X]
9480 | SmartCare Rated Delta Support Pressure | mbar | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. Rated Dpressure support by internal controller in SmartCare, SC-DPsupp rated. For Evita Infinity V500: SC-Rated ΔPsupp SC-Δpsupp. | Double | VitalSigns | Evita Infinity V500 [MEDIBUS.X]
9481 | SmartCare Duration of Patient Session | - | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. Duration of session (hours and minutes) in SmartCare, SC-duration. Variable is in Multiset format (need combine the first value in hours with the second value in minutes to obtain the exact duration). Please use variable 4346 instead. For Evita Infinity V500: Multiset variable (need combine the first value in hours with the second value in minutes to obtain the exact duration). | Double (Multiset) | VitalSigns | Evita Infinity V500 [MEDIBUS.X]
9482 | SmartCare Spontaneous Respiration Rate | 1/min | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. Spontaneous respiratory rate in SmartCare, SC-Rrspon. For Evita Infinity V500: SC-Rrspon. | Double | VitalSigns | Evita Infinity V500 [MEDIBUS.X]
9483 | SmartCare Tidal Volume | mL | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. Tidal volume in SmartCare, SC-VT. For Evita Infinity V500: SC-VT. | Double | VitalSigns | Evita Infinity V500 [MEDIBUS.X]
9484 | SmartCare EtCO2 | mmHg | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. End-tidal CO2 concentration in SmartCare, SC-etCO2. For Evita Infinity V500: SC-etCO2. | Double | VitalSigns | Evita Infinity V500 [MEDIBUS.X]
9485 | Maximum Inspiratory Time Setting | s | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Carina, Evita V300. Maximal inspiratory time for flow cycled strokes, Timax. For Carina: Maximal inspiration time for CPAP, ASB. | Double | Settings | Evita Infinity V500 [MEDIBUS.X]
9486 | Inspiratory Flow Termination Level Setting | % | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Savina 300, Oxylog 3000 plus, Evita V300. Inspiratory termination criterion based on peak inspiratory flow, Insp. term. For Evita Infinity V500: Inspiratory termination. | Double | Settings | Evita Infinity V500 [MEDIBUS.X]
9487 | Expiratory Flow Termination Setting | % | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. Expiratory termination criterion during APRV/AutoRelease, Exp. term. For Evita Infinity V500, Babylog VN500: Expiratory termination. | Double | Settings | Evita Infinity V500 [MEDIBUS.X]
9488 | SmartCare Body Weight | kg | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. Ideal body weight in SmartCare, SC-Wt. For Evita Infinity V500: SC-Body weight. | Double | Settings | Evita Infinity V500 [MEDIBUS.X]
9489 | SmartCare Start of Night Rest | h | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. Start of night rest in SmartCare, SC-startrest. For Evita Infinity V500: SC-Start of night rest. | Double | Settings | Evita Infinity V500 [MEDIBUS.X]
9490 | SmartCare End of Night Rest | h | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. End of night rest in SmartCare, SC-endrest. For Evita Infinity V500: SC-End of night rest. | Double | Settings | Evita Infinity V500 [MEDIBUS.X]
50021 | Device Alarm Inactivation State | unitless |  | String | DeviceInfo | Evita Infinity V500 [MEDIBUS.X]

## Evita V800
ID | Label | Unit | Description | Type | Group | Device | Source
20 | End Tidal CO2 (EtCO2) | - | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Evita V300, Fabius Family. This variable is associated with UpperAlarmLimit attribute, which is available for Perseus A500, Atlan, Evita Infinity V500, Babylog VN500, Savina 300, Primus, Primus IE, Apollo, Zeus IE, Oxylog 3000 plus, Evita V300. This variable is associated with LowerAlarmLimit attribute, which is available for Perseus A500, Atlan, Evita Infinity V500, Babylog VN500, Savina 300, Primus, Primus IE, Apollo, Zeus IE, Oxylog 3000 plus, Evita V300. Possible units: mmHg, %, kPa. This variable is in Multiset. See variables 2355 and 6551 for concentration and partial pressure respectively. For Evita Infinity V500, Babylog VN500: This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. end tidal CO2 etCO2 Note: etCO2 values and alarm limits are transmitted according to the configuration on the device screen (mmHg or Vol%). If kPa is configured be aware that the Medibus alarm limits will be transmitted with mmHg. | Double (Multiset) | VitalSigns | Evita Infinity V800
60 | Tidal Volume | mL | For Evita 4, Evita 2 dura, Evita XL: Expiratory tidal volume. This variable is associated with UpperAlarmLimit attribute, which is also sent in separate variable 2763. Values below 1 L will be transmitted in units of mL, otherwise in units of L. Device firmware version 04.00 and higher: Limit is not transmitted, if alarm is switched off. For Savina: This variable is associated with UpperAlarmLimit attribute.Device SW version 02.00 and higher: Limit is not transmitted, if alarm is switched off. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300, Fabius Family. This variable is associated with UpperAlarmLimit attribute, which is available for Perseus A500, Evita Infinity V500, Savina 300, Evita V300. This variable is associated with LowerAlarmLimit attribute, which is available for Evita Infinity V500, Evita V300. For Evita Infinity V500, Babylog VN500: This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. For Zeus, Zeus IE, Primus, Primus IE, Apollo, Pallas: Exp. Tidal Volume. | Double | VitalSigns | Evita Infinity V800
61 | Dynamic Compliance | L/bar | For Evita 4, Evita 2 dura, Evita XL: Values below 100 will be transmitted with one decimal place, otherwise without a decimal place. For Savina: Available since device SW version 02.00. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Evita V300. | Double | VitalSigns | Evita Infinity V800
114 | CO2 Concentration Waveform | - | For Evita 4, Evita 2 dura, Evita XL: Possible units: mmHg, kPa, and % For Evita 2: Possible units: mmHg, kPa, and % Firmware version 2.00 or higher. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Evita V300. Possible units: mmHg, %, kPa. Unit "mmHg" is available for Perseus A500, Atlan, Evita Infinity V500, Babylog VN500, Savina 300, Primus, Primus IE, Apollo, Zeus IE, Oxylog 3000 plus, Evita V300. Unit "kPa" and "%" are available for Perseus A500, Atlan, Savina 300, Primus, Primus IE, Apollo, Zeus IE, Oxylog 3000 plus. See variables 8866 and 7872 for concentration and partial pressure respectively. For Evita Infinity V500, Babylog VN500: Possible units: mmHg, %. Expiratory CO2 concentration. For Zeus, Zeus IE, Primus, Primus IE, Apollo, Pallas, Vamos, Vamos plus: Possible units: mmHg, %, kPa. | Double | Waveforms | Evita Infinity V800
118 | Airway Flow Wave | L/min | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300, Fabius Family. For Evita Infinity V500, Babylog VN500, Carina, Zeus, Zeus IE, Primus, Primus IE, Apollo, Pallas: Flow (insp./exp.) | Double | Waveforms | Evita Infinity V800
582 | Resistance | mbar/L/s | For Evita 4, Evita 2 dura, Evita XL: Values below 100 will be transmitted with one decimal place, otherwise without a decimal place. For Savina: Available since device SW version 02.00. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Zeus IE, Evita V300. | Double | VitalSigns | Evita Infinity V800
584 | Ventilation Mode | unitless | The ventilation mode is reported in variable 584 and 2756. Variable 2756 contains mode ID numeric value, while variable 584 contains string representation which is the complete ventilation mode information concatenated from basic ventilation mode together with available options such as "Mode VC-CMV /AF". See variable 2756 for full list of supported ventilation mode. | String | DeviceInfo | Evita Infinity V800
635 | Inspired O2 (FiO2) | % | For Savina: This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. Available since device SW version 03.00. Limits are transmitted in LPO-Mode (Option) only. For Fabius GS, Fabius GS premium, Fabius Tiro, Fabius plus, Fabius plus XL, Fabius OS, Fabius MRI: This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Evita V300, Fabius Family. This variable is associated with UpperAlarmLimit attribute, which is available for Perseus A500, Atlan, Primus, Primus IE, Apollo, Zeus IE, Fabius Family. This variable is associated with LowerAlarmLimit attribute, which is available for Perseus A500, Atlan, Primus, Primus IE, Apollo, Zeus IE, Fabius Family. Inspiratory oxygen fraction, FiO2. For Zeus, Zeus IE: This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. For Primus, Primus IE, Apollo, Pallas: This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. | Double | VitalSigns | Evita Infinity V800
754 | Device Name | unitless | This variable contains the device name defined by the manufacture. See variable 3591 for device model identification number. | String | DeviceInfo | Common | Device Information
775 | Inspiratory Component of I:E Ratio | unitless | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: Obsolete variable that may be removed in future release. Partial information of I:E ratio. For Evita Infinity V500, Savina: Partial information. Please use variable 1190 instead. Obsolete variable that may be removed in future release. | Double | VitalSigns | Evita Infinity V800
776 | Expiratory Component of I:E Ratio | unitless | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: Obsolete variable that may be removed in future release. Partial information of I:E ratio. For Evita Infinity V500, Savina: Partial information. Please use variable 1190 instead. Obsolete variable that may be removed in future release. | Double | VitalSigns | Evita Infinity V800
1189 | Positive End Expiratory Pressure (PEEP) | mbar | For Fabius GS, Fabius GS premium, Fabius Tiro, Fabius plus, Fabius plus XL, Fabius OS, Fabius MRI: 1mbar = 1cmH2O. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300, Fabius Family. | Double | VitalSigns | Evita Infinity V800
1190 | I:E Ratio | unitless | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Evita V300. Format: 1:x if E > I, x:1 if I > E, or 1:1 if I = E. The useless leading or ending “0” or “.” is not transmitted, and the maximum digits after the floating point is 2. For Evita Infinity V500, Babylog VN500: Format: 1:x if E > I, x:1 if I > E, or 1:1 if I = E. The useless leading or ending “0” or “.” is not transmitted, and the maximum digits after the floating point is 2. | String | VitalSigns | Evita Infinity V800
1307 | Minute Volume | L/min | For Evita 4, Evita 2 dura, Evita XL: This variable is associated with UpperAlarmLimit attribute, which is also sent in separate variable 1400. This variable is associated with LowerAlarmLimit attribute, which is also sent in separate variable 1399. For device firmware version 04.00 and higher: Lower alarm limit is not transmitted, if alarm is switched off. Values may be transmitted with one or two decimal place, depended on firmware version. Note that the value sent by the device is less precise than what is displayed on device screen. For example, a displayed value of 5.12 is sent by the device and thus by DDI as 5.1. For Evita 2, Evita, Fabius GS, Fabius GS premium, Fabius Tiro, Fabius plus, Fabius plus XL, Fabius OS, Fabius MRI, Evita Infinity V500, Babylog VN500, Carina: This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. This variable is available only with a precision to the tenth. For example: for a minute volume value of 7.43, the device displays 7.43 L/min but the DDI will output it as 7.4 L/min. For Savina: This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300, Fabius Family. This variable is associated with UpperAlarmLimit attribute, which is available for Perseus A500, Atlan, Savina 300, Primus, Primus IE, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Fabius Family. This variable is associated with LowerAlarmLimit attribute, which is available for Perseus A500, Atlan, Savina 300, Primus, Primus IE, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Fabius Family. This variable is available only with a precision to the tenth. For example: for a minute volume value of 7.43, the device displays 7.43 L/min but the DDI will output it as 7.4 L/min. For Zeus, Zeus IE, Primus, Primus IE, Apollo, Pallas: This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. Respiratory Minute Volume This variable is available only with a precision to the tenth. For example: for a minute volume value of 7.43, the device displays 7.43 L/min but the DDI will output it as 7.4 L/min. For Evita Infinity V500 and Babylog VN500: This variable is only output if the device version is less than 2.30, otherwise variable 2326 is exported. | Double | VitalSigns | Evita Infinity V800
1314 | Spontaneous Respiration Rate | 1/min | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Zeus IE, Oxylog 3000 plus, Evita V300. This variable is associated with UpperAlarmLimit attribute, which is available for Oxylog 3000 plus. | Double | VitalSigns | Evita Infinity V800
1318 | Static Compliance | mL/mbar | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500 Evita Infinity V600, Evita Infinity V800. For Evita Infinity V500: Cstat (Low Flow) Cstat. | Double | VitalSigns | Evita Infinity V800
1320 | Mandatory Respiration Rate Setting | 1/min | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300, Fabius Family. Respiratory rate, RR. For Evita Infinity V500, Babylog VN500: Respiratory Rate RR. | Double | Settings | Evita Infinity V800
1324 | Sigh Interval Setting | s | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. Time interval between sighs, Interval sigh. | Double | Settings | Evita Infinity V800
1326 | Sigh Number Setting | unitless | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. Number of sigh cycles, Cycles sigh. | Double | Settings | Evita Infinity V800
1328 | Apnea Tidal Volume Setting | - | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Oxylog 3000 plus, Carina, Evita V300. Tidal volume for apnea ventilation, Vtapn. Possible units: L, mL. Unit "L" is available for Evita Infinity V500, Babylog VN500, Savina 300, Oxylog 3000 plus, Carina, Evita V300. Unit "mL" is available for Evita Infinity V500, Babylog VN500, Evita V300. For Evita Infinity V500, Babylog VN500: Possible units: L, mL. VT Apnea VTapn. | Double (Multiset) | Settings | Evita Infinity V800
1329 | Apnea Respiration Rate Setting | 1/min | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Oxylog 3000 plus, Carina, Evita V300. Respiratory rate for apnea ventilation, Rrapn. | Double | Settings | Evita Infinity V800
1332 | Pressure Support Level Above PEEP Setting | mbar | For Evita 4, Evita 2 dura, Evita XL: This variable is calculated from variables 1392 and 2104. Available only with Evita XL. For Medibus.X compatible device, Perseus A500, Atlan, Evita V300, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300, Fabius Family. Pressure amplitude above PEEP in Pressure Support, DPsupp. For Evita Infinity V500, Babylog VN500: ΔPsupp. For Carina: Assisted spontaneous breathing. | Double | Settings | Evita Infinity V800
1381 | Inspiratory Flow Setting | L/min | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300, Fabius Family. | Double | Settings | Evita Infinity V800
1383 | Inspiratory Flow Trigger Setting | L/min | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Evita V300, Fabius Family. Flow trigger threshold, Flow trigger. For Primus, Primus IE, Apollo, Pallas: Flow trigger. | Double | Settings | Evita Infinity V800
1386 | Intermittent PEEP Setting | mbar | For Fabius GS, Fabius GS premium, Fabius Tiro, Fabius plus, Fabius plus XL, Fabius OS, Fabius MRI: 1mbar = 1cmH2O. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Evita V300. Additional intermittent PEEP for sighs, PEEPsigh. For Evita Infinity V500, Babylog VN500: Interm. PEEP ΔintPEEP. For Zeus, Zeus IE: PEEP for all vent mode different from Man/Spont. | Double | Settings | Evita Infinity V800
1387 | APRV Low Pressure Setting | mbar | For Evita: BIPAP low pressure For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Evita V300. Lower pressure level during APRV, Plow. For Evita Infinity V500, Babylog VN500: APRV Low Pressure Plow. | Double | Settings | Evita Infinity V800
1388 | APRV Low Time Setting | s | For Evita: BIPAP low time For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Evita V300. Time of lower pressure level during APRV, Tlow. For Evita Infinity V500, Babylog VN500: APRV Low Time Tlow. | Double | Settings | Evita Infinity V800
1389 | APRV High Pressure Setting | mbar | For Evita: BIPAP high pressure For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Evita V300. Upper pressure level during APRV, Phigh. For Evita Infinity V500, Babylog VN500: APRV High Pressure Phigh. | Double | Settings | Evita Infinity V800
1390 | APRV High Time Setting | s | For Evita: BIPAP high time For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Evita V300. Time of upper pressure level during APRV, Thigh. For Evita Infinity V500, Babylog VN500: APRV High Time Thigh. | Double | Settings | Evita Infinity V800
1391 | Apnea Time Setting | s | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Oxylog 3000 plus, Carina, Evita V300. For Evita Infinity V500, Babylog VN500: Apnea Time Tapn. For Carina: Apnea alarm time. | Double | Settings | Evita Infinity V800
1393 | Maximum Inspired Pressure Setting | mbar | For Fabius GS, Fabius GS premium, Fabius Tiro, Fabius plus, Fabius plus XL, Fabius OS, Fabius MRI: 1mbar = 1cmH2O. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300, Fabius Family. Pressure limitation, Pmax. For Evita Infinity V500, Babylog VN500: Max Pressure Pmax. For Carina: Pressure limitation. For Primus, Primus IE, Apollo, Pallas: Max. insp. Airway pressure. | Double | Settings | Evita Infinity V800
1415 | Mean Airway Pressure | mbar | For Fabius GS, Fabius GS premium, Fabius Tiro, Fabius plus, Fabius plus XL, Fabius OS, Fabius MRI: This variable is associated with LowerAlarmLimit attribute. 1mbar = 1cmH2O. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300, Fabius Family. For Zeus, Zeus IE, Primus, Primus IE, Apollo, Pallas: Mean Breathing Pressure. | Double | VitalSigns | Evita Infinity V800
1422 | Rapid Shallow Breathing Index | - | For Evita 4, Evita 2 dura, Evita XL: Firmware version 04.00 or higher For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Atlan, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Evita V300. Possible units: 1/L * min, 1/mL*min. Unit "1/L * min" is available for Evita Infinity V500, Babylog VN500, Savina 300, Evita V300. Unit "1/mL*min" is available for Evita Infinity V500, Babylog VN500, Evita V300 (added with SW 2.30 for Babylog VN500, added with SW 2.30 for Evita Infinity V500). | Double (Multiset) | VitalSigns | Evita Infinity V800
1426 | Intrinsic PEEP | mbar | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. | Double | VitalSigns | Evita Infinity V800
1569 | Minimum Airway Pressure | mbar | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. | Double | VitalSigns | Evita Infinity V800
1570 | P0.1 Occlusion Pressure | mbar | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. Occlusion pressure in P0.1 maneuver. For Evita Infinity V500, Babylog VN500: Only available for Evita Infinity V500 | Double | VitalSigns | Evita Infinity V800
1571 | Lung Volume Trapped (TrpVol) | mL | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. Volume trapped in the lung by intrinsic PEEP, Vtrap. For Evita Infinity V500: Trapped Volume Vtrap. | Double | VitalSigns | Evita Infinity V800
1575 | Respired Volume Wave | mL | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Evita V300. Respiratory volume since start of inspiration, Volume. For Evita Infinity V500, Babylog VN500: Volume since inspiration began. | Double | Waveforms | Evita Infinity V800
1637 | Ventilation Pressure Variability Setting | % | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800. For Evita Infinity V500, Babylog VN500: Supported through Medibus.X protocol. | Double | Settings | Evita Infinity V800
1638 | APRV Low Time Limitation | s | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. Maximum time of low pressure level during APRV/AutoRelease, Tlow max. For Evita Infinity V500, Babylog VN500: Tlow max. | Double | Settings | Evita Infinity V800
1640 | Spontaneous Mean Tidal Volume | mL | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500 Evita Infinity V600, Evita Infinity V800. Spontaneous mean tidal volume, VTspon mean. For Evita Infinity V500, Babylog VN500: VTspon mean (Not available on Babylog VN500). Only available for Evita Infinity V500 | Double | VitalSigns | Evita Infinity V800
1641 | Spontaneous Inspired Mean Tidal Volume | mL | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500 Evita Infinity V600, Evita Infinity V800. For Evita Infinity V500, Babylog VN500: VTispon mean (Not available on Babylog VN500). Only available for Evita Infinity V500 | Double | VitalSigns | Evita Infinity V800
1642 | Spontaneous Expired Mean Tidal Volume | mL | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800. For Evita Infinity V500, Babylog VN500: VTespon mean (Not available on Babylog VN500). Only available for Evita Infinity V500 | Double | VitalSigns | Evita Infinity V800
1650 | SmartCare FiO2 Limitation | % | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. FiO2 in SmartCare - upper limit, SC-FiO2high For Evita Infinity V500, Babylog VN500: SC-Upper limit inspiratory O2 FiO2 max. | Double | Settings | Evita Infinity V800
1651 | SmartCare PEEP Limitation | mbar | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. PEEP in SmartCare - upper limit, SC-PEEPmax. For Evita Infinity V500, Babylog VN500: SC-Upper limit PEEP, PEEP max. | Double | Settings | Evita Infinity V800
1652 | SmartCare Spontaneous Respiration Rate Lower Limitation | 1/min | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. Spontaneous respiratory rate in SmartCare - lower limit, SC-RRspon low. For Evita Infinity V500, Babylog VN500: SC-Lower limit respiratory rate in the guideline, RRspon low. | Double | Settings | Evita Infinity V800
1653 | SmartCare Spontaneous Respiration Rate Upper Limitation | 1/min | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. Spontaneous respiratory rate in SmartCare - upper limit, SC-RRspon high. For Evita Infinity V500, Babylog VN500: SC-Upper limit respiratory rate in the guideline, RRspon high. | Double | Settings | Evita Infinity V800
1654 | SmartCare Tidal Volume Limitation | mL/kg | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. Tidal volume in SmartCare - lower limit, SC-VTlow. For Evita Infinity V500, Babylog VN500: SC-Lower limit tidal volume, VT low. | Double | Settings | Evita Infinity V800
1655 | SmartCare EtCO2 Limitation | mmHg | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. End-tidal CO2 in SmartCare - upper limit, SC-etCO2high. For Evita Infinity V500, Babylog VN500: SC-Upper limit endtidal CO2 in the guideline etCO2 high. | Double | Settings | Evita Infinity V800
1656 | SmartCare Patient Height | cm | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. Patient height in SmartCare, SC-Ht. For Evita Infinity V500: SC-Body height. | Double | Settings | Evita Infinity V800
1761 | Airway Resistance | mbar/L/s | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. | Double | VitalSigns | Evita Infinity V800
2044 | P01 Repetition Interval Setting | min | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. P0.1 repetition interval. | Double | Settings | Evita Infinity V800
2092 | Inspired Minute Volume | L/min | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Carina, Evita V300. This variable is associated with UpperAlarmLimit attribute, which is available for Carina. This variable is associated with LowerAlarmLimit attribute, which is available for Carina. | Double | VitalSigns | Evita Infinity V800
2102 | Inspiratory Time Setting | s | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300, Fabius Family. | Double | Settings | Evita Infinity V800
2104 | Positive End Expiratory Pressure (PEEP) Setting | mbar | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300, Fabius Family. For Evita Infinity V500, Babylog VN500: PEEP / CPAP PEEP. For Zeus, Zeus IE: PEEP (CPAP). PEEP for mode Man/Spont. | Double | Settings | Evita Infinity V800
2110 | Inspired O2 (FiO2) Setting | % | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Zeus IE, Oxylog 3000 plus, Carina, Evita V300. For Primus, Primus IE, Apollo, Pallas: Not in device Apollo and Pallas | Double | Settings | Evita Infinity V800
2217 | Manual Breath Ventilation Pressure Setting | mbar | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300(added with SW 2.30 for Babylog VN500, added with SW 2.30 for Evita Infinity V500). Manual inspiratory pressure for Neonates during NIV - Non-Invasive Ventilation, PmanInsp. | Double | Settings | Evita Infinity V800
2245 | Manual Breath Maximum Time Limit | s | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300(added with SW 2.30 for Babylog VN500, added with SW 2.30 for Evita Infinity V500). Maximum time of pressure for manual inspiration pressure for Neonates during NIV - Non-Invasive Ventilation, TmanInsp. | Double | Settings | Evita Infinity V800
2250 | CO2 Tidal Production | mL | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. CO2 production volume per breath, VTCO2. For Evita Infinity V500, Babylog VN500: Supported through Medibus.X protocol. | Double | VitalSigns | Evita Infinity V800
2251 | CO2 Minute Production | mL/min | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Evita V300. CO2 production, V'CO2. | Double | VitalSigns | Evita Infinity V800
2255 | Device Date/Time | unitless | Format: YYYYMMDDhhmmss, where ss is always present but might be 00. | String | DeviceInfo | Evita Infinity V800 | Device Information
2324 | Inspired Tidal Volume | mL | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Carina, Evita V300. This variable is associated with UpperAlarmLimit attribute, which is available for Evita Infinity V500, Zeus IE, Evita V300. This variable is associated with LowerAlarmLimit attribute, which is available for Evita Infinity V500, Evita V300. For Primus, Primus IE, Apollo, Pallas: Insp. Tidal Volume. | Double | VitalSigns | Evita Infinity V800
2325 | Expired Tidal Volume | mL | For Evita XL: Neo expired tidal volume. Available only with firmware version 7.01 or higher. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Evita V300. | Double | VitalSigns | Evita Infinity V800
2326 | Expired Minute Volume | L/min | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. This variable is associated with UpperAlarmLimit attribute, which is available for Evita Infinity V500, Babylog VN500, Evita V300. This variable is associated with LowerAlarmLimit attribute, which is available for Evita Infinity V500, Babylog VN500, Evita V300. For Evita Infinity V500, Babylog VN500: This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. | Double | VitalSigns | Evita Infinity V800
2355 | End Tidal CO2 Percentage | % | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Evita V300, Fabius Family. This variable is associated with UpperAlarmLimit attribute, which is available for Perseus A500, Atlan, Evita Infinity V500, Babylog VN500, Savina 300, Primus, Primus IE, Apollo, Zeus IE, Oxylog 3000 plus, Evita V300. This variable is associated with LowerAlarmLimit attribute, which is available for Perseus A500, Atlan, Evita Infinity V500, Babylog VN500, Savina 300, Primus, Primus IE, Apollo, Zeus IE, Oxylog 3000 plus, Evita V300. end tidal CO2 etCO2 | Double | VitalSigns | Evita Infinity V800
2377 | Fresh Gas O2 Flow Rate | L/min | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Evita V300. O2 therapy flow, Flow. For Evita Infinity V500, Babylog VN500: Flow O2 Therapy Flow. | Double | DeviceInfo | Evita Infinity V800
2382 | Ventilator Operation Mode | unitless | This variable contains ventilation operation mode in string format, which is also send as a numerical value in variable 2758. See detailed description of variable 2758 for the list of supported values. For Evita 4, Evita 2 dura, Evita XL: Firmware version 04.00 or higher For Evita 2: Variable label is not correct. Please use variable 3853 instead. Obsolete variable that may be removed in future release. | String | DeviceInfo | Evita Infinity V800
2416 | Spontaneous Inspired Tidal Volume | mL | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. For Zeus, Zeus IE: This variable is associated with UpperAlarmLimit attribute. Insp. Tidal volume. | Double | VitalSigns | Evita Infinity V800
2421 | Airway Dead Space | mL | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. Serial dead space volume, Vds. For Evita Infinity V500, Babylog VN500: Serial dead space Vds. | Double | VitalSigns | Evita Infinity V800
2423 | Dynamic Compliance Over Last 20% Breath | unitless | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. Ratio of compliance during the last 20 % of inspiration over dynamic compliance, C20/Cdyn. For Evita Infinity V500, Babylog VN500: C20/Cdyn. | Double | VitalSigns | Evita Infinity V800
2432 | Expired Spontaneous Tidal Volume | mL | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. | Double | VitalSigns | Evita Infinity V800
2433 | Mechanical Expired Minute Volume | L/min | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Primus, Primus Infinity Empowered, Apollo, Evita V300. Mandatory expiratory minute volume, MVemand | Double | VitalSigns | Evita Infinity V800
2434 | Spontaneous Expired Minute Volume | L/min | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Primus, Primus Infinity Empowered, Apollo, Evita V300. | Double | VitalSigns | Evita Infinity V800
2467 | Patient Type ID | unitless | Consult detailed help for more information on this variable | UByte | PatientInfo | Evita Infinity V800
2542 | Airway Pressure Wave | mbar | For Fabius GS, Fabius GS premium, Fabius Tiro, Fabius plus, Fabius plus XL, Fabius OS, Fabius MRI: 1mbar = 1cmH2O. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300, Fabius Family. For Evita Infinity V500, Babylog VN500: Airway pressure Paw. | Double | Waveforms | Evita Infinity V800
2564 | Peak Airway Pressure | mbar | For Evita 4, Evita 2 dura, Evita XL: This variable is associated with UpperAlarmLimit attribute, which is also sent in separate variable 7040. For Evita 2, Evita, Savina: This variable is associated with UpperAlarmLimit attribute. For Fabius GS, Fabius GS premium, Fabius Tiro, Fabius plus, Fabius plus XL, Fabius OS, Fabius MRI: 1mbar = 1cmH2O. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300, Fabius Family. This variable is associated with UpperAlarmLimit attribute, which is available for Perseus A500, Atlan, Evita Infinity V500, Babylog VN500, Savina 300, Zeus IE, Oxylog 3000 plus, Carina, Evita V300, Fabius Family. This variable is associated with LowerAlarmLimit attribute, which is available for Zeus IE, Fabius Family. Peak inspiratory pressure, PIP. For Evita Infinity V500, Babylog VN500, Carina: This variable is associated with UpperAlarmLimit attribute. For Zeus, Zeus IE: This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. Peak Breathing Press. For Primus, Primus IE, Apollo, Pallas: Peak Breathing Pressure. | Double | VitalSigns | Evita Infinity V800
2565 | Plateau Airway Pressure | mbar | For Fabius GS, Fabius GS premium, Fabius Tiro, Fabius plus, Fabius plus XL, Fabius OS, Fabius MRI: 1mbar = 1cmH2O. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Evita V300, Fabius Family. For Evita Infinity V500, Babylog VN500: Only available for Evita Infinity V500 | Double | VitalSigns | Evita Infinity V800
2570 | Spontaneous Tidal Volume | - | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Zeus IE, Evita V300. Possible units: L, mL. For Evita Infinity V500, Babylog VN500: Possible units: L, mL. | Double (Multiset) | VitalSigns | Evita Infinity V800
2583 | I:E Ratio Setting | unitless | String format: 1:x if E > I, x:1 if I > E, or 1:1 if I = E. The useless leading or ending “0” or “.” is not transmitted, and the maximum digits after the floating point is 2. For Carina, Primus, Primus IE, Apollo, Pallas, Perseus A500: This variable is calculated by the DDI from respiration rate and inspiratory time settings. For Medibus.X compatible device: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300, Fabius Family. For Primus, Primus IE, Apollo, Perseus A500, Carina and Oxylog 3000 plus, it is calculated by the DDI from respiration rate and inspiratory time settings. For Atlan, Evita Infinity V500, Babylog VN500, Savina 300, Zeus IE, Evita V300, and Fabius Family, it is calculated by the DDI from inspiratory and expiratory components. For Evita, Evita 2, Evita 4, Evita 2 dura, and Evita XL, Fabius GS, Fabius GS premium, Fabius Tiro, Fabius plus, Fabius plus XL, Fabius OS, Fabius MRI, Evita Infinity V500, Babylog VN500, Zeus IE, Evita V300: This variable is calculated by the DDI from inspiratory and expiratory components of I:E ratio sent by the devices. | String | Settings | Evita Infinity V800
2596 | Alarm Severity Level | unitless | This variable is in multiset format. This variable contains alarm priority level, variable 6164 contains alarm codes, and variable 3878 contains alarm messages. They need to be interpreted together. One alarm code may have one or more priority values and one or more alarm messages associated with it. The possible alarms are depended on device firmware version. Please see device user manuals for details. See variables 177 and 588 and 4131 for string format. | UByte (Multiset) | Alarms | Evita Infinity V800 | Alarm And Text Message
2754 | Ventilator General Status | unitless | Ventilator Standby mode status. | UByte | DeviceInfo | Evita Infinity V800
Consult detailed help for more information on this variable
2756 | Ventilation Mode ID | unitless | This variable contains ID of basic ventilation mode ID without options. See also variable 584. | UByte | DeviceInfo | Evita Infinity V800
Consult detailed help for more information on this variable
2758 | Ventilator Operation Mode ID | unitless | For Evita 4, Evita 2 dura, Evita XL: Firmware version 04.00 or higher For Evita 2: Variable label is not correct. Please use variable 2467 instead. Obsolete variable that may be removed in future release. | UByte | DeviceInfo | Evita Infinity V800
Consult detailed help for more information on this variable
2771 | Tidal Volume Setting | L | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300, Fabius Family. Inspiratory tidal volume, Vti. For Evita Infinity V500, Babylog VN500: See also variable 1321 and 5044. For Zeus, Zeus IE: This variable contains same information as variable 1321, except that the value of different units (L or mL) are sent in separate variables. See also variable 5044. | Double | Settings | Evita Infinity V800
2784 | ASB Ramp Time Setting | s | For Evita 2: Firmware version 2.00 or higher. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300. Pressure rise time, Slope. For Carina: ASB ramp. For Primus, Primus IE, Apollo, Pallas: Slope time (Ramp). | Double | Settings | Evita Infinity V800
2952 | PS Activation State | unitless | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300, Fabius Family. A value of 1 indicates that Pressure Support (PS) is added to the current ventilation mode. A value of 0 indicates that Pressure Support (PS) is not added to the current ventilation mode. Text message displayed on device screen: "/PS" | Boolean | DeviceInfo | Evita Infinity V800
2953 | VS Activation State | unitless | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. A value of 1 indicates that Volume support (VS) is added to the current ventilation mode. A value of 0 indicates that Volume support (VS) is not added to the current ventilation mode. Text messages displayed on device screen: "/VS" | Boolean | DeviceInfo | Evita Infinity V800
2954 | APRV High Pressure (Phigh) | mbar | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. Upper pressure level during APRV, Phigh. | Double | VitalSigns | Evita Infinity V800
2959 | APRV Low Pressure (Plow) | mbar | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. Lower pressure level during APRV, Plow. | Double | VitalSigns | Evita Infinity V800
2987 | Inspired Pressure Setting | mbar | For Evita 4, Evita 2 dura, Evita XL: Available only with EvitaXL firmware version 7.00 or higher. For Fabius GS, Fabius GS premium, Fabius Tiro, Fabius plus, Fabius plus XL, Fabius OS, Fabius MRI: 1mbar = 1cmH2O. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300, Fabius Family. | Double | Settings | Evita Infinity V800
3326 | Device Average Gas Flow | L/min | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. Average device flow, FlowDev. | Double | OtherMeasuredData | Evita Infinity V800
3503 | Auto Flow Active Status | unitless | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300. A value of 1 indicates that AutoFlow (AF) is added to the current ventilation mode. A value of 0 indicates that AutoFlow (AF) is not added to the current ventilation mode. Text message displayed on device screen: "/AF" | Boolean | DeviceInfo | Evita Infinity V800
3555 | ATC Activation Status | unitless | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Evita V300. A value of 1 indicates that automatic tube compensation (ATC) is added to the current ventilation mode. A value of 0 indicates that automatic tube compensation (ATC) is not added to the current ventilation mode. Text message displayed on device screen: "/ATC" For Savina 300 (MEDIBUS.X): added with SW 5.00 | Boolean | DeviceInfo | Evita Infinity V800
3591 | Device Number | unitless | Consult detailed help for more information on this variable | String | DeviceInfo | Common | Device Information
3593 | Device Communications Status | unitless | Value "stop" sent when a request to stop the communication is received from the device | String | DeviceInfo | Common | Device Information
3853 | Patient Type | unitless | This variable contains patient type in string format, which is also send as a numerical value in variable 2467. See detailed description of variable 2467 for the list of supported values. | String | PatientInfo | Evita Infinity V800
3878 | Generic Alarm Message | unitless | This variable is in multiset format. This variable contains alarm messages, variable 2596 contains alarm priority level, variable 6164 contains alarm codes. They need to be interpreted together. One alarm code may have one or more priority values and one or more alarm messages associated with it. The possible alarms are dependent on device firmware version. Please see device user manuals for details. See variables 177 and 588 and 4131 for string format. The HL7 output can contain an '&' character to indicate a LOW alarm. In order to avoid issues, the '&' character will be replaced by a '?' character in the DDI output. | String (Multiset) | Alarms | Evita Infinity V800 | Alarm And Text Message
3966 | Inspired Mandatory Tidal Volume | mL | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. | Double | VitalSigns | Evita Infinity V800
3968 | Spontaneous Fraction of Minute Volume | % | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. Spontaneous breathing portion of minute volume, % Mvspon. For Evita Infinity V500, Babylog VN500: Spontaneous Fraction Min. Vol. %MVspon (MVspon : MVtotal). | Double | VitalSigns | Evita Infinity V800
3971 | Negative Inspiratory Force | mbar | For Evita 4, Evita 2 dura, Evita XL: The value is not transmitted, if it is displayed on Evita 4 / Evita 2 dura as "[less than] –45 mbar". Firmware version 04.00 or higher For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. | Double | VitalSigns | Evita Infinity V800
3998 | Auto Release Activation Status | unitless | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. A value of 1 indicates that AutoRelease is added to the current ventilation mode. A value of 0 indicates that AutoRelease is not added to the current ventilation mode. Text messages displayed on device screen: "/AutoRelease" | Boolean | DeviceInfo | Evita Infinity V800
4003 | Generic Respiration Rate | 1/min | For Evita 4, Evita 2 dura, Evita XL: This variable is associated with UpperAlarmLimit attribute, which is also sent in separate variable 1339. Total Respiratory Rate (RRtotal). For Evita 2, Evita, Savina, Carina: This variable is associated with UpperAlarmLimit attribute. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300, Fabius Family. This variable is associated with UpperAlarmLimit attribute, which is available for Evita Infinity V500, Babylog VN500, Savina 300, Carina, Evita V300. Respiratory rate, RR. For Evita Infinity V500, Babylog VN500: This variable is associated with UpperAlarmLimit attribute. Respiratory Rate (RRtotal) RR. | Double | VitalSigns | Evita Infinity V800
4010 | Minute Volume Alarm Delay | s | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. Delay for minute volume high alarm, MVhigh delay. This variable is associated with a ChannelID identifier. | Double | Settings | Evita Infinity V800
4011 | Flow Assist Level Setting | 0.1 mbar/L/s | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. Flow assist during PPS ventilation, Flow Assist. For Evita Infinity V500, Babylog VN500: Flow Assist. | Double | Settings | Evita Infinity V800
4012 | Volume Assist Level Setting | 0.1 mbar/L | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. Volume assist during PPS ventilation, Vol. Assist. For Evita Infinity V500, Babylog VN500: Vol. Assist. | Double | Settings | Evita Infinity V800
4015 | Disconnection Alarm Delay Time | s | For Evita 4, Evita 2 dura, Evita XL: Firmware version 4.00 or higher. For Savina: Tdeconnect. Available since Device SW version 02.00. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Carina, Evita V300. Disconnection alarm time, Tdiscon. For Evita Infinity V500, Babylog VN500: Tdisconnect. For Carina: Disconnection detection timeout. | Double | Settings | Evita Infinity V800
4044 | Medibus Revision | unitless |  | String | DeviceInfo | Common | Device Information
4045 | Device Revision | unitless | Format: mm.nn | String | DeviceInfo | Common | Device Information
4150 | VG Activation Status | unitless | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Carina, Evita V300. A value of 1 indicates that Volume Guarantee (VG) is added to the current ventilation mode. A value of 0 indicates that Volume Guarantee (VG) is not added to the current ventilation mode. Text message displayed on device screen: "/VG" | Boolean | DeviceInfo | Evita Infinity V800
4732 | Inspiratory Flow Waveform | L/min | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. | Double | Waveforms | Evita Infinity V800
4889 | Spontaneous Inspiration Time | s | For Evita 4, Evita 2 dura, Evita XL: Available only with EvitaXL firmware version 7.02 or higher. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. | Double | VitalSigns | Evita Infinity V800
4891 | Lung Time Constant | s | For Evita 4, Evita 2 dura, Evita XL: Available only with EvitaXL firmware version 7.00 or higher. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. | Double | VitalSigns | Evita Infinity V800
4892 | Airway Correlation Coefficient | unitless | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. Correlation factor, r2. For Evita Infinity V500, Babylog VN500: r². | Double | VitalSigns | Evita Infinity V800
4966 | Sigh Activation Status | unitless | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. A value of 1 indicates that the sigh function has been added to the current ventilation mode. A value of 0 indicates that the sigh function has not been added to the current ventilation mode. Text message displayed on device screen: "/Sigh" | Boolean | DeviceInfo | Evita Infinity V800
5044 | Inspired Tidal Volume Setting | mL | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Zeus IE, Evita V300. Inspiratory tidal volume, Vti. For Evita Infinity V500, Babylog VN500: See also variable 1321 and 2771. For Zeus, Zeus IE: This variable contains same information as variable 1321, except that the value of different units (L or mL) are sent in separate variables. See also variable 2771. | Double | Settings | Evita Infinity V800
5087 | Dead Space to Tidal Volume Ratio | % | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. Ratio of serial dead space volume to expiratory tidal volume, Vds/VTe. For Evita Infinity V500, Babylog VN500: Dead space ventilation Vds/VTe. | Double | VitalSigns | Evita Infinity V800
5340 | Expired Mandatory Tidal Volume | mL | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. | Double | VitalSigns | Evita Infinity V800
5374 | Expiratory Time Constant (RCexp) | s | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. | Double | VitalSigns | Evita Infinity V800
5728 | EtCO2 Phase III Increment | - | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. Possible units: mmHg/L, kPa/L, Vol%/L. Increase of measured CO2 value in phase III of the capnogram, CO2 slope. | Double (Multiset) | VitalSigns | Evita Infinity V800
5815 | Spontaneous I:E Ratio | unitless | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. Format: 1:x if E > I, x:1 if I > E, or 1:1 if I = E. The useless leading or ending “0” or “.” is not transmitted, and the maximum digits after the floating point is 2. For Evita Infinity V500, Babylog VN500: Format: 1:x if E > I, x:1 if I > E, or 1:1 if I = E. The useless leading or ending “0” or “.” is not transmitted, and the maximum digits after the floating point is 2. | String | VitalSigns | Evita Infinity V800
6164 | General Alarm Code | unitless | This variable is in multiset format. This variable contains alarm codes, variable 3878, alarm messages, variable 2596 contains alarm priority level. They need to be interpreted together. One alarm code may have one or more priority values and one or more alarm messages associated with it. The possible alarms are depended on device firmware version. Please see device user manuals for details. See variables 177 and 588 and 4131 for string format. | UShort (Multiset) | Alarms | Evita Infinity V800 | Alarm And Text Message
Consult detailed help for more information on this variable
6389 | Endotracheal Tube Size | mm | For Evita 4, Evita 2 dura, Evita XL: Possible units: "0.1mm" in firmware versions 7.00 and 7.01, "mm" otherwise. Available only with EvitaXL firmware version 7.00 or higher. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Evita V300. Inner diameter of the endotracheal/tracheostoma tube. For Evita Infinity V500, Babylog VN500: ATC Tube Diameter. | Double | DeviceInfo | Evita Infinity V800
6551 | End Tidal CO2 Pressure | - | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Evita V300, Fabius Family. This variable is associated with UpperAlarmLimit attribute, which is available for Perseus A500, Atlan, Evita Infinity V500, Babylog VN500, Savina 300, Primus, Primus IE, Apollo, Zeus IE, Oxylog 3000 plus, Evita V300. This variable is associated with LowerAlarmLimit attribute, which is available for Perseus A500, Atlan, Evita Infinity V500, Babylog VN500, Savina 300, Primus, Primus IE, Apollo, Zeus IE, Oxylog 3000 plus, Evita V300. Possible units: mmHg, kPa. end tidal CO2 etCO2 | Double | VitalSigns | Evita Infinity V800
6701 | Expiratory Time Setting | s | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Zeus IE, Evita V300. | Double | Settings | Evita Infinity V800
6702 | Inspiratory Component of I:E Ratio Setting | unitless | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Zeus IE, Evita V300, Fabius Family, Atlan For Evita Infinity V500: Partial information. Please use variable 2583 instead. Obsolete variable that may be removed in future release. | Double | Settings | Evita Infinity V800
6703 | Expiratory Component of I:E Ratio Setting | unitless | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800,Savina 300, Zeus IE, Evita V300, Fabius Family, Atlan For Evita Infinity V500: Partial information. Please use variable 2583 instead. Obsolete variable that may be removed in future release. | Double | Settings | Evita Infinity V800
6735 | Lung Elastance | mbar/L | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. | Double | VitalSigns | Evita Infinity V800
7076 | Generic Ventilation Mode | unitless | Consult detailed help for more information on this variable | String | Settings | Evita Infinity V800
7077 | Ventilation Mode Primary | unitless | Consult detailed help for more information on this variable | String | Settings | Evita Infinity V800
7078 | Ventilation Mode Options | unitless | Consult detailed help for more information on this variable | String | Settings | Evita Infinity V800
7318 | Variable PS Activation Status | unitless | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800. A value of 1 indicates that Variable PS (VPS) has been added to the current ventilation mode. A value of 0 indicates that Variable PS (VPS) has not been added to the current ventilation mode. Text message displayed on device screen: "VariablePS" | Boolean | Settings | Evita Infinity V800
7364 | Patient Operation Mode | unitless | Possible values: 1 for Invasive, 2 for Noninvasive | UByte | PatientInfo | Evita Infinity V800
7637 | Artificial Airway Compensation Status | % | For Evita 4, Evita 2 dura, Evita XL: Tube Compensation. Available only with EvitaXL firmware version 7.00 or higher. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. Compensation rate during ATC, % comp. For Evita Infinity V500, Babylog VN500: ATC Compensation Compens. | Double | DeviceInfo | Evita Infinity V800
7640 | Mandatory Respiration Rate | 1/min | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Zeus IE, Evita V300. | Double | VitalSigns | Evita Infinity V800
7642 | Mandatory Tidal Volume | - | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Babylog VN500, Babylog VN600, Babylog VN800, Zeus IE, Evita V300(added with SW 2.30 for Babylog VN500, added with SW 2.30 for Evita Infinity V500). Possible units: L, mL. Unit "L" available for Evita Infinity V500, Babylog VN500, Evita V300 (added with SW 2.30 for Babylog VN500, added with SW 2.30 for Evita Infinity V500). Unit "mL" available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Zeus IE, Evita V300. | Double (Multiset) | VitalSigns | Evita Infinity V800
7872 | CO2 Partial Pressure Waveform | - | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Evita V300. Possible units: mmHg, kPa. Unit "mmHg" is available for Perseus A500, Atlan, Evita Infinity V500, Babylog VN500, Savina 300, Primus, Primus IE, Apollo, Zeus IE, Oxylog 3000 plus, Evita V300. Unit "kPa" is available for Perseus A500, Atlan, Savina 300, Primus, Primus IE, Apollo, Zeus IE, Oxylog 3000 plus. | Double | Waveforms | Evita Infinity V800
7879 | Time Duration at Low Pressure (Tlow) | s | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. Effective expiratory time during APRV/AutoRelease, Tlow. For Evita Infinity V500, Babylog VN500: Tlow (resulting Tlow from expiratory flow cycling). | Double | VitalSigns | Evita Infinity V800
7904 | Spontaneous Tidal Volume | - |  | Double (Multiset) | VitalSigns | Evita Infinity V800
8366 | Generic Message Code | unitless | This variable is in multiset format. This variable, along with variable 8696, contains text message code and alarm text respectively. They need to be interpreted together. | UShort (Multiset) | Messages | Evita Infinity V800 | Alarm And Text Message
8696 | Generic Message Text | unitless | This variable is in multiset format. This variable, along with variable 8366, contains text message code and alarm text respectively. They need to be interpreted together. | String (Multiset) | Messages | Evita Infinity V800 | Alarm And Text Message
9146 | End Inspiratory Pressure | mbar | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. End-inspiratory pressure, EIP. | Double | VitalSigns | Evita Infinity V800
9163 | Tracheal Pressure Waveform | mbar | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Evita V300. For Evita Infinity V500, Babylog VN500: Ptrach. | Double | Waveforms | Evita Infinity V800
9288 | Tidal Volume Indexed By Body Weight | mL/kg | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. Tidal volume per kg body weight, VT/Wt. The leakage is determined by the difference between the delivered flow on the inspiratory side and the measured flow on the expiratory side, which is displayed as the leakage minute volume MVleak and relative as % leak (MVleak to MV). For Evita Infinity V500, Babylog VN500: VT/kg BW. | Double | VitalSigns | Evita Infinity V800
9466 | Mandatory Triggered Respiration Rate | 1/min | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300(added with SW 2.42 for Evita Infinity V500, Evita V300 and Babylog VN500). Respiratory rate of triggered, mandatory breaths, RRtrig For Evita 4, Evita 2 dura, Evita XL: Available only with Evita XL firmware version 7.00 or higher. | Double | VitalSigns | Evita Infinity V800
9475 | Volume of Leakage (MVLeak) | L/min | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Carina, Evita V300. Leakage minute volume, Mvleak. The leakage is determined by the difference between the delivered flow on the inspiratory side and the measured flow on the expiratory side, which is displayed as the leakage minute volume MVleak and relative as % leak (MVleak to MV). For Evita Infinity V500, Babylog VN500: Mvleak. The leakage is determined by the difference between the delivered flow on the inspiratory side and the measured flow on the expiratory side, which is displayed as the leakage minute volume MVleak and relative as % leak (MVleak to MV). For Carina: Leakage minute volume. For Evita XL: Leakage minute volume. Available since device SW version 7.02. | Double | OtherMeasuredData | Evita Infinity V800
9476 | Percent of Leakage | % | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Evita V300. Leakage minute volume in % of inspiratory minute volume, % leak Leakage rel. % leak - Infinity V500 determines the difference between the delivered flow on the inspiratory side and the measured flow on the expiratory side. This difference provides a measure of the amount of leakage and is displayed by Infinity V500 as the leakage minute volume MVleak and relative as % leak (MVleak to MV). For Evita Infinity V500, Babylog VN500: Leakage rel. % leak. | Double | OtherMeasuredData | Evita Infinity V800
9477 | Inspiratory Component of Spontaneous I:E Ratio | unitless | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: Obsolete variable that may be removed in future release. Partial information of I:E ratio. For Evita Infinity V500: Partial information. Please use variable 5815 instead. Obsolete variable that may be removed in future release. | Double | VitalSigns | Evita Infinity V800
9478 | Expiratory Component of Spontaneous I:E Ratio | unitless | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: Obsolete variable that may be removed in future release. Partial information of I:E ratio. For Evita Infinity V500: Partial information. Please use variable 5815 instead. Obsolete variable that may be removed in future release. | Double | VitalSigns | Evita Infinity V800
9479 | SmartCare Target Delta Support Pressure | mbar | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. Dpressure support goal in SmartCare, SC-DPsupp goal. For Evita Infinity V500: SC-Target ΔPsupp ΔPsupp goal. | Double | VitalSigns | Evita Infinity V800
9480 | SmartCare Rated Delta Support Pressure | mbar | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. Rated Dpressure support by internal controller in SmartCare, SC-DPsupp rated. For Evita Infinity V500: SC-Rated ΔPsupp SC-Δpsupp. | Double | VitalSigns | Evita Infinity V800
9481 | SmartCare Duration of Patient Session | - | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. Duration of session (hours and minutes) in SmartCare, SC-duration. Variable is in Multiset format (need combine the first value in hours with the second value in minutes to obtain the exact duration). Please use variable 4346 instead. For Evita Infinity V500: Multiset variable (need combine the first value in hours with the second value in minutes to obtain the exact duration). | Double (Multiset) | VitalSigns | Evita Infinity V800
9482 | SmartCare Spontaneous Respiration Rate | 1/min | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. Spontaneous respiratory rate in SmartCare, SC-Rrspon. For Evita Infinity V500: SC-Rrspon. | Double | VitalSigns | Evita Infinity V800
9483 | SmartCare Tidal Volume | mL | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. Tidal volume in SmartCare, SC-VT. For Evita Infinity V500: SC-VT. | Double | VitalSigns | Evita Infinity V800
9484 | SmartCare EtCO2 | mmHg | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. End-tidal CO2 concentration in SmartCare, SC-etCO2. For Evita Infinity V500: SC-etCO2. | Double | VitalSigns | Evita Infinity V800
9485 | Maximum Inspiratory Time Setting | s | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Carina, Evita V300. Maximal inspiratory time for flow cycled strokes, Timax. For Carina: Maximal inspiration time for CPAP, ASB. | Double | Settings | Evita Infinity V800
9486 | Inspiratory Flow Termination Level Setting | % | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Savina 300, Oxylog 3000 plus, Evita V300. Inspiratory termination criterion based on peak inspiratory flow, Insp. term. For Evita Infinity V500: Inspiratory termination. | Double | Settings | Evita Infinity V800
9487 | Expiratory Flow Termination Setting | % | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. Expiratory termination criterion during APRV/AutoRelease, Exp. term. For Evita Infinity V500, Babylog VN500: Expiratory termination. | Double | Settings | Evita Infinity V800
9488 | SmartCare Body Weight | kg | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. Ideal body weight in SmartCare, SC-Wt. For Evita Infinity V500: SC-Body weight. | Double | Settings | Evita Infinity V800
9489 | SmartCare Start of Night Rest | h | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. Start of night rest in SmartCare, SC-startrest. For Evita Infinity V500: SC-Start of night rest. | Double | Settings | Evita Infinity V800
9490 | SmartCare End of Night Rest | h | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. End of night rest in SmartCare, SC-endrest. For Evita Infinity V500: SC-End of night rest. | Double | Settings | Evita Infinity V800
50021 | Device Alarm Inactivation State | unitless |  | String | DeviceInfo | Evita Infinity V800

## Applix Smart
ID | Label | Unit | Description | Type | Group | Device | Source
1450 | Infusion Pump Type | unitless | 0: Smart or Go 2: Vision or Go+ | UByte | DeviceInfo | Common
1451 | Pump Infusion Status | unitless | Consult detailed help for more information on this variable | UByte | DeviceInfo | Common
1913 | Device Serial Number | unitless | Format: 8 characters string. This variable is buffered and sent repeatedly by the DDI with each data message. | String | DeviceInfo | Common
1927 | Instrument Code | unitless | Format: 6 characters string | String | DeviceInfo | Common
2879 | Device Software Revision | unitless | Format: vv.r.c vv = software version r = software revision c = software correction | String | DeviceInfo | Common
3615 | Battery Charge | unitless | 58112 => battery charged | UShort | DeviceInfo | Common
3636 | Device Alarm Status | unitless | The alarms code are output as decimal values. Consult detailed help for more information on this variable | UShort | Alarms | Common
3849 | Alarm History | unitless | 16 last alarms | UShort (Multiset) | Alarms | Common
4201 | Product Descritpion | unitless | 4 characters for product identification | String | DeviceInfo | Common
4350 | Current Delivery Rate | mL/h | If the pump is OFF, the return rate is 0. | UShort | Settings | Common
5490 | Cumulated Infused Volume | mL |  | UShort | VitalSigns | Common
5508 | Current Volume Infused | mL |  | UShort | OtherMeasuredData | Common
9221 | Error Response From Device | unitless | Error code sent by the pump when an incorrect transmission happens. | UByte | DeviceInfo | Common
9267 | Device Pre-alarm Status | unitless | The alarms code are output as decimal values. Consult detailed help for more information on this variable | UShort | Alarms | Common
50201 | Device Type | unitless | This variable contains a fixed value "48 (0x30)" for Infusion Pump. | UByte | DeviceInfo | Common | Recommended Variables
53001 | Infusion Pump Type | unitless | Consult detailed help for more information on this variable | UByte | DeviceInfo | Common | Recommended Variables
53012 | Infusion Pump Serial Number | unitless | 8 characters string | String | DeviceInfo | Common | Recommended Variables
53021 | Infusion Pump Rack Index | unitless | The variable has a fixed value "1" because it is a standalone pump. | UByte | DeviceInfo | Common | Recommended Variables
53022 | Infusion Pump Slot Index | unitless | The variable has a fixed value "1" because it is a standalone pump. | UByte | DeviceInfo | Common | Recommended Variables
53101 | Generic Infusion Pump Mode | unitless | Consult detailed help for more information on this variable | UShort | OtherMeasuredData | Common | Recommended Variables
53111 | Actual Infusion Volumetric Rate | mL/h |  | Double | VitalSigns | Common | Recommended Variables
53121 | Total Volume Infused | mL | This variable is reset to "0" only if users do so manually on the pump panel. It is not automatically reset if the infusion has been interrupted and/or the pump has been switched off. | Double | VitalSigns | Common | Recommended Variables
53211 | Infusion Drug Name | unitless | This variable contains a fixed nonsignificant value "NA" | String | PatientInfo | Common | Recommended Variables
53221 | Infusion Volumetric Rate Setting | mL/h |  | Double | Settings | Common | Recommended Variables
53801 | Infusion Event | unitless | Consult detailed help for more information on this variable | UByte | Messages | Common | Recommended Variables

## Multifiltrate PRO
ID | Label | Unit | Description | Type | Group | Device | Source
754 | Device Name | unitless | This variable is obsolete for MultiFiltrate and might be removed in a future version. | String | DeviceInfo | Common | Treatment Status/Data Management
965 | Cumulated Therapy Time | min | Treatment time, running, total | Long | OtherMeasuredData | Multifiltrate PRO | Dialysis Course
1060 | Anticoagulant Cumulated Continuous Infused Volume | 0.1 mL | Anticoagulation, continuous, total | Long | VitalSigns | Multifiltrate PRO | Dialysis Course
1090 | Anticoagulant Cumulated Bolus Infused Volume | 0.1 mL | Anticoagulation bolus, total | Long | VitalSigns | Multifiltrate PRO | Dialysis Course
1728 | Ultrafiltration/Blood Flow Ratio | % |  | Long | VitalSigns | Multifiltrate PRO | Dialysis Course
1842 | Arterial Pressure Upper Alarm Limit | mmHg | This variable is also sent as alarm limit of variableID 7438. | Long | Settings | Multifiltrate PRO | Dialysis Settings
1843 | Arterial Pressure Lower Alarm Limit | mmHg | This variable is also sent as alarm limit of variableID 7438. | Long | Settings | Multifiltrate PRO | Dialysis Settings
1913 | Device Serial Number | unitless |  | String | DeviceInfo | Common | Treatment Status
1929 | Patient ID | unitless |  | String | PatientInfo | Common | Administrative Data
2022 | Device System Code | unitless |  | String | DeviceInfo | Multifiltrate PRO | Treatment Status
2537 | Citrate Calcium Activation Status | unitless | Possible values: | Boolean | DeviceInfo | Multifiltrate PRO | Dialysis Settings
1: CiCa requested
0: CiCa not requested.
2539 | Therapy Pause Start Time | unitless | Format: YYYYMMDDhhmmss | String | OtherMeasuredData | Multifiltrate PRO | Treatment Status
2855 | Device Status | unitless | Consult detailed help for more information on this variable | UShort | DeviceInfo | Multifiltrate PRO | Treatment Status
2879 | Device Software Revision | unitless |  | String | DeviceInfo | Common | Treatment Status
2961 | Therapy Pause End Time | unitless | Format: YYYYMMDDhhmmss | String | OtherMeasuredData | Multifiltrate PRO | Treatment Status
2969 | Disposable Lifetime | min | The resolution of this variable is 10 min. This variable is sent from the device to the DDI on a regular basis, but the device will keep on sending the same value until a minimal change of 10 min is observed. | Long | DeviceInfo | Multifiltrate PRO | Dialysis Course
3003 | Blood Flow Rate Setting | mL/min |  | Long | Settings | Common | Dialysis Settings
3004 | Replacement Flow Rate Setting | mL/h | Substitution rate setting | Long | Settings | Multifiltrate PRO | Dialysis Settings
3005 | Dialysate Flow Rate Setting | mL/h |  | Long | Settings | Common | Dialysis Settings
3006 | Fluid Removal Volume Setting | mL | UF goal setting | Long | Settings | Common | Dialysis Settings
3016 | Transmembrane Pressure Upper Alarm Limit | mmHg | This variable is also sent as alarm limit of variableID 3029. | Long | Settings | Multifiltrate PRO | Dialysis Settings
3026 | Filter Pod Pressure | mmHg | Filtrate pressure | Long | OtherMeasuredData | Multifiltrate PRO | Dialysis Course
For Multifiltrate PRO: The resolution of this variable is 10 mmHg. This variable is sent from the device to the DDI on a regular basis, but the device will keep on sending the same value until a minimal change of 10 mmHg is observed.
3029 | Transmembrane Pressure | mmHg | This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. | Long | VitalSigns | Multifiltrate PRO | Dialysis Course
For Multifiltrate PRO: The resolution of this variable is 10 mmHg. This variable is sent from the device to the DDI on a regular basis, but the device will keep on sending the same value until a minimal change of 10 mmHg is observed.
3050 | Therapy Type | unitless | Consult detailed help for more information on this variable | String | DeviceInfo | Multifiltrate PRO | Treatment Status
3248 | Generic Venous Line Pressure | mmHg | This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. | Long | VitalSigns | Multifiltrate PRO | Dialysis Course
For Multifiltrate PRO: The resolution of this variable is 10 mmHg. This variable is sent from the device to the DDI on a regular basis, but the device will keep on sending the same value until a minimal change of 10 mmHg is observed.
3626 | Therapy Ending Time | unitless |  | String | OtherMeasuredData | Multifiltrate PRO | Treatment Status
3628 | Therapy Starting Time | unitless | Date/time in the format of "YYYYMMDDhhmmss" | String | OtherMeasuredData | Common | Treatment Status
3829 | Device Total Operation Time | unitless | Format: [hours]:[minutes] | String | DeviceInfo | Multifiltrate PRO | Treatment Status
4226 | Replacement Cumulated Volume | mL | Substituate-Total, total | Long | VitalSigns | Multifiltrate PRO | Dialysis Course
4229 | Fluid Removal Cumulated Volume | mL | Effective total withdrawal, total | Long | VitalSigns | Common | Dialysis Course
4573 | Dialysis Filter Lifetime | min | Filter Lifetime | Long | DeviceInfo | Multifiltrate PRO | Dialysis Course
For Multifiltrate PRO: The resolution of this variable is 10 min. This variable is sent from the device to the DDI on a regular basis, but the device will keep on sending the same value until a minimal change of 10 min is observed.
5860 | Blood Sensed Status | unitless | 0=no blood sensed 1=blood sensed | Long | DeviceInfo | Multifiltrate PRO | Dialysis Course
5868 | Fluid Removal Rate Setting | mL/h | UF rate setting | Long | Settings | Common | Dialysis Settings
5869 | Dialysate Temperature Setting | 0.1 °C |  | Long | Settings | Common | Dialysis Settings
5870 | Plasma Volume Setting | mL |  | Long | Settings | Multifiltrate PRO | Dialysis Settings
5871 | Citrate Dose Setting | 0.01 mmol/L |  | Long | Settings | Multifiltrate PRO | Dialysis Settings
5872 | Calcium Dose Setting | 0.01 mmol/L |  | Long | Settings | Multifiltrate PRO | Dialysis Settings
5875 | Citrate Flow Rate | mL/h |  | Long | VitalSigns | Multifiltrate PRO | Dialysis Course
5876 | Calcium Flow Rate | mL/h |  | Long | VitalSigns | Multifiltrate PRO | Dialysis Course
5878 | Citrate Cumulated Volume | 0.1 mL |  | Long | VitalSigns | Multifiltrate PRO | Dialysis Course
5880 | Calcium Cumulated Volume | 0.1 mL |  | Long | VitalSigns | Multifiltrate PRO | Dialysis Course
5882 | Filtration Fluid Cumulated Balance | mL | DF-Total, total balance | Long | VitalSigns | Multifiltrate PRO | Dialysis Course
5884 | Replacement Cumulated Bolus Volume | mL | Bolus-Substituate-Total, total | Long | VitalSigns | Multifiltrate PRO | Dialysis Course
6063 | Central Venous Pressure Lower Alarm Limit | mmHg | This variable is also sent as alarm limit of variableID 3248. | Long | Settings | Multifiltrate PRO | Dialysis Settings
6064 | Central Venous Pressure Upper Alarm Limit | mmHg | This variable is also sent as alarm limit of variableID 3248. | Long | Settings | Multifiltrate PRO | Dialysis Settings
6164 | General Alarm Code | unitless | Consult detailed help for more information on this variable | Long | Alarms | Multifiltrate PRO | Dialysis Course
6768 | Device Quick Test Status | unitless | T0 Test Status. Possible values: Start, End, E1, E2, E3. | String | DeviceInfo | Multifiltrate PRO | Treatment Status
7438 | Generic Arterial Line Pressure | mmHg | This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. | Long | VitalSigns | Multifiltrate PRO | Dialysis Course
For Multifiltrate PRO: The resolution of this variable is 10 mmHg. This variable is sent from the device to the DDI on a regular basis, but the device will keep on sending the same value until a minimal change of 10 mmHg is observed.
8809 | Anticoagulant Infusion Rate Setting | 0.1 mL/h | Heparin continuous rate | Long | Settings | Common | Dialysis Settings
8817 | Patient Weight Cumulated Value | g | Absolute balancing error | Long | VitalSigns | Multifiltrate PRO | Dialysis Course
8840 | Patient Sensor Value 1 | g | Weight of scale 1 | Long | OtherMeasuredData | Multifiltrate PRO | Dialysis Course
For Multifiltrate PRO: The resolution of this variable is 10 g. This variable is sent from the device to the DDI on a regular basis, but the device will keep on sending the same value until a minimal change of 10 g is observed. The value of "0" will always be sent regardless of the delta change.
8841 | Patient Sensor Value 2 | g | Weight of scale 2 | Long | OtherMeasuredData | Multifiltrate PRO | Dialysis Course
For Multifiltrate PRO: The resolution of this variable is 10 g. This variable is sent from the device to the DDI on a regular basis, but the device will keep on sending the same value until a minimal change of 10 g is observed. The value of "0" will always be sent regardless of the delta change.
8842 | Patient Sensor Value 3 | g | Weight of scale 3 | Long | OtherMeasuredData | Multifiltrate PRO | Dialysis Course
For Multifiltrate PRO: The resolution of this variable is 10 g. This variable is sent from the device to the DDI on a regular basis, but the device will keep on sending the same value until a minimal change of 10 g is observed. The value of "0" will always be sent regardless of the delta change.
8843 | Patient Sensor Value 4 | g | Weight of scale 4 | Long | OtherMeasuredData | Multifiltrate PRO | Dialysis Course
For Multifiltrate PRO: The resolution of this variable is 10 g. This variable is sent from the device to the DDI on a regular basis, but the device will keep on sending the same value until a minimal change of 10 g is observed. The value of "0" will always be sent regardless of the delta change.
8853 | Patient Fluid Balance | mL | Balance, total | Long | VitalSigns | Multifiltrate PRO | Dialysis Course
8856 | Therapy Operation Status | unitless | Consult detailed help for more information on this variable | String | DeviceInfo | Common | Treatment Status
8967 | Case ID | unitless |  | String | PatientInfo | Common | Administrative Data
9643 | Blood Flow Rate | mL/min | Actual blood flow This variable contains value "0" when blood pump is stopped. | Long | VitalSigns | Multifiltrate PRO | Dialysis Course
For Multifiltrate PRO: The resolution of this variable is 10 mL/min. This variable is sent from the device to the DDI on a regular basis, but the device will keep on sending the same value until a minimal change of 10 mL/min is observed. The value of "0" will always be sent regardless of the delta change.
9644 | Prefilter Pod Pressure | mmHg | Pre-Hemofilter pressure | Long | OtherMeasuredData | Multifiltrate PRO | Dialysis Course
For Multifiltrate PRO: This variable is associated with LowerAlarmLimit and UpperAlarmLimit attributes.
The resolution of this variable is 10 mmHg. This variable is sent from the device to the DDI on a regular basis, but the device will keep on sending the same value until a minimal change of 10 mmHg is observed.
9648 | Transmembrane Pressure Lower Alarm Limit | mmHg | This variable is also sent as alarm limit of variableID 3029. | Long | Settings | Multifiltrate PRO | Dialysis Settings

## 5008
ID | Label | Unit | Description | Type | Group | Device | Source
1 | Heart Rate | 1/min | This variable is associated with LowerAlarmLimit and UpperAlarmLimit attributes. | Long | VitalSigns | Common
2 | NIBP Systolic | mmHg | This variable is associated with LowerAlarmLimit and UpperAlarmLimit attributes. | Long | VitalSigns | Common
3 | NIBP Diastolic | mmHg | This variable is associated with LowerAlarmLimit and UpperAlarmLimit attributes. | Long | VitalSigns | Common
4 | NIBP Mean | mmHg | This variable is associated with LowerAlarmLimit and UpperAlarmLimit attributes. | Long | VitalSigns | Common
7 | Patient Delta Temperature | °C |  | Double | VitalSigns | Common
356 | Generic Dialysate Alarm | unitless | Value of "1" when alarm is active. | Boolean | Alarms | Common
357 | Urea Clearance Alarm | unitless | Value of "1" when alarm is active. | Boolean | Alarms | Common
358 | Blood Volume Alarm | unitless | Value of "1" when alarm is active. | Boolean | Alarms | Common
359 | Mixed Dilution Alarm | unitless | Value of "1" when alarm is active. | Boolean | Alarms | Common
393 | Single Needle Ratio Setting | % |  | Long | Settings | Common
444 | Target Treatement Effectiveness (Kt/V) Remaining Time | s |  | Long | VitalSigns | Common
448 | Shunt Blood Flow Rate | mL/min |  | Short | VitalSigns | Common
449 | Recirculation Measurement Active Status | unitless | Possible values: 0=off, 1=on | Boolean | DeviceInfo | Common
450 | Isolated Ultrafiltration Therapy Time | s |  | Long | VitalSigns | Common
451 | Isolated Ultrafiltration Rate | mL/h |  | Long | VitalSigns | Common
452 | Initial Hemoglobin | g/dL |  | Double | VitalSigns | Common
453 | Initial Protein Concentration | g/dL |  | Double | VitalSigns | Common
454 | Single Needle Time | min |  | Long | VitalSigns | Common
455 | Blood Outlet Pressure | mmHg |  | Short | VitalSigns | Common
456 | Dialysate Inlet Pressure | mmHg |  | Short | VitalSigns | Common
457 | Dialysate Outlet Pressure | mmHg |  | Short | VitalSigns | Common
458 | Maximum Transmembrane Pressure | mmHg |  | Short | VitalSigns | Common
459 | Auto Dialysate Flow Enabled Status | unitless | Possible values: 0=disabled, 1=enabled | Boolean | DeviceInfo | Common
460 | Auto Dialysate Flow Factor Setting | unitless |  | Double | Settings | Common
461 | Auto AntiCoagulant Enabled Status | unitless | Possible values: 0=disabled, 1=enabled | Boolean | DeviceInfo | Common
462 | Auto Anticoagulant Bolus Enabled Status | unitless | Possible values: 0=off, 1=on | Boolean | DeviceInfo | Common
463 | Auto Anticoagulant Bolus Rate Enabled Status | unitless | Possible values: 0=disabled, 1=enabled | Boolean | DeviceInfo | Common
464 | Single Needle Auto Rate Enabled Status | unitless | Possible values: 0=disabled, 1=enabled | Boolean | DeviceInfo | Common
468 | Replacement Pump Running Status | unitless | Possible values: 0=off, 1=on | Boolean | DeviceInfo | Common
508 | Clearance Measurement Interval Setting | s |  | UShort | Settings | Common
540 | Low Venous Line Pressure Alarm | unitless | Value of "1" when alarm is active. | Boolean | Alarms | Common
543 | High Venous Line Pressure Alarm | unitless | Value of "1" when alarm is active. | Boolean | Alarms | Common
546 | High Arterial Line Pressure Alarm | unitless | Value of "1" when alarm is active. | Boolean | Alarms | Common
547 | Low Arterial Line Pressure Alarm | unitless | Value of "1" when alarm is active. | Boolean | Alarms | Common
558 | Heparin Pump Alarm | unitless | Value of "1" when alarm is active. | Boolean | Alarms | Common
569 | Patient Specific Filtration Ratio Setting | % |  | UShort | Settings | Common
576 | Average Dialysate Temperature | °C |  | Double | VitalSigns | Common
610 | Cuff Inflation Pressure | mmHg |  | UShort | OtherMeasuredData | Common
676 | Bicarbonate Concentration (HCO3) | mmol/L |  | Double | VitalSigns | Common
754 | Device Name | unitless |  | String | DeviceInfo | Common
757 | Patient Weight | kg |  | Double | PatientInfo | Common
762 | Predilution Enable Status | unitless | Internal status for the mixed HDF control | UShort | DeviceInfo | Common
Consult detailed help for more information on this variable
845 | Clearance Measurement Enabled Status | unitless | Possible values: 0=off, 1=on | Boolean | DeviceInfo | Common
846 | Maximum Ultrafiltration Rate Setting | mL/h |  | Long | Settings | Common
847 | Ultrafiltration Goal Positive Deviation Setting | mL |  | Long | Settings | Common
848 | Ultrafiltration Goal Negative Deviation Setting | mL |  | Long | Settings | Common
851 | Residual Blood Volume Active Status | unitless | Possible values: 0=off, 1=on | Boolean | DeviceInfo | Common
852 | Temperature Change Rate Setting | °C/h |  | Double | Settings | Common
859 | Reinfusion Type Setting | unitless | Consult detailed help for more information on this variable | UShort | Settings | Common
860 | Auto Reinfusion Active Status | unitless | Possible values: 0=off, 1=on | Boolean | DeviceInfo | Common
914 | Low Volume Treatment Active Status | unitless | Possible values: 0=off, 1=on | Boolean | DeviceInfo | Common
955 | Battery Running Alarm | unitless | Value of "1" when alarm is active. | Boolean | Alarms | Common
965 | Cumulated Therapy Time | s |  | Long | OtherMeasuredData | Common
966 | Current Therapy Time | s |  | Long | OtherMeasuredData | Common
967 | Current Therapy ID | unitless |  | String | PatientInfo | Common
982 | Device IP Address | unitless |  | String | DeviceInfo | Common
1011 | Therapy Running Status | unitless | Consult detailed help for more information on this variable | String | DeviceInfo | Common
1043 | Manually Entered Hematocrit | % |  | UShort | Settings | Common
1090 | Anticoagulant Cumulated Bolus Infused Volume | mL |  | Double | VitalSigns | Common
1100 | Average Anticoagulant Rate | mL/h |  | Double | VitalSigns | Common
1101 | Single Needle Stroke Volume | mL |  | Long | Settings | Common
1102 | Average Urea Clearance | mL/min |  | Long | VitalSigns | Common
1103 | Online Bolus Volume | mL |  | Long | VitalSigns | Common
1104 | Online Bolus Rate Setting | mL/min |  | Long | Settings | Common
1220 | Patient Birth Date | unitless | Format: YYYYMMDD | String | PatientInfo | Common
1246 | Hematocrit | % |  | Double | VitalSigns | Common
1463 | Hemoglobin (Hb) | g/dL |  | Double | VitalSigns | Common
1510 | Concentrate Source | unitless | Consult detailed help for more information on this variable | Long | DeviceInfo | Common
1525 | Maximum Fluid Removal Rate Setting | mL/h |  | Long | Settings | Common
1591 | Average Blood Flow Rate | mL/min |  | Long | VitalSigns | Common
1602 | Reinfusion Volume | mL |  | Long | VitalSigns | Common
1803 | Device Alarm Global Status | unitless | Format: Bit 2: 0 --> Alarm | String | Alarms | Common
1822 | Blood Pump Alarm | unitless | Value of "1" when alarm is active. | Boolean | Alarms | Common
1913 | Device Serial Number | unitless |  | String | DeviceInfo | Common
1929 | Patient ID | unitless |  | String | PatientInfo | Common
2003 | Therapy Options | unitless | Consult detailed help for more information on this variable | Long | DeviceInfo | Common
2486 | NIBP Interval Time | min |  | Long | DeviceInfo | Common
2499 | Relative Blood Volume Change Limit | % |  | Long | Settings | Common
2590 | Generic Device Alarm | unitless | Format: list of maximum 16 alarm IDs separated with a "¦". | String | Alarms | Common
Consult detailed help for more information on this variable
2847 | Urea Distribution Volume Setting | L |  | Double | Settings | Common
2848 | Auto Ultrafiltration Control Activation Status | unitless | Possible values: 0=off, 1=on | Boolean | Settings | Common
2879 | Device Software Revision | unitless |  | String | DeviceInfo | Common
2901 | Patient First Name | unitless |  | String | PatientInfo | Common
3003 | Blood Flow Rate Setting | mL/min |  | UShort | Settings | Common
3004 | Replacement Flow Rate Setting | mL/min |  | Long | Settings | Common
3005 | Dialysate Flow Rate Setting | mL/min |  | ULong | Settings | Common
3006 | Fluid Removal Volume Setting | mL |  | Long | Settings | Common
3029 | Transmembrane Pressure | mmHg |  | Long | VitalSigns | Common
3070 | Patient Profile | unitless | Na profile number | UByte | PatientInfo | Common
3071 | Therapy Profile | unitless |  | UByte | PatientInfo | Common
3100 | Dialysate Sodium Start Concentration Setting | mmol/L |  | Double | Settings | Common
3168 | Fluid Removal Actual Rate | mL/h |  | Long | VitalSigns | Common
3189 | Average Replacement Flow Rate | mL/min |  | Long | VitalSigns | Common
3190 | Minimum Relative Treated Blood Volume | % |  | Double | VitalSigns | Common
3194 | Auto Replacement Control Activation Status | unitless | Possible values: 0=disabled, 1=enabled | Boolean | Settings | Common
3237 | Relative Blood Volume Change | % |  | Double | VitalSigns | Common
3239 | Replacement Volume Setting | L |  | Double | Settings | Common
3248 | Generic Venous Line Pressure | mmHg |  | Long | VitalSigns | Common
3252 | Total Replacement Cumulated Volume | L | Expected total substitution volume (pre + post) at the end of the treatment. | Double | VitalSigns | Common
3253 | Blood Circuit Pressure Alarm | unitless | Value of "1" when alarm is active. | Boolean | Alarms | Common
3263 | Recirculation Percentage | % |  | Double | VitalSigns | Common
3269 | Filtration Ratio | % |  | UShort | VitalSigns | Common
3417 | Ultrafiltration Remaining Time | min | Remaining ISO UF time. | Long | VitalSigns | Common
3430 | Total Replacement Flow Rate | mL/min |  | Long | VitalSigns | Common
3437 | Replacement Post-delivered Percentage | % |  | UShort | VitalSigns | Common
3575 | Arterial Temperature | °C |  | Double | VitalSigns | Common
3576 | Venous Temperature | °C |  | Double | VitalSigns | Common
3590 | Device Type | unitless |  | String | DeviceInfo | Common
3626 | Therapy Ending Time | unitless | Format: YYYYMMDDhhmm | String | OtherMeasuredData | Common
3628 | Therapy Starting Time | unitless | Format: YYYYMMDDhhmm | String | OtherMeasuredData | Common
4229 | Fluid Removal Cumulated Volume | mL |  | ULong | VitalSigns | Common
4238 | Single Needle Max Pressure Setting | mmHg |  | Long | Settings | Common
4239 | Single Needle Min Pressure Setting | mmHg |  | Long | Settings | Common
4240 | Plasma Sodium (Na) | mmol/L | For 6008 only: This variable will contain the Online Clearance Monitoring plasma sodium concentration when sodium management is being used. | Double | VitalSigns | Common
4347 | Therapy Remaining Time | min | Remaining UF time including ISO UF time. | ULong | OtherMeasuredData | Common
4365 | Device Communication Mode | unitless | Consult detailed help for more information on this variable | UByte | DeviceInfo | Common
4570 | Dialyzer Name | unitless |  | String | DeviceInfo | Common
4579 | Dialysate Name | unitless |  | String | PatientInfo | Common
4592 | Dialysate Average Flow Rate | mL/min |  | ULong | VitalSigns | Common
4601 | Dialysate Sodium (Na+) Concentration Setting | mmol/L |  | Double | Settings | Common
4605 | Dialysate Bicarbonate (HCO3) Concentration Setting | mmol/L |  | Double | Settings | Common
4608 | Urea Clearance | mL/min |  | Short | VitalSigns | Common
4609 | Fluid Removal Average Rate | mL/h |  | Long | VitalSigns | Common
4610 | Anticoagulant Stopping Interval Setting | s |  | Long | Settings | Common
4614 | Target Treatement Effectiveness (Kt/V) Setting | unitless |  | Double | Settings | Common
4615 | Initial Hematocrit | % |  | Double | VitalSigns | Common
4685 | Serum Protein | g/dL |  | Double | VitalSigns | Common
4705 | Total Protein | g/dL |  | Double | VitalSigns | Common
5318 | Blood Temperature Alarm | unitless | Value of "1" when alarm is active. | Boolean | Alarms | Common
5845 | Total Replacement Cumulated Volume | mL | Cumulated total substitution volume (pre + post). | Long | VitalSigns | Common
5846 | Single Needle Blood Pump Rate | mL/min |  | Long | VitalSigns | Common
5849 | Therapy Duration Setting | s |  | Long | Settings | Common
5850 | Anticoagulant Cumulated Volume | mL |  | Double | VitalSigns | Common
5853 | Delivered Treatment Effectiveness (Kt/V) | unitless |  | Double | VitalSigns | Common
5856 | Dialysate Flow Rate | mL/min |  | ULong | VitalSigns | Common
5858 | Disinfect Mode Active Status | unitless | Consult detailed help for more information on this variable | UByte | DeviceInfo | Common
5859 | Dialysis Mode Active Status | unitless | Consult detailed help for more information on this variable | Long | DeviceInfo | Common
5861 | Effective Blood Volume Processed (Kt) | L |  | Double | VitalSigns | Common
5863 | Dialysate Temperature | °C |  | Double | VitalSigns | Common
5868 | Fluid Removal Rate Setting | mL/h |  | Long | Settings | Common
5869 | Dialysate Temperature Setting | °C |  | Double | Settings | Common
5884 | Replacement Cumulated Bolus Volume | mL |  | Long | VitalSigns | Common
5899 | Actual Anticoagulant Rate | mL/h |  | Double | VitalSigns | Common
5905 | Therapy Processed Percentage | % |  | Byte | VitalSigns | Common
6328 | Blood Plasma Inlet Pressure | mmHg |  | Short | VitalSigns | Common
6334 | Temperature Module General Status | unitless | Possible values: 0=off, 1=on | Boolean | DeviceInfo | Common
6427 | Last Desinfection Time | unitless | Format: YYYYMMDDhhmm | String | DeviceInfo | Common
6429 | Last Therapy End Time | unitless | Format: YYYYMMDDhhmm | String | DeviceInfo | Common
6430 | Dialysis Pump Stop Status | unitless | Consult detailed help for more information on this variable | UShort | DeviceInfo | Common
7351 | Isolated Ultrafiltration Volume | mL |  | Long | VitalSigns | Common
7352 | Single Needle Stroke Volume | mL | Average single needle stroke volume | Long | VitalSigns | Common
7353 | Initial Ultrafiltration Factor Setting | unitless |  | UShort | Settings | Common
7438 | Generic Arterial Line Pressure | mmHg |  | Long | VitalSigns | Common
7979 | Dialysate Conductivity | mS/cm |  | Double | VitalSigns | Common
7980 | Dialysate Conductivity Alarm | unitless | Value of "1" when alarm is active. | Boolean | Alarms | Common
8244 | Bolus Manual Enable Status | unitless | Possible values: 0=disabled, 1=enabled | Boolean | DeviceInfo | Common
8262 | Plasma Filtration Fraction Setting | % |  | UShort | Settings | Common
8340 | Patient Family Name | unitless |  | String | PatientInfo | Common
8803 | Replacement Pre-delivered Rate Setting | mL/min |  | UShort | Settings | Common
8804 | Replacement Post-delivered Rate Setting | mL/min |  | UShort | Settings | Common
8807 | Anticoagulant Bolus Volume Setting | mL |  | Double | Settings | Common
8809 | Anticoagulant Infusion Rate Setting | mL/h |  | Double | Settings | Common
8813 | Processed Cumulated Blood Volume | L |  | Double | VitalSigns | Common
8815 | Replacement Post-delivered Cumulated Volume | mL |  | Long | VitalSigns | Common
8856 | Therapy Operation Status | unitless | Consult detailed help for more information on this variable | Short | DeviceInfo | Common
9522 | Device Manufacturer | unitless |  | String | DeviceInfo | Common
9643 | Blood Flow Rate | mL/min |  | Long | VitalSigns | Common
9662 | Device Location | unitless |  | String | DeviceInfo | Common

## Total Scale
ID | Label | Unit | Description | Type | Group | Device | Source
168 | Device ID | unitless | The globally unique identifier for a Bed entity as defined by Hill-Rom (e.g. “VersaCare-0x0001”). | String | DeviceInfo | Bed Data Server (Common) | BED
754 | Device Name | unitless | The bed model name. | String | DeviceInfo | Bed Data Server (Common) | BED
757 | Patient Weight | kg | The last patient weight captured by a caregiver from the bedside displayed in kilograms. -9999.9 -- 9999.9 | Double | PatientInfo | Bed Data Server (Common) | BED
980 | Base Device ID | unitless | The globally unique identifier for the BedConnector entity associated with this Bed entity or an empty string if no association currently exists. | String | DeviceInfo | Bed Data Server (Common) | BED
1831 | Bed Head Angle Alarm | unitless | The current status for the head angle alarm system. | Byte | Alarms | Bed Data Server (Common) | BED
1891 | Power Supply Status | unitless | The current status of the AC power system. | String | DeviceInfo | Bed Data Server (Common) | BED
1913 | Device Serial Number | unitless | The serial number of the bed communication controller. 0x000000000000 - 0xFFFFFFFFFFFF | String | DeviceInfo | Bed Data Server (Common) | BED
1929 | Patient ID | unitless | The globally unique identifier for the Patient entity associated with this Bed entity or an empty string if no association currently exists. | String | PatientInfo | Bed Data Server (Common) | BED
1930 | Patient Name | unitless | Easily recognizable patient name that can be used by authorized personnel to locate the specific physical patient (e.g. j_doe…). | String | PatientInfo | Bed Data Server (Common) | PATIENT
1931 | Patient Location ID | unitless | The globally unique identifier for the Location entity associated with this Bed entity or an empty string if no association currently exists. | String | PatientInfo | Bed Data Server (Common) | BED
1932 | Patient Location Text | unitless | Easily recognizable location name that can be used by authorized personnel to locate the specific physical location (e.g. Tower1_WingC_Floor2_Room103). | String | PatientInfo | Bed Data Server (Common) | LOCATION
1946 | Remote Alarm Application Mode | unitless | The current configuration mode of the SafeView alarm system. | String | DeviceInfo | Bed Data Server (Common) | BED
1947 | Remote Alarm Application Indicator Status | unitless | The current status of the SafeView alarm system indicators. | String | DeviceInfo | Bed Data Server (Common) | BED
1948 | Patient Weight Scale Activity | unitless | The last command to be successfully processed by the weigh scale system. | String | DeviceInfo | Bed Data Server (Common) | BED
1957 | Remote CIS Application Status | unitless | The current status for the NaviCare alerts switch. | String | DeviceInfo | Bed Data Server (Common) | BED
1958 | Remote CIS Application Indicator Status | unitless | The current status for the NaviCare alerts indicator. | String | DeviceInfo | Bed Data Server (Common) | BED
1963 | Nurse Call Indicator Status | unitless | The current status for the nurse call indicator. | String | DeviceInfo | Bed Data Server (Common) | BED
1964 | Nurse Answer Indicator Status | unitless | The current status for the nurse answer indicator. | String | DeviceInfo | Bed Data Server (Common) | BED
1986 | Bed Height Position | unitless | The bed’s physical articulated position. | String | DeviceInfo | Bed Data Server (Common) | BED
1987 | Bed Lockout Status | unitless |  | String | DeviceInfo | Bed Data Server (Common) | BED
1988 | Bed Head Rails Position | unitless | The physical positions of the head-end siderails. | String | DeviceInfo | Bed Data Server (Common) | BED
1989 | Bed Foot RailsPosition | unitless | The physical positions of the foot-end siderails. | String | DeviceInfo | Bed Data Server (Common) | BED
1990 | Bed Surface Support Mode | unitless | The current configuration mode of the patient support surface system. | String | DeviceInfo | Bed Data Server (Common) | BED
1991 | Bed Emergency CPR Status | unitless | The current status for the CPR switch. | String | DeviceInfo | Bed Data Server (Common) | BED
1992 | Bed Clearance Status | unitless | The current status for the Bed Cleaned switch. | String | DeviceInfo | Bed Data Server (Common) | BED
1993 | Bed Rotation Therapy Status | unitless | The current status for the rotation therapy system. | String | DeviceInfo | Bed Data Server (Common) | BED
1994 | Bed Percussion Therapy Status | unitless | The current status for the percussion therapy system. | String | DeviceInfo | Bed Data Server (Common) | BED
1995 | Bed Vibration Therapy Status | unitless | The current status for the vibration therapy system. | String | DeviceInfo | Bed Data Server (Common) | BED
1996 | Patient Environment Control Activity | unitless | The last command completed by the patient environment control system. | String | DeviceInfo | Bed Data Server (Common) | BED
1997 | Patient Environment Activity Cleared Status | unitless | The last command completed by the patient environment control system. | Byte | DeviceInfo | Bed Data Server (Common) | BED
1998 | Bed Head Angle Alarm Mode | unitless | The current mode configuration for the head angle alarm system. | String | DeviceInfo | Bed Data Server (Common) | BED
1999 | Bed Head Angle Alarm Silence Status | unitless | The current mode configuration for the audible annunciation portion of the head angle alarm system. | String | DeviceInfo | Bed Data Server (Common) | BED
2000 | Bed Clearance Indicator Status | unitless | The current status for the bed cleaned indicator. | String | DeviceInfo | Bed Data Server (Common) | BED
2001 | SmartBed Patient Movement State | unitless | The current direction of patient activity as detected by the patient positioning monitor (PPM) alarm system. | String | OtherMeasuredData | Bed Data Server (Common) | BED
2477 | Battery Module Status | unitless | The current status of the battery power system. | String | DeviceInfo | Bed Data Server (Common) | BED
2879 | Device Software Revision | unitless | The software revision of the bed communication controller. 0.00 -- 9.99 | String | DeviceInfo | Bed Data Server (Common) | BED
3593 | Device Communications Status | unitless | The current connection state between the Bed entity and the Bed Data Server application. | String | DeviceInfo | Bed Data Server (Common) | BED
3833 | Device Protocol Mode | unitless | The connectivity protocol supported by the bed communication controller. | String | DeviceInfo | Bed Data Server (Common) | BED
4274 | Service Message | unitless | The current status of the service required monitor system. | String | Group not set yet | Bed Data Server (Common) | BED
4281 | Nurse Call Status | unitless | The current status for the nurse call switch. | String | DeviceInfo | Bed Data Server (Common) | BED
9526 | Bed Head Angle | ° | The articulated angle of the bed’s head section in degrees. -360.0 -- 360.0 | Double | DeviceInfo | Bed Data Server (Common) | BED
9529 | Current Patient Weight | kg | The last patient weight captured by the dynamic weight scale system displayed in kilograms. -9999.9 -- 9999.9 | Double | VitalSigns | Bed Data Server (Common) | BED
9530 | Bed Brake Status | unitless | The current status for the brake switch. | String | DeviceInfo | Bed Data Server (Common) | BED
9536 | SmartBed Working Mode | unitless | The current configuration mode of the patient positioning monitor (PPM) alarm system. | String | Settings | Bed Data Server (Common) | BED
9537 | SmartBed General Alarm | unitless | The current status of the patient positioning monitor (PPM) alarm system. | Byte | Alarms | Bed Data Server (Common) | BED
9543 | SmartBed Patient Position State | unitless | The current level of patient activity as detected by the patient positioning monitor (PPM) alarm system. | String | OtherMeasuredData | Bed Data Server (Common) | BED

## Expression IP5 et MR400
ID | Label | Unit | Description | Type | Group | Device | Source
1 | Heart Rate | 1/min | This variable is associated with a SourceID identifier. For Expression, the SourceID is always ""MEAS_UNKNOWN" if using Expression Plus protocol. | Double | VitalSigns | Common | HR
2 | NIBP Systolic | - | Possible units: mmHg or kPa. | Double | VitalSigns | Common | NIBP
If PCMCIA protocol is used, the unit code of this variable is always 0 (unit unknown). A MeasurementTime attribute is also associated with this variable that is generated by the DDI which corresponds to the Smartlinx Server time when a new measurement is detected (i.e. either systolic, diastolic, or mean value has changed). When two consecutive NIBP measurements on the device yield identical values the DDI is not able to detect that a new measurement occurred and the MeasurementTime attribute value of this variable is not updated.
If Expression Plus protocol is used, this variable is associated with a MeasurementTime attribute, which is transmitted by the device.
3 | NIBP Diastolic | - | Possible units: mmHg or kPa. | Double | VitalSigns | Common | NIBP
If PCMCIA protocol is used, the unit code of this variable is always 0 (unit unknown). A MeasurementTime attribute is also associated with this variable that is generated by the DDI which corresponds to the Smartlinx Server time when a new measurement is detected (i.e. either systolic, diastolic, or mean value has changed). When two consecutive NIBP measurements on the device yield identical values the DDI is not able to detect that a new measurement occurred and the MeasurementTime attribute value of this variable is not updated.
If Expression Plus protocol is used, this variable is associated with a MeasurementTime attribute, which is transmitted by the device.
4 | NIBP Mean | - | Possible units: mmHg or kPa. | Double | VitalSigns | Common | NIBP
If PCMCIA protocol is used, the unit code of this variable is always 0 (unit unknown). A MeasurementTime attribute is also associated with this variable that is generated by the DDI which corresponds to the Smartlinx Server time when a new measurement is detected (i.e. either systolic, diastolic, or mean value has changed). When two consecutive NIBP measurements on the device yield identical values the DDI is not able to detect that a new measurement occurred and the MeasurementTime attribute value of this variable is not updated.
If Expression Plus protocol is used, this variable is associated with a MeasurementTime attribute, which is transmitted by the device.
8 | INVP1 Systolic | - | Possible units: mmHg or kPa. If PCMCIA protocol is used, the unit code of this variable is always 0 (unit unknown). | Double | VitalSigns | Common | IBP
9 | INVP1 Diastolic | - | Possible units: mmHg or kPa. If PCMCIA protocol is used, the unit code of this variable is always 0 (unit unknown). | Double | VitalSigns | Common | IBP
10 | INVP1 Mean | - | Possible units: mmHg or kPa. If PCMCIA protocol is used, the unit code of this variable is always 0 (unit unknown). | Double | VitalSigns | Common | IBP
11 | INVP2 Systolic | - | Possible units: mmHg or kPa. If PCMCIA protocol is used, the unit code of this variable is always 0 (unit unknown). | Double | VitalSigns | Common | IBP
12 | INVP2 Diastolic | - | Possible units: mmHg or kPa. If PCMCIA protocol is used, the unit code of this variable is always 0 (unit unknown). | Double | VitalSigns | Common | IBP
13 | INVP2 Mean | - | Possible units: mmHg or kPa. If PCMCIA protocol is used, the unit code of this variable is always 0 (unit unknown). | Double | VitalSigns | Common | IBP
14 | SpO2 | % |  | Double | VitalSigns | Common | SpO2
20 | End Tidal CO2 (EtCO2) | - | Possible units: mmHg or kPa. If PCMCIA protocol is used, the unit code of this variable is always 0 (unit unknown). | Double | VitalSigns | Common | Resp
21 | Inspired CO2 (InCO2) | - | Possible units: mmHg or kPa. If PCMCIA protocol is used, the unit code of this variable is always 0 (unit unknown). | Double | VitalSigns | Common | Resp
22 | Respiration Rate | 1/min |  | Double | VitalSigns | Common | Resp
168 | Device ID | unitless | 1–8 characters length alphanumeric string. For MR Patient Care Portal 5000, the length can be up to 31 characters. | String | DeviceInfo | Expression IP5 | Monitor Info (EXPRESS Plus, Express Connect Only)
419 | Inspired O2 (FiO2) | % |  | Double | VitalSigns | Common | Resp
426 | Inspired N2O | % |  | Double | VitalSigns | Expression IP5 | Agent (EXPRESS Plus, Express Connect Only)
427 | End Tidal N2O | % |  | Double | VitalSigns | Common | Agent
1929 | Patient ID | unitless | 1–16 characters length alphanumeric string. | String | PatientInfo | Expression IP5 | Patient Info (EXPRESS Plus, Express Connect Only)
2255 | Device Date/Time | unitless | Format: YYYY-MM-DD HH:MM:SS | String | DeviceInfo | Common | Device Info
2889 | Inspired Agent Concentration | % | This Variable is associated with a SourceID identifier. | Double | VitalSigns | Common | Agent
2890 | Expired Agent Concentration | % | This Variable is associated with a SourceID identifier. | Double | VitalSigns | Common | Agent
2907 | Patient Temperature | - | Possible units: °C or °F. | Double | VitalSigns | Common | Temp
4008 | Inspiratory 2nd Agent | % | This Variable is associated with a SourceID identifier. | Double | VitalSigns | Expression IP5 | Agent (EXPRESS Plus, Express Connect Only)
4009 | Expiratory 2nd Agent | % | This Variable is associated with a SourceID identifier. | Double | VitalSigns | Expression IP5 | Agent (EXPRESS Plus, Express Connect Only)

## Criticool
ID | Label | Unit | Description | Type | Group | Device | Source
1378 | Device Configuration | unitless | For CritiCool: This variable contains information on the fixed step rate at which the patient gradually warms up. For Allon: This variable contains information on if the temperature alarms are on or off.Consult detailed help for more information on this variable | String | DeviceInfo | Common
2255 | Device Date/Time | unitless | Format: YYYYMMDDHHMMSS | String | DeviceInfo | Common
2855 | Device Status | unitless | This variable indicated if the water flow is through the patient wrap (garment) or is circulating internally (bypass). Possible values: 0 = Bypass 1 = Garment | Boolean | DeviceInfo | Common
2879 | Device Software Revision | unitless | Format: XY.ZZ | String | DeviceInfo | Common
2907 | Patient Temperature | °C | This variable is associated with SourceID identifier. | Double | VitalSigns | Common
3291 | Patient Temperature Setting | °C |  | Double | Settings | Common
4596 | System Pressure High Alarm | unitless | Possible values: 0 = Active 1 = Not Active | Boolean | Alarms | Common
5235 | Low Water Level Alarm | unitless | Low water level alarm Possible values: 0 = On 1 = Off | Boolean | Alarms | Common
7452 | Device Operation Mode | unitless | Consult detailed help for more information on this variable | String | DeviceInfo | Common
8678 | General Temperature Alarm | unitless | Consult detailed help for more information on this variable | Double | Alarms | Common

## X3
ID | Label | Unit | Description | Type | Group | Device | Source
2 | NIBP Systolic | - |  | Double | VitalSigns | Common | NBP
3 | NIBP Diastolic | - |  | Double | VitalSigns | Common | NBP
4 | NIBP Mean | - |  | Double | VitalSigns | Common | NBP
7 | Patient Delta Temperature | - |  | Double | VitalSigns | Common | DeltaTemp
14 | SpO2 | - |  | Double | VitalSigns | Common | SpO2
16 | NIBP Pulse Rate | - |  | Double | VitalSigns | IntelliVue X3 | Pulse
19 | ECG Heart Rate | - |  | Double | VitalSigns | Common | ECG: HR
22 | Respiration Rate | - |  | Double | VitalSigns | Common | RR
26 | Device Inactivation Status | unitless | Possible values: "enabled" and "audio-paused". See "Multiset Alarms" on help file. | String | DeviceInfo | Common | Alarms
28 | Active Alarms | unitless | See "Multiset Alarms" on help file. | Double (Multiset) | Alarms | Common | Alarms
Consult detailed help for more information on this variable
29 | Latched Alarms | unitless | See "Multiset Alarms" on help file. | Double (Multiset) | Alarms | Common | Alarms
Consult detailed help for more information on this variable
33 | Inactive Alarms | unitless | See "Multiset Alarms" on help file. | Double (Multiset) | Alarms | Common | Alarms
Consult detailed help for more information on this variable
34 | General ST Level 1 | - | Physiological identifier: NOM_ECG_AMPL_ST_I | Double | VitalSigns | Common | ECG: ST
35 | General ST Level 2 | - | Physiological identifier: NOM_ECG_AMPL_ST_II | Double | VitalSigns | Common | ECG: ST
36 | General ST Level 3 | - | Physiological identifier: NOM_ECG_AMPL_ST_III | Double | VitalSigns | Common | ECG: ST
45 | Temperature 3 | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
46 | Temperature 4 | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
48 | O2 Saturation (SO2) | - |  | Double | VitalSigns | IntelliVue X3 | SO2
113 | Pleth Waveform | - |  | Double | Waveforms | IntelliVue X3 | Waves: Pleth
127 | ECG Measurement Status | - | ECG Rhythm Status | Double | OtherMeasuredData | Common | ECG Status (on IntelliVue)
Consult detailed help for more information on this variable
169 | Patient Age | - | Patient age. The Monitor device calculates the age of the patient from the birth data manually entered on the Monitor. | Double | PatientInfo | Common | Patient Demographics
170 | Patient Sex | unitless | Possible values: "unknown", "male", "female", and "unspecified". | String | PatientInfo | Common | Patient Demographics
405 | Generic O2 Saturation (1) | - |  | Double | VitalSigns | IntelliVue X3 | SO2 1
406 | Generic O2 Saturation (2) | - |  | Double | VitalSigns | IntelliVue X3 | SO2 2
414 | Cardiac Output | - |  | Double | VitalSigns | IntelliVue X3 | C.O.
465 | Fetal Heart Rate (FHR) | - | This variable is associated with ChannelID identifier. | Double | VitalSigns | IntelliVue X3 | DFHR/DFHR1/DFHR2/DFHR3
467 | Fetal Beat-to-Beat Rate (FHR) | - | This variable is associated with ChannelID identifier. | Double | VitalSigns | IntelliVue X3 | bDFHR/bDFHR1/bDFHR2/bDFHR3
469 | TOCO | - |  | Double | VitalSigns | IntelliVue X3 | Toco
497 | ECG Lead Name | - | Physiological identifier: NOM_ECG_AMPL_ST | Double | OtherMeasuredData | Common | Available only in revF or earlier
501 | Arterial Pressure (ART) Systolic | - |  | Double | VitalSigns | Common | IBP: ART
502 | Arterial Pressure (ART) Diastolic | - |  | Double | VitalSigns | Common | IBP: ART
503 | Arterial Pressure (ART) Mean | - |  | Double | VitalSigns | Common | IBP: ART
512 | Left Atrial Mean Pressure (LAP) | - |  | Double | VitalSigns | Common | IBP: LAP
519 | Central Venous Pressure (CVP) | - | Available on IntelliVue devices with firmware revision G or later. | Double | VitalSigns | Common | IBP: CVP
538 | Device Language | unitless | Consult detailed help for more information on this variable | String | DeviceInfo | Common | System Information
591 | ECG Lead II | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
592 | Airway N2O Concentration | - | Measurement from VueLink | Double | VitalSigns | IntelliVue X3 | N2O
593 | Inspired Desflurane | - |  | Double | VitalSigns | IntelliVue X3 | DES
595 | Inspired Sevoflurane | - |  | Double | VitalSigns | IntelliVue X3 | SEV
597 | Fresh Gas Agent Value | - |  | Double | VitalSigns | IntelliVue X3 | fgAGT
598 | Fresh Gas Halothane | - |  | Double | VitalSigns | IntelliVue X3 | fgHAL
599 | Fresh Gas Enflurane | - |  | Double | VitalSigns | IntelliVue X3 | fgENF
600 | Fresh Gas Isoflurane | - |  | Double | VitalSigns | IntelliVue X3 | fgISO
601 | Inspired Halothane | - |  | Double | VitalSigns | IntelliVue X3 | HAL
603 | Inspired Enflurane | - |  | Double | VitalSigns | IntelliVue X3 | ENF
605 | Inspired Isoflurane | - |  | Double | VitalSigns | IntelliVue X3 | ISO
608 | End Tidal CO2 | - |  | Double | VitalSigns | Common | CO2
616 | Airway Pressure | - | Measured in the inspiratory path. | Double | VitalSigns | IntelliVue X3 | AWPin
618 | Airway N2O Wave | - |  | Double | Waveforms | IntelliVue X3 | Waves: N2O
620 | Halothane Concentration Wave | - |  | Double | Waveforms | IntelliVue X3 | Waves: HAL
621 | Enflurane Concentration Wave | - |  | Double | Waveforms | IntelliVue X3 | Waves: ENF
622 | Isoflurane Concentration Wave | - |  | Double | Waveforms | IntelliVue X3 | Waves: ISO
623 | Desflurane Concentration Wave | - |  | Double | Waveforms | IntelliVue X3 | Waves: DES
624 | Sevoflurane Concentration Wave | - |  | Double | Waveforms | IntelliVue X3 | Waves: SEV
625 | Apnea Duration | - | This is the device's measurement with label="NLS_NOM_TIME_PD_APNEA" and physio_id="NOM_TIME_PD_APNEA". | Double | VitalSigns | IntelliVue X3 | ApneaD
626 | End Tidal CO2 | - |  | Double | VitalSigns | IntelliVue X3 | CRT
630 | Cardiac Beat-to-Beat Rate | - | Cardiac Beat-to-Beat Rate (NOM_ECG_CARD_BEAT_RATE_BTB). | Double | VitalSigns | Common | ECG: btbHR
633 | Airway Delta O2 Concentration | - |  | Double | VitalSigns | IntelliVue X3 | VENT: DeltaO2
649 | Cardiac Output Index | - |  | Double | VitalSigns | IntelliVue X3 | C.I.
650 | Stroke Volume | - |  | Double | VitalSigns | IntelliVue X3 | SV
651 | Stroke Volume Index | - |  | Double | VitalSigns | IntelliVue X3 | SI
654 | Systemic Vascular Resistance | - |  | Double | VitalSigns | IntelliVue X3 | SVR
669 | Normalized Ionized Calcium Concentration (Ca) | - |  | Double | VitalSigns | IntelliVue X3 | iCa-N
678 | Potassium Ion Concentration (K+) | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
696 | QT Interval (QT) | - |  | Double | VitalSigns | IntelliVue X3 | ECG: QT
764 | Device Specification ID | unitless | Variables 764, 765 and 4201 must be interpreted conjointly. | String (Multiset) | DeviceInfo | Common | System Information
Consult detailed help for more information on this variable
765 | Device Specification Type | unitless | Variables 764, 765 and 4201 must be interpreted conjointly. | String (Multiset) | DeviceInfo | Common | System Information
Consult detailed help for more information on this variable
794 | Pleth Wave (2) | - |  | Double | Waveforms | IntelliVue X3 | Waves: PLETH2
854 | Total Number of Used Disposable | - |  | Double | DeviceInfo | IntelliVue X3 | nFilt
946 | Cumulated Syringe Infused Volume | - |  | Double | OtherMeasuredData | IntelliVue X3 | SyrC
966 | Current Therapy Time | - |  | Double | OtherMeasuredData | IntelliVue X3 | FiltTi
1043 | Manually Entered Hematocrit | - |  | Double | Settings | IntelliVue X3 | scHCTi
1055 | Infusion Drug Name | - |  | Double | PatientInfo | IntelliVue X3 | sDrug
1072 | Infusion Time Remaining | - |  | Double | VitalSigns | IntelliVue X3 | InfTim
1095 | Infusion Dose Rate Setting | - |  | Double | Settings | IntelliVue X3 | sDosRt
1096 | Volume To Be Infused Original Setting | - |  | Double | Settings | IntelliVue X3 | sVtot
1110 | KVO Rate Setting | - |  | Double | Settings | IntelliVue X3 | sKVO
1186 | Inspired Minute Volume | - |  | Double | VitalSigns | IntelliVue X3 | MINVOL
1190 | I:E Ratio | - | This variable contains expiratory component of I:E Ratio assuming that the inspiratory component is "1". | Double | VitalSigns | IntelliVue X3 | I:E 1:
1208 | SpO2 | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
1220 | Patient Birth Date | unitless | Format: 14 ascii digits string of "YYYYMMDDhhmmss" | String | PatientInfo | Common | Patient Demographics
1234 | SpO2 (Pre-Ductal) | - |  | Double | VitalSigns | IntelliVue X3 | SpO2pr
1236 | SpO2 (Post-Ductal) | - |  | Double | VitalSigns | IntelliVue X3 | SpO2po
1244 | ECG Lead III | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
1246 | Hematocrit | - |  | Double | VitalSigns | IntelliVue X3 | Hct
1250 | Pulmonary Vascular Resistance | - |  | Double | VitalSigns | IntelliVue X3 | PVR
1251 | Body Surface Area | - |  | Double | PatientInfo | IntelliVue X3 | BSA
1255 | ECG Lead MCL1 | - | Physiological identifier: NOM_ECG_AMPL_ST_MCL | Double | VitalSigns | Common | ECG: ST
1257 | ECG Lead V1 | - | Physiological identifier: NOM_ECG_AMPL_ST_V1 | Double | VitalSigns | Common | ECG: ST
1258 | ECG Lead V6 | - | Physiological identifier: NOM_ECG_AMPL_ST_V6 | Double | VitalSigns | Common | ECG: ST
1259 | ECG Lead aVL | - | Physiological identifier: NOM_ECG_AMPL_ST_AVL | Double | VitalSigns | Common | ECG: ST
1260 | ECG Lead aVF | - | Physiological identifier: NOM_ECG_AMPL_ST_AVF | Double | VitalSigns | Common | ECG: ST
1261 | ECG Lead aVR | - | Physiological identifier: NOM_ECG_AMPL_ST_AVR | Double | VitalSigns | Common | ECG: ST
1262 | ECG Lead V2 | - | Physiological identifier: NOM_ECG_AMPL_ST_V2 | Double | VitalSigns | Common | ECG: ST
1263 | ECG Lead V3 | - | Physiological identifier: NOM_ECG_AMPL_ST_V3 | Double | VitalSigns | Common | ECG: ST
1264 | ECG Lead V4 | - | Physiological identifier: NOM_ECG_AMPL_ST_V4 | Double | VitalSigns | Common | ECG: ST
1265 | ECG Lead V5 | - | Physiological identifier: NOM_ECG_AMPL_ST_V5 | Double | VitalSigns | Common | ECG: ST
1283 | Patient Temperature | - | This variable is associated with a SourceID Identifier. | Double | VitalSigns | IntelliVue X3 | Infrared Temp
1286 | Micro Environment Air Temperature | - |  | Double | OtherMeasuredData | IntelliVue X3 | INCUBATOR: Air T
1306 | Tidal Volume | - |  | Double | VitalSigns | IntelliVue X3 | TV
1319 | Static Resistance | - |  | Double | VitalSigns | IntelliVue X3 | Rstat
1321 | Tidal Volume Setting | - |  | Double | Settings | IntelliVue X3 | sTV
1323 | Pressure Sensitivity Setting | - |  | Double | Settings | IntelliVue X3 | sSenPr
1324 | Sigh Interval Setting | - |  | Double | Settings | IntelliVue X3 | sTsigh
1325 | Sigh Volume Setting | - |  | Double | Settings | IntelliVue X3 | sVsigh
1332 | Pressure Support Level Above PEEP Setting | - |  | Double | Settings | IntelliVue X3 | sdPEEP
1333 | Spontaneous Minute Volume | - |  | Double | VitalSigns | IntelliVue X3 | VENT: SpMV
1340 | Negative Inspiratory Pressure | - |  | Double | VitalSigns | IntelliVue X3 | NgInsP
1342 | Peak Spontaneous Flow | - |  | Double | VitalSigns | IntelliVue X3 | SpPkFl
1343 | Flow-by Base Flow Setting | - |  | Double | Settings | IntelliVue X3 | VENT: sBasFl
1344 | Flow Sensitivity Setting | - |  | Double | Settings | IntelliVue X3 | VENT: sSenFl
1348 | Expiratory Time | - |  | Double | VitalSigns | IntelliVue X3 | ExpTi
1381 | Inspiratory Flow Setting | - |  | Double | Settings | IntelliVue X3 | sInsFl
1382 | Expiratory Flow Setting | - |  | Double | Settings | IntelliVue X3 | sExpFl
1383 | Inspiratory Flow Trigger Setting | - | Delivered by the Evita 2 VueLink Driver. | Double | Settings | IntelliVue X3 | VENT: sTrgFl
1387 | APRV Low Pressure Setting | - |  | Double | Settings | IntelliVue X3 | sPSVl
1389 | APRV High Pressure Setting | - |  | Double | Settings | IntelliVue X3 | sPSVh
1391 | Apnea Time Setting | - |  | Double | Settings | IntelliVue X3 | sAlnTi
1398 | Minute Volume Setting | - |  | Double | Settings | IntelliVue X3 | sMV
1418 | Positive End Expiratory Pressure (PEEP) | - |  | Double | VitalSigns | IntelliVue X3 | VENT: PEEP
1419 | Ventilator Work of Breathing | - |  | Double | VitalSigns | IntelliVue X3 | WOBv
1422 | Rapid Shallow Breathing Index | - |  | Double | VitalSigns | IntelliVue X3 | RSBI
1424 | Patient Work of Breathing | - |  | Double | VitalSigns | IntelliVue X3 | WOBp
1435 | Systemic Vascular Resistance Index | - |  | Double | VitalSigns | IntelliVue X3 | SVRI
1436 | Pulmonary Vascular Resistance Index | - |  | Double | VitalSigns | IntelliVue X3 | PVRI
1442 | Continuous Cardiac Output | - |  | Double | VitalSigns | IntelliVue X3 | CCO
1443 | Continuous Cardiac Output Index | - |  | Double | VitalSigns | IntelliVue X3 | CCI
1451 | Pump Infusion Status | - |  | Double | DeviceInfo | IntelliVue X3 | PuStat
1463 | Hemoglobin (Hb) | - | This variable is associated with SourceID identifier. | Double | VitalSigns | IntelliVue X3 | sHb
1464 | Pleth Wave (Pre-Ductal) | - |  | Double | Waveforms | IntelliVue X3 | Waves: PLTHpr
1465 | Pleth Wave (Post-Ductal) | - |  | Double | Waveforms | IntelliVue X3 | Waves: PLTHpo
1466 | Arterial Venous O2 Content Difference | - |  | Double | VitalSigns | IntelliVue X3 | Ca-vO2
1467 | O2 Delivery (DO2) | - |  | Double | VitalSigns | IntelliVue X3 | DO2
1469 | O2 Extraction Ratio (O2ER) | - |  | Double | VitalSigns | IntelliVue X3 | O2ER
1471 | O2 Delivery Index (DO2I) | - |  | Double | VitalSigns | IntelliVue X3 | DO2I
1472 | O2 Consumption Index (VO2I) | - |  | Double | VitalSigns | IntelliVue X3 | VO2I
1473 | Oxygenation Ratio (PaO2/FiO2) | - | PaO2 to FIO2 ratio expressed in mmHg to % ratio. | Double | VitalSigns | IntelliVue X3 | PaFIO2
1475 | Physiologic Dead Space | - |  | Double | VitalSigns | IntelliVue X3 | VENT: Vd
1477 | Alveolar Ventilation Volume (VA) | - |  | Double | VitalSigns | IntelliVue X3 | VENT: ALVENT
1478 | Urea Nitrogen Concentration (BUN) | - |  | Double | VitalSigns | IntelliVue X3 | BUN
1479 | Creatinine Concentration (Crea) | - | Measured by the i-Stat module. | Double | VitalSigns | IntelliVue X3 | Crea
1480 | Plasma Osmolarity (Osm) | - |  | Double | VitalSigns | IntelliVue X3 | PlOsm
1482 | Urine Creatinine | - |  | Double | VitalSigns | IntelliVue X3 | UrCrea
1484 | Urine Potassium (K+) | - |  | Double | VitalSigns | IntelliVue X3 | UrK
1486 | Urine Osmolarity | - |  | Double | VitalSigns | IntelliVue X3 | UrOsm
1487 | Blood Urea Nitrogen Creatinine Ratio (BUN/cr) | - |  | Double | VitalSigns | IntelliVue X3 | BUN/cr
1488 | Osmolar Clearance (COSM) | - |  | Double | VitalSigns | IntelliVue X3 | COsm
1489 | Creatinine Clearance (CreaCl) | - |  | Double | VitalSigns | IntelliVue X3 | CreaCl
1490 | Fractional Excretion of Sodium (FeNa) | - |  | Double | VitalSigns | IntelliVue X3 | FeNa
1493 | Urine Serum Creatinine Ratio | - |  | Double | VitalSigns | IntelliVue X3 | U/SCr
1495 | Urine Potassium Excretion | - |  | Double | VitalSigns | IntelliVue X3 | UrKEx
1496 | Urine Sodium Excretion | - |  | Double | VitalSigns | IntelliVue X3 | UrNaEx
1497 | Free Water Clearance (CH2O) | - |  | Double | VitalSigns | IntelliVue X3 | CH2O
1520 | Inspired N2O | - |  | Double | VitalSigns | IntelliVue X3 | N2O
1523 | Premature Ventricular Contractions | - |  | Double | VitalSigns | Common | ECG: PVC
1533 | LAP Systolic | - |  | Double | VitalSigns | Common | IBP: LAP
1534 | LAP Diastolic | - |  | Double | VitalSigns | Common | IBP: LAP
1540 | TcGas Heating Power | - |  | Double | DeviceInfo | IntelliVue X3 | TCUT: HeatPw
1556 | Abdominal Fetal Heart Rate (FHR) | - |  | Double | VitalSigns | IntelliVue X3 | aFHR
1557 | Abdominal Fetal Beat-to-Beat Rate (FHR) | - |  | Double | VitalSigns | IntelliVue X3 | baFHR
1558 | Patient Activity Level (Movement) | - |  | Double | VitalSigns | IntelliVue X3 | mAct
1559 | Patient Activity Level (Wearable) | - |  | Double | VitalSigns | IntelliVue X3 | wAct
1569 | Minimum Airway Pressure | - |  | Double | VitalSigns | IntelliVue X3 | AWPmin
1570 | P0.1 Occlusion Pressure | - |  | Double | VitalSigns | IntelliVue X3 | VENT: Poccl
1573 | Patient CO2 Production | - | Volume of CO2 produced by respiration. | Double | VitalSigns | IntelliVue X3 | VCO2
1574 | Relative Dead Space | - | Relative Dead Space Volume. | Double | VitalSigns | IntelliVue X3 | VENT: relVd
1590 | Procalcitonin (PCT) | - |  | Double | VitalSigns | IntelliVue X3 | PCT
1607 | Fetal Heart Rate (USound) | - | This variable is associated with ChannelID identifier. | Double | VitalSigns | IntelliVue X3 | FHR/FHR1/FHR2/FHR3
1609 | Fetal Beat-to-Beat Rate (USound) | - | This variable is associated with ChannelID identifier. | Double | VitalSigns | IntelliVue X3 | btbFHR/bFHR1/bFHR2/bFHR3
1630 | HFV Mean Pressure Setting | - |  | Double | Settings | IntelliVue X3 | sHFMAP
1631 | HFV Tidal Volume Setting | - |  | Double | Settings | IntelliVue X3 | sHFVTV
1645 | PAV-based Resistance Setting | - |  | Double | Settings | IntelliVue X3 | sRppv
1646 | PAV-based Elastance Setting | - |  | Double | Settings | IntelliVue X3 | sEppv
1703 | Abdominal TOCO | - |  | Double | VitalSigns | IntelliVue X3 | aToco
1704 | Chest Movement Respiration Rate | - |  | Double | VitalSigns | IntelliVue X3 | cmRR
1705 | Patient Movement Respiration Rate | - |  | Double | VitalSigns | IntelliVue X3 | mRR
1706 | Patient Movement Pulse Rate | - |  | Double | VitalSigns | IntelliVue X3 | mPulse
1707 | Respiration Signal Quality Index (Movement) | - |  | Double | OtherMeasuredData | IntelliVue X3 | mSQI
1710 | Cardiac Power Index (CCO) | - |  | Double | VitalSigns | IntelliVue X3 | CPI
1711 | Cardiac Power Output (CCO) | - |  | Double | VitalSigns | IntelliVue X3 | CPO
1715 | Central Venous Pressure (Thermodilution) | - |  | Double | VitalSigns | IntelliVue X3 | iCVP
1716 | Mean Arterial Pressure (Thermodilution) | - |  | Double | VitalSigns | IntelliVue X3 | iARTm
1718 | Average Blood Temperature (Thermodilution) | - |  | Double | VitalSigns | IntelliVue X3 | iAvgBT
1723 | O2 Delivery DO2 (Thermodilution) | - |  | Double | VitalSigns | IntelliVue X3 | iDO2
1724 | O2 Consumption (Thermodilution) | - |  | Double | VitalSigns | IntelliVue X3 | iVO2
1725 | O2 Consumption Index (Thermodilution) | - |  | Double | VitalSigns | IntelliVue X3 | iVO2I
1733 | Global End-Diastolic Volume (Thermodilution) | - |  | Double | VitalSigns | IntelliVue X3 | iGEDV
1734 | Global End-Diastolic Volum Index (Thermodilution) | - |  | Double | VitalSigns | IntelliVue X3 | iGEDVI
1735 | Pulmonary Vascular Permeability Index (Thermodilution) | - |  | Double | VitalSigns | IntelliVue X3 | iPVPI
1737 | Stroke Volume (Thermodilution) | - |  | Double | VitalSigns | IntelliVue X3 | iSV
1738 | Stroke Volume Index (Thermodilution) | - |  | Double | VitalSigns | IntelliVue X3 | iSI
1739 | O2 Consumption Estimate | - |  | Double | VitalSigns | IntelliVue X3 | iVO2e
1745 | EEG1 Wave | - |  | Double | Waveforms | IntelliVue X3 | Waves: EEG1
1746 | EEG2 Wave | - |  | Double | Waveforms | IntelliVue X3 | Waves: EEG2
1747 | EEG3 Wave | - |  | Double | Waveforms | IntelliVue X3 | Waves: EEG3
1748 | EEG4 Wave | - |  | Double | Waveforms | IntelliVue X3 | Waves: EEG4
1751 | Spectral Edge Frequency | - |  | Double | VitalSigns | IntelliVue X3 | EEG: SEF
1757 | Burst Suppression Ratio | - | This variable is associated with ChannelID identifier. | Double | VitalSigns | IntelliVue X3 | BSR1/BSR2/BSR3/BSR4
1759 | Patient Energy Expenditure | - |  | Double | VitalSigns | IntelliVue X3 | EE
1760 | Patient Respiratory Quotient | - |  | Double | VitalSigns | IntelliVue X3 | RQ
1782 | Spectral Edge Frequency (CH3) | - |  | Double | VitalSigns | IntelliVue X3 | EEG: Frequ3
1784 | Relative Power In Delta Band (CH3) | - |  | Double | VitalSigns | IntelliVue X3 | Delta3
1785 | Relative Power In Theta Band (CH3) | - |  | Double | VitalSigns | IntelliVue X3 | Theta3
1786 | Relative Power In Alpha Band (CH3) | - |  | Double | VitalSigns | IntelliVue X3 | Alpha3
1787 | Relative Power In Beta Band (CH3) | - |  | Double | VitalSigns | IntelliVue X3 | Beta3
1790 | Spectral Edge Frequency (CH4) | - |  | Double | VitalSigns | IntelliVue X3 | SEF4
1792 | Relative Power In Delta Band (CH4) | - |  | Double | VitalSigns | IntelliVue X3 | Delta4
1793 | Relative Power In Theta Band (CH4) | - |  | Double | VitalSigns | IntelliVue X3 | Theta4
1794 | Relative Power In Alpha (CH4) | - |  | Double | VitalSigns | IntelliVue X3 | Alpha4
1795 | Relative Power In Beta Band (CH4) | - |  | Double | VitalSigns | IntelliVue X3 | Beta4
1911 | Device Unique ID | unitless |  | String | DeviceInfo | Common | System Information
1914 | Patient Bed ID | unitless | This variable contains a printable string identifying the system location. | String | PatientInfo | Common | System Information
1929 | Patient ID | unitless |  | String | PatientInfo | Common | Patient Demographics
1930 | Patient Name | unitless |  | String | PatientInfo | Common | Patient Demographics
1942 | Patient Comment Text | unitless |  | String (Multiset) | PatientInfo | Common | Patient Demographics
1955 | Patient Stress Index | - |  | Double | PatientInfo | IntelliVue X3 | StrssI
1982 | Generic Patient Temperature | - |  | Double | VitalSigns | IntelliVue X3 | Tbody
2049 | Breath Cycle Type | - |  | Double | DeviceInfo | IntelliVue X3 | sCycTy
2077 | Airway Gas Flow Rate | - | Measured in the inspiratory path. | Double | VitalSigns | IntelliVue X3 | VENT: AWFin
2092 | Inspired Minute Volume | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
2104 | Positive End Expiratory Pressure (PEEP) Setting | - |  | Double | Settings | IntelliVue X3 | sEPAP
2163 | Spectral Edge Frequency - 1st EEG | - |  | Double | VitalSigns | IntelliVue X3 | EEG: Frequ1
2167 | Bispectral Index | - | This variable is associated with ChannelID identifier. | Double | VitalSigns | IntelliVue X3 | BIS L/BIS R
2168 | Total Power (Ch1) | - |  | Double | VitalSigns | IntelliVue X3 | EEG: TP1
2169 | Spike Rate | - | This variable is associated with ChannelID identifier. | Double | VitalSigns | IntelliVue X3 | BurstL/BurstR
2173 | Spectral Edge Frequency - 2nd EEG | - |  | Double | VitalSigns | IntelliVue X3 | EEG: Frequ2
2178 | Total Power (Ch2) | - |  | Double | VitalSigns | IntelliVue X3 | EEG: TP2
2182 | Narcotrend Index | - |  | Double | VitalSigns | IntelliVue X3 | NI
2183 | Sharp Transient Intensity | - |  | Double | VitalSigns | IntelliVue X3 | STI
2188 | Total Power Relative | - |  | Double | VitalSigns | IntelliVue X3 | TPrel
2200 | Device Hardware Revision | unitless | This variable is obsolete because it may contains incorrect information. Please use variables 764, 765 and 4201 must be used instead. | String | DeviceInfo | Common | Obsolete Variables
2201 | Serial Number | unitless | This variable is obsolete because it may contains incorrect information. Please use variables 764, 765 and 4201 must be used instead. | String | DeviceInfo | Common | Obsolete Variables
2246 | O2 Flush Percentage Setting | - |  | Double | Settings | IntelliVue X3 | sO2max
2250 | CO2 Tidal Production | - |  | Double | VitalSigns | IntelliVue X3 | VCO2ti
2254 | Inspiratory Resistance | - |  | Double | VitalSigns | IntelliVue X3 | Rinsp
2255 | Device Date/Time | unitless | Format: 16 ascii digits string of "YYYYMMDDhhmmssff", where the "ff" is second_fractions that is not used by IntelliVue monitors. | String | DeviceInfo | Common | System Information
2259 | Arterial BP Systolic | - |  | Double | VitalSigns | Common | IBP: ABP
2260 | Arterial BP Diastolic | - |  | Double | VitalSigns | Common | IBP: ABP
2261 | Arterial BP Mean | - |  | Double | VitalSigns | Common | IBP: ABP
2262 | Pulmonary Artery BP - Systolic | - |  | Double | VitalSigns | Common | IBP: PAP
2263 | Pulmonary Artery BP - Diastolic | - |  | Double | VitalSigns | Common | IBP: PAP
2264 | Pulmonary Artery BP Mean | - |  | Double | VitalSigns | Common | IBP: PAP
2271 | Right Atrial Pressure (RAP) | - | Available on IntelliVue devices with firmware revision G or later. | Double | VitalSigns | Common | IBP: RAP
2272 | Left Atrial Pressure (LAP) | - | Available on IntelliVue devices with firmware revision G or later. | Double | VitalSigns | Common | IBP: LAP
2273 | Intra Cranial Pressure (ICP) | - | Available on IntelliVue devices with firmware revision G or later. | Double | VitalSigns | Common | IBP: ICP
2317 | Pulmonary Artery Wedge Pressure | - |  | Double | VitalSigns | IntelliVue X3 | IBP: PAWP
2321 | Generic NIBP Systolic Pressure | - | This variable is same as variable 2, except that the early estimated values are marked as NonSignificant in this variable. | Double | VitalSigns | Common | NBP
2322 | Generic NIBP Diastolic Pressure | - | This variable is same as variable 3, except that the early estimated values are marked as NonSignificant in this variable. | Double | VitalSigns | Common | NBP
2323 | Generic NIBP Mean Pressure | - | This variable is same as variable 4, except that the early estimated values are marked as NonSignificant in this variable. | Double | VitalSigns | Common | NBP
2324 | Inspired Tidal Volume | - |  | Double | VitalSigns | IntelliVue X3 | TVin
2325 | Expired Tidal Volume | - |  | Double | VitalSigns | IntelliVue X3 | TVexp
2326 | Expired Minute Volume | - |  | Double | VitalSigns | IntelliVue X3 | MINVOL
2341 | Patient Blood Temperature | - |  | Double | VitalSigns | IntelliVue X3 | Tblood
2342 | Injectate Temperature | - |  | Double | OtherMeasuredData | IntelliVue X3 | Tinj
2343 | Inspired O2 | - |  | Double | VitalSigns | IntelliVue X3 | VENT: FIO2
2350 | Inspired O2 Pressure | - | This variable is marked as obsolete because the label is incorrect. Please see variable 52523 for concentration and 52527 for partial pressure. | Double | VitalSigns | IntelliVue X3 | O2
2351 | End Tidal O2 | - |  | Double | VitalSigns | IntelliVue X3 | O2
2352 | CO2 (Gas Analyzer) | - |  | Double | VitalSigns | IntelliVue X3 | CO2
2357 | End Tidal N2O | - |  | Double | VitalSigns | IntelliVue X3 | N2O
2360 | End Tidal Agent Concentration | - |  | Double | VitalSigns | IntelliVue X3 | AGT
2362 | End Tidal Desflurane | - |  | Double | VitalSigns | IntelliVue X3 | DES
2364 | End Tidal Sevoflurane | - |  | Double | VitalSigns | IntelliVue X3 | SEV
2366 | End Tidal Halothane | - |  | Double | VitalSigns | IntelliVue X3 | HAL
2368 | End Tidal Enflurane | - |  | Double | VitalSigns | IntelliVue X3 | ENF
2370 | End Tidal Isoflurane | - |  | Double | VitalSigns | IntelliVue X3 | ISO
2378 | Fresh Gas N2O Flow | - | Mixer parameter. Setting of fresh gas N2O flow. | Double | DeviceInfo | IntelliVue X3 | sfgN2O
2379 | Fresh Gas Air Flow | - |  | Double | DeviceInfo | IntelliVue X3 | fgAir
2391 | Transcutaneous O2 Pressure | - |  | Double | VitalSigns | IntelliVue X3 | TCUT: tcpO2
2392 | Transcutaneous CO2 Pressure | - |  | Double | VitalSigns | IntelliVue X3 | TCUT: tcpCO2
2399 | Arterial O2 Saturation (SaO2) | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
2408 | EEG Total Power | - |  | Double | VitalSigns | IntelliVue X3 | EEG: TP
2409 | Bispectral Index (BIS) | - |  | Double | VitalSigns | IntelliVue X3 | EEG: BIS
2411 | EEG Burst Suppression Ratio | - |  | Double | VitalSigns | IntelliVue X3 | EEG: BSI
2420 | Spontaneous Alveolar Minute Volume | - |  | Double | VitalSigns | IntelliVue X3 | SpMVal
2421 | Airway Dead Space | - |  | Double | VitalSigns | IntelliVue X3 | awVd
2432 | Expired Spontaneous Tidal Volume | - |  | Double | VitalSigns | IntelliVue X3 | SpTVex
2434 | Spontaneous Expired Minute Volume | - |  | Double | VitalSigns | IntelliVue X3 | %SpMV
2440 | Total Alveolar Tidal Volume | - |  | Double | VitalSigns | IntelliVue X3 | TValv
2447 | Respiration Waveform | - |  | Double | Waveforms | IntelliVue X3 | Waves: cmResp
2456 | EEG Signal Quality Index | - |  | Double | VitalSigns | IntelliVue X3 | EEG: SQI
2517 | Wearable Respiration Rate | - |  | Double | VitalSigns | IntelliVue X3 | wRR
2518 | Wearable Pulse Rate | - |  | Double | VitalSigns | IntelliVue X3 | wPulse
2519 | Wearable Patient Temperature | - |  | Double | VitalSigns | IntelliVue X3 | wTemp
2520 | Respiration Signal Quality Index (Wearable) | - |  | Double | OtherMeasuredData | IntelliVue X3 | wSQI
2521 | Pulse Rate Signal Quality Index (Wearable) | - |  | Double | OtherMeasuredData | IntelliVue X3 | wSQI
2542 | Airway Pressure Wave | - |  | Double | Waveforms | IntelliVue X3 | Waves: AWP
2543 | Citrate Blood Concentration Setting | - |  | Double | Settings | IntelliVue X3 | sCitDo
2547 | Neutrophil Count | - |  | Double | VitalSigns | IntelliVue X3 | Bands
2580 | Fresh Gas Total Flow Setting | - | Mixer parameter. | Double | Settings | IntelliVue X3 | sfgFl
2583 | I:E Ratio Setting | unitless | This variable is associated with ChannelID identifier. This variable is calculated from variable 2773, which is sent as a string in the format "1:xx.xx". | String | Settings | IntelliVue X3 | Calculated Variables
2598 | Patient Admission Status | unitless | Possible values: "empty", "pre-admitted", "admitted", and "discharged". | String | PatientInfo | Common | Patient Demographics
2602 | ECG Lead V | - | Physiological identifier: NOM_ECG_AMPL_ST_V | Double | VitalSigns | Common | ECG: ST
2738 | Plateau Pressure | - | This variable contains parameter from Philips label "NLS_NOM_VENT_PRESS_RESP_PLAT" and observed value "NLS_NOM_VENT_PRESS_RESP_PLAT". | Double | VitalSigns | IntelliVue X3 | VENT: PEinsp
2773 | I:E Ratio Setting | - |  | Double | Settings | IntelliVue X3 | sIE 1:
2776 | Total PEEP | - |  | Double | VitalSigns | IntelliVue X3 | VENT: tPEEP
2787 | APRV Low Time (Tlow) | - |  | Double | VitalSigns | IntelliVue X3 | Tlow
2821 | System General Status | unitless | This variable identifies the current operating mode of the device. Possible values: "not specified", "monitoring", "demo", and "service". Moreover, "- standby" or "- config" can be appended | String | DeviceInfo | Common | System Information
2845 | Frequency | Hz | This variable describes the frequency of the main power supply. Possible values: "unspecified", "50", and "60". | String | DeviceInfo | Common | System Information
2855 | Device Status | unitless | Possible values: "disconnected", "unassociated", "operating". | String | DeviceInfo | Common | System Information
2879 | Device Software Revision | unitless | This variable is obsolete because it may contains incorrect information. Please use variables 764, 765 and 4201 must be used instead. | String | DeviceInfo | Common | Obsolete Variables
2880 | Maximum Apnea Time Setting | - |  | Double | DeviceInfo | IntelliVue X3 | sTAmax
2889 | Inspired Agent Concentration | - |  | Double | VitalSigns | IntelliVue X3 | inAGT
2898 | Patient Weight | - |  | Double | PatientInfo | IntelliVue X3 | Weight
2901 | Patient First Name | unitless |  | String | PatientInfo | Common | Patient Demographics
2907 | Patient Temperature | - |  | Double | VitalSigns | Common | Temp
2915 | Ejection Fraction | - |  | Double | VitalSigns | IntelliVue X3 | REF
2916 | End Diastolic Volume | - |  | Double | VitalSigns | IntelliVue X3 | EDV
2917 | End Diastolic Volume Index | - |  | Double | VitalSigns | IntelliVue X3 | EDVI
2918 | End Systolic Volume | - |  | Double | VitalSigns | IntelliVue X3 | ESV
2919 | End Systolic Volume Index | - |  | Double | VitalSigns | IntelliVue X3 | ESVI
2922 | EEG Electromyography (EMG) | - | This variable is associated with ChannelID identifier. | Double | VitalSigns | IntelliVue X3 | EMG L/EMG R
2923 | Generic EEG Wave | - |  | Double | Waveforms | IntelliVue X3 | Waves: EEG
2924 | ECG Pad Wave | - |  | Double | Waveforms | IntelliVue X3 | Waves: Pads
2925 | ECG Paddle Wave | - |  | Double | Waveforms | IntelliVue X3 | Waves: Paddle
2930 | Unspecific Pressure Pulse Rate | - |  | Double | VitalSigns | IntelliVue X3 | Pulse
2932 | Patient Temperature 1 | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
2933 | Patient Temperature 2 | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
2954 | APRV High Pressure (Phigh) | - |  | Double | VitalSigns | IntelliVue X3 | Phigh
2959 | APRV Low Pressure (Plow) | - |  | Double | VitalSigns | IntelliVue X3 | Plow
2964 | I:E Ratio | unitless | This variable is associated with ChannelID identifier This variable is calculated from variable 1190 and sent as a string in the format "1:xx.xx". | String | VitalSigns | IntelliVue X3 | Calculated Variables
2968 | Plateau Airway Pressure | - |  | Double | VitalSigns | IntelliVue X3 | Pplat
2984 | Plateau Pressure Setting | - |  | Double | Settings | IntelliVue X3 | sPplat
2991 | FiO2 Upper Alarm Limit | - |  | Double | Settings | IntelliVue X3 | AlarmLimit: highO2
2992 | FiO2 Lower Alarm Limit | - |  | Double | Settings | IntelliVue X3 | AlarmLimit: lowO2
2995 | Device Altitude | m |  | Double | DeviceInfo | Common | System Information
3003 | Blood Flow Rate Setting | - |  | Double | Settings | IntelliVue X3 | sBFl
3005 | Dialysate Flow Rate Setting | - |  | Double | Settings | IntelliVue X3 | sDialF
3007 | Effluent Flow Rate Setting | - |  | Double | Settings | IntelliVue X3 | sEfflF
3024 | Access Pod Pressure | - |  | Double | OtherMeasuredData | IntelliVue X3 | AccP
3025 | Return Pod Pressure | - |  | Double | OtherMeasuredData | IntelliVue X3 | RetP
3026 | Filter Pod Pressure | - |  | Double | OtherMeasuredData | IntelliVue X3 | FiltP
3027 | Effluent Pod Pressure | - |  | Double | OtherMeasuredData | IntelliVue X3 | EfflP
3029 | Transmembrane Pressure | - |  | Double | VitalSigns | IntelliVue X3 | TMP
3062 | Hb Calculation Offset | - |  | Double | VitalSigns | IntelliVue X3 | OffsHb
3091 | Patient Time Since Last Ward Admission | - |  | Double | PatientInfo | IntelliVue X3 | TWrdAd
3092 | Patient Total Admission Time | - |  | Double | PatientInfo | IntelliVue X3 | WardAd
3101 | Patient Drainage Volume | - |  | Double | VitalSigns | IntelliVue X3 | DrnVol
3168 | Fluid Removal Actual Rate | - |  | Double | VitalSigns | IntelliVue X3 | Vrem/h
3204 | Sum of Minimum Alveolar Concentration (MAC) | - |  | Double | VitalSigns | IntelliVue X3 | MAC
3266 | HFV Frequency Setting | - |  | Double | Settings | IntelliVue X3 | sHFFrq
3268 | Ventilator Base Flow Setting | - |  | Double | Settings | IntelliVue X3 | sfgFmn
3271 | Blood Pressure Difference | - |  | Double | VitalSigns | IntelliVue X3 | IBP: DeltaP
3272 | Generic Blood Pressure Difference 1 | - |  | Double | VitalSigns | IntelliVue X3 | IBP: DeltaP1
3273 | Generic Blood Pressure Difference 2 | - |  | Double | VitalSigns | IntelliVue X3 | IBP: DeltaP2
3278 | Patient Height | - |  | Double | PatientInfo | IntelliVue X3 | Height
3291 | Patient Temperature Setting | - |  | Double | Settings | IntelliVue X3 | sTskin
3385 | Arterial O2 Content (CaO2) | - |  | Double | VitalSigns | IntelliVue X3 | CaO2
3391 | Alveolar Arterial O2 Difference (AaDO2) | - |  | Double | VitalSigns | IntelliVue X3 | AaDO2
3400 | Venous O2 Content (CvO2) | - |  | Double | VitalSigns | IntelliVue X3 | CvO2
3415 | Ionized Calcium Concentration (Ca) | - |  | Double | VitalSigns | IntelliVue X3 | iCa
3421 | Chloride Concentration (Cl-) | - |  | Double | VitalSigns | IntelliVue X3 | Cl
3435 | Activated Clotting Time | - | Measured by the i-Stat module. | Double | VitalSigns | IntelliVue X3 | ACT
3446 | Generic Temperature | - |  | Double | VitalSigns | IntelliVue X3 | T1
3447 | Generic Temperature 2 | - |  | Double | VitalSigns | IntelliVue X3 | T2
3448 | Generic Temperature 3 | - |  | Double | VitalSigns | IntelliVue X3 | T3
3449 | Generic Temperature 4 | - |  | Double | VitalSigns | IntelliVue X3 | T4
3474 | Expiratory Resistance | - |  | Double | VitalSigns | IntelliVue X3 | Rexp
3490 | TcGas Site Timer | - |  | Double | OtherMeasuredData | IntelliVue X3 | TCUT: SitTim
3491 | O2 Saturation (SO2) | - |  | Double | VitalSigns | IntelliVue X3 | SaO2
3496 | TcGas Transducer Temperature | - |  | Double | DeviceInfo | IntelliVue X3 | TCUT: SensrT
3502 | Peak Expiratory Flow Setting | - |  | Double | Settings | IntelliVue X3 | s%PEF
3503 | Auto Flow Active Status | - |  | Double | DeviceInfo | IntelliVue X3 | sFlCtr
3506 | Inspiratory Minimum CO2 | - |  | Double | VitalSigns | Common | CO2
3507 | Expired Tidal Volume Indexed by Body Weight | - |  | Double | VitalSigns | IntelliVue X3 | TVe/kg
3555 | ATC Activation Status | - |  | Double | DeviceInfo | IntelliVue X3 | sTubeC
3570 | Skin Temperature | - |  | Double | VitalSigns | Common | Tskin
3571 | Rectal Temperature | - |  | Double | VitalSigns | Common | Trect
3572 | Nasopharyngeal Temperature | - |  | Double | VitalSigns | Common | Tnaso
3573 | Esophageal Temperature | - |  | Double | VitalSigns | Common | Tesoph
3574 | Core Temperature | - |  | Double | VitalSigns | Common | Tcore
3575 | Arterial Temperature | - |  | Double | VitalSigns | Common | Tart
3576 | Venous Temperature | - |  | Double | VitalSigns | Common | Tven
3583 | Generic Core Temperature 1 | - |  | Double | VitalSigns | IntelliVue X3 | T1Core
3584 | Generic Core Temperature 2 | - |  | Double | VitalSigns | IntelliVue X3 | T2Core
3591 | Device Number | unitless | Device Model Number (example: "M8000") | String | DeviceInfo | Common | System Information
3635 | Drug Delivery Rate | - |  | Double | DeviceInfo | IntelliVue X3 | sDlvRt
3637 | Drug Concentration | - |  | Double | PatientInfo | IntelliVue X3 | sDConc
3639 | Manufacturers Name | unitless | "Philips" | String | DeviceInfo | Common | System Information
3640 | Pressure Alarm Delay | - |  | Double | DeviceInfo | IntelliVue X3 | sPADel
3646 | Arterial Blood Pressure (ART) | - | Available on IntelliVue devices with firmware revision G or later. Arterial Blood Pressure (NOM_PRESS_BLD_ART). | Double | VitalSigns | Common | IBP: ART
3647 | Pulmonary Artery Blood Pressure (PA) | - | Available on IntelliVue devices with firmware revision G or later. | Double | VitalSigns | Common | IBP: PAP
3704 | Arterial/Main Pump Flow | - |  | Double | VitalSigns | IntelliVue X3 | HLMfl
3710 | Slave Pump Flow | - |  | Double | VitalSigns | IntelliVue X3 | SlvPfl
3715 | Suction Pump Flow | - |  | Double | VitalSigns | IntelliVue X3 | SucPfl
3720 | Auxiliary Pump Flow | - |  | Double | VitalSigns | IntelliVue X3 | AuxPfl
3723 | Auxiliary Pump Time Since Start of Current Plegia Session | - |  | Double | VitalSigns | IntelliVue X3 | AxOnTi
3724 | Auxiliary Pump Time Since Stop of Plegia Session | - |  | Double | VitalSigns | IntelliVue X3 | AxOffT
3725 | Auxiliary Pump Delivered Volume in Current Plegia Session | - |  | Double | VitalSigns | IntelliVue X3 | AxDVol
3726 | Auxiliary Pump Total Volume | - |  | Double | VitalSigns | IntelliVue X3 | AxTVol
3727 | Auxiliary Pump Total Plegia Time | - |  | Double | VitalSigns | IntelliVue X3 | AxPlTi
3730 | Cardioplegia Main Pump Flow | - |  | Double | VitalSigns | IntelliVue X3 | PlePfl
3733 | Cardioplegia Main Pump Delivered Volume of Current Plegia Session | - |  | Double | VitalSigns | IntelliVue X3 | CpDVol
3734 | Cardioplegia Main Pump Total Volume | - |  | Double | VitalSigns | IntelliVue X3 | CpTVol
3737 | Cardioplegia Slave Pump Flow | - |  | Double | VitalSigns | IntelliVue X3 | SplPfl
3740 | Cardioplegia Slave Pump Delivered Volume in Current Plegia Session | - |  | Double | VitalSigns | IntelliVue X3 | CsDVol
3741 | Cardioplegia Slave Pump Total Volume | - |  | Double | VitalSigns | IntelliVue X3 | CsTVol
3742 | Cycle Flow Sensitivity Setting | - |  | Double | Settings | IntelliVue X3 | sFlSeC
3761 | Hydrocarbon Plasma Concentration (HCO3) | - |  | Double | VitalSigns | IntelliVue X3 | HCO3
3781 | Corpuscular Mean Hemoglobin Concentration | - |  | Double | VitalSigns | IntelliVue X3 | MCHC
3785 | Arterial O2 Partial Pressure | - |  | Double | VitalSigns | IntelliVue X3 | PaO2
3786 | Venous O2 Partial Pressure | - |  | Double | VitalSigns | IntelliVue X3 | PvO2
3787 | Arterial CO2 Partial Pressure | - |  | Double | VitalSigns | IntelliVue X3 | PaCO2
3789 | Arterial O2 Saturation (SaO2) | - | Obsolete variable that may be removed in a future release. The label of this variable is not correct. Please use variable 14 instead. | Double | VitalSigns | Common | Obsolete Variables
3793 | Hematocrit (HCT) | - | Post-Dialysis Hematocrit | Double | VitalSigns | IntelliVue X3 | cHCTo
3804 | Respiratory Rate (Flow) | - |  | Double | VitalSigns | IntelliVue X3 | RRawf
3805 | Respiratory Rate (CO2) | - |  | Double | VitalSigns | IntelliVue X3 | RRCO2
3853 | Patient Type | unitless | Possible values: "unspecified", "adult", "pediatric", and "neonatal". | String | PatientInfo | Common | Patient Demographics
3854 | Patient Pace Maker Status | unitless | Possible values: "not paced", "paced gen". | String | PatientInfo | Common | Patient Demographics
3869 | Heart Rate Signal Quality | - |  | Double | OtherMeasuredData | IntelliVue X3 | mSQI
3917 | Tracheal Pressure | - |  | Double | VitalSigns | IntelliVue X3 | Ptrach
3921 | Ventilator Leakage | - |  | Double | VitalSigns | IntelliVue X3 | Lksys
3961 | Plateau Time Setting | - | This variable is obsolete that may be removed in a future release. The label of this variable is incorrect, please use variable 9585 instead. | Double | Settings | IntelliVue X3 | Obsolete Variables
3963 | Warmer Mattress Temperature | - |  | Double | OtherMeasuredData | IntelliVue X3 | Tmatt
3968 | Spontaneous Fraction of Minute Volume | - |  | Double | VitalSigns | IntelliVue X3 | SpMVe
4004 | Breathing System Leakage | - |  | Double | OtherMeasuredData | IntelliVue X3 | VENT: Leak
4006 | Inspired MAC Concentration | - |  | Double | VitalSigns | IntelliVue X3 | MAC
4008 | Inspiratory 2nd Agent | - |  | Double | VitalSigns | IntelliVue X3 | inAGTs
4009 | Expiratory 2nd Agent | - | This is the device's measurement with label="NLS_NOM_EMFC_etAGTs" and physio_id="NOM_CONC_AWAY_AGENT_ET". | Double | VitalSigns | IntelliVue X3 | etAGTs
4012 | Volume Assist Level Setting | - |  | Double | Settings | IntelliVue X3 | sWOBva
4014 | Backup Tidal Volume Setting | - |  | Double | Settings | IntelliVue X3 | sTVbak
4016 | Flow Acceleration Setting | - |  | Double | Settings | IntelliVue X3 | sFlacc
4017 | Cumulated Halothane Consumption | - |  | Double | OtherMeasuredData | IntelliVue X3 | AccHAL
4018 | Cumulated Enflurane Consumption | - |  | Double | OtherMeasuredData | IntelliVue X3 | AccENF
4019 | Cumulated Isoflurane Consumption | - |  | Double | OtherMeasuredData | IntelliVue X3 | AccISO
4020 | Cumulated Desflurane Consumption | - |  | Double | OtherMeasuredData | IntelliVue X3 | AccDes
4021 | Cumulated Sevoflurane Consumption | - |  | Double | OtherMeasuredData | IntelliVue X3 | AccSEV
4026 | Air Humidity Setting | - |  | Double | Settings | IntelliVue X3 | sHum
4028 | Mattress Temperature Setting | - |  | Double | Settings | IntelliVue X3 | sTmatt
4029 | Inspired Agent Setting | - |  | Double | Settings | IntelliVue X3 | sInAGT
4030 | Expired Agent Setting | - | End-Tidal agent concentration | Double | Settings | IntelliVue X3 | sEtAGT
4037 | Inspired Desflurane Setting | - |  | Double | Settings | IntelliVue X3 | sInDES
4038 | Expired Desflurane Setting | - | End-Tidal desflurane concentration | Double | Settings | IntelliVue X3 | sEtDES
4039 | Inspired Sevoflurane Setting | - |  | Double | Settings | IntelliVue X3 | sInSEV
4040 | Expired Sevoflurane Setting | - | End-Tidal sevoflurane concentration | Double | Settings | IntelliVue X3 | sEtSEV
4079 | Fresh Gas Sevoflurane | - |  | Double | VitalSigns | IntelliVue X3 | fgSEV
4080 | Fresh Gas Desflurane | - |  | Double | VitalSigns | IntelliVue X3 | fgDES
4081 | Ambient Temperature | - |  | Double | OtherMeasuredData | IntelliVue X3 | Tamb
4082 | Ambient Pressure | - |  | Double | OtherMeasuredData | IntelliVue X3 | PB
4084 | Pulse Rate ( Unspecified Source ) | - |  | Double | VitalSigns | Common | Pulse
4089 | Central Venous Pressure Diastolic | - |  | Double | VitalSigns | Common | IBP: CVP
4093 | Central Venous Mean Pressure | - |  | Double | VitalSigns | Common | IBP: CVP
4094 | Central Venous Pressure Systolic | - |  | Double | VitalSigns | Common | IBP: CVP
4120 | Expiratory Volume Wave | - |  | Double | Waveforms | IntelliVue X3 | Waves: AWVex
4143 | Apnea Ventilation I:E Ratio Setting (Number) | - |  | Double | Settings | IntelliVue X3 | sBIE1:
4145 | Dynamic Plateau Airway Pressure | - |  | Double | VitalSigns | IntelliVue X3 | PplDyn
4158 | Transcutaneous Relative Heater Power | - |  | Double | OtherMeasuredData | IntelliVue X3 | HeatPr
4171 | Spontaneous Respiratory Rate Setting | - |  | Double | Settings | IntelliVue X3 | sRRmax
4201 | Product Description | unitless | Variables 764, 765 and 4201 must be interpreted conjointly. | String (Multiset) | DeviceInfo | Common | System Information
Consult detailed help for more information on this variable
4226 | Replacement Cumulated Volume | - |  | Double | VitalSigns | IntelliVue X3 | Vrep/h
4227 | Dialysate Cumulated Volume | - | This variable is associated with ChannelID identifier. | Double | VitalSigns | IntelliVue X3 | DialC/Vdly/h
4228 | Effluent Cumulated Volume | - | This variable is associated with ChannelID identifier. | Double | VitalSigns | IntelliVue X3 | Veff/h / EfflC
4229 | Fluid Removal Cumulated Volume | - |  | Double | VitalSigns | IntelliVue X3 | Vufr
4352 | Base Pressure | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
4369 | PEEP/CPAP Setting | - |  | Double | Settings | IntelliVue X3 | VENT: sPEEP
4370 | Sigh Rate Setting | - |  | Double | Settings | IntelliVue X3 | VENT: sSghR
4371 | Sigh Tidal Volume Setting | - |  | Double | Settings | IntelliVue X3 | VENT: sSghTV
4375 | Data Averaging Interval Setting | - |  | Double | Settings | IntelliVue X3 | Available only in revF or earlier
4376 | Chart Interval Setting | - |  | Double | Settings | IntelliVue X3 | Available only in revF or earlier
4383 | Static Lung Compliance | - |  | Double | VitalSigns | IntelliVue X3 | Cstat
4384 | Static Lung Resistance | - |  | Double | VitalSigns | IntelliVue X3 | Raw
4385 | Dynamic Lung Compliance | - |  | Double | VitalSigns | IntelliVue X3 | Cdyn
4386 | Dynamic Lung Resistance | - | NLS_NOM_RES_AWAY_DYN | Double | VitalSigns | IntelliVue X3 | Rdyn
4388 | Apnea Inspiratory Pressure Setting | - |  | Double | Settings | IntelliVue X3 | sIPApn/sAInP
4393 | Patient Hose Compliance | - |  | Double | OtherMeasuredData | IntelliVue X3 | Chos
4401 | Oximeter Averaging Interval Setting | - |  | Double | Settings | IntelliVue X3 | Available only in revF or earlier
4457 | System Alarm Status | unitless | Overall alert state of the system. | Double | Alarms | Common | Alarms
Consult detailed help for more information on this variable
4458 | Patient Alarm Priority | unitless | Overall Patient alarm priority. This is the current highest priority value of currently active Patient Alarms. | Double | Alarms | Common | Alarms
Consult detailed help for more information on this variable
4459 | Technical Alarm Priority | unitless | Overall Technical alarm priority. This is the current highest priority value of currently active Technical Alarms. | Double | Alarms | Common | Alarms
Consult detailed help for more information on this variable
4460 | Audio Alarm Priority | unitless | Overall audible alarm priority. This is the current highest priority value of currently active audible alarms. | Double | Alarms | Common | Alarms
Consult detailed help for more information on this variable
4461 | Technical Alarm Source ID | unitless |  | Double (Multiset) | Alarms | Common | Alarms
4462 | Technical Alarm Global State | unitless | Consult detailed help for more information on this variable | Double (Multiset) | Alarms | Common | Alarms
4463 | Technical Alarm Priority | unitless | Consult detailed help for more information on this variable | Double (Multiset) | Alarms | Common | Alarms
4464 | Technical Alarm Flag | unitless | Consult detailed help for more information on this variable | Double (Multiset) | Alarms | Common | Alarms
4465 | Technical Alarm Priority | unitless |  | Double (Multiset) | Alarms | Common | Alarms
4466 | Technical Alarm Description | unitless |  | String (Multiset) | Alarms | Common | Alarms
4467 | Patient Alarm Source ID | unitless |  | Double (Multiset) | Alarms | Common | Alarms
4468 | Patient Alarm Global State | unitless | Consult detailed help for more information on this variable | Double (Multiset) | Alarms | Common | Alarms
4469 | Patient Alarm Priority | unitless | Consult detailed help for more information on this variable | Double (Multiset) | Alarms | Common | Alarms
4470 | Patient Alarm Flag | unitless | Consult detailed help for more information on this variable | Double (Multiset) | Alarms | Common | Alarms
4471 | Patient Alarm Priority | unitless |  | Double (Multiset) | Alarms | Common | Alarms
4472 | Patient Alarm Description | unitless |  | String (Multiset) | Alarms | Common | Alarms
4473 | Device Nomenclature Revision | unitless | Format: major.minor (such as 1.0) | String | DeviceInfo | Common | System Information
4474 | Device Text Catalog Revision | unitless | Format: "MM.NN mm.nn", where "MM.NN" is text catalog revision and "mm.nn" is language revision. | String | DeviceInfo | Common | System Information
4475 | Device Data Format | unitless | Always 11 (Unicode). | String | DeviceInfo | Common | System Information
4476 | Device Operating Area | unitless | Possible values: "unspecified", "operating room", "intensive care unit", "neonatal intensive care unit", and "cardiology care unit". | String | DeviceInfo | Common | System Information
4496 | Patient O2 Consumption (VO2) | - |  | Double | VitalSigns | IntelliVue X3 | VO2
4497 | Patient CO2 Production (VCO2) | - |  | Double | VitalSigns | IntelliVue X3 | bbvCO2
4507 | Arterial O2 Saturation (Left) | - |  | Double | VitalSigns | IntelliVue X3 | SpO2 l
4508 | Arterial O2 Saturation (Right) | - |  | Double | VitalSigns | IntelliVue X3 | SpO2 r
4509 | SpO2 Difference | - |  | Double | VitalSigns | IntelliVue X3 | DeltaSpO2
4510 | Perfusion Indicator | - |  | Double | VitalSigns | IntelliVue X3 | Perf
4511 | Relative Perfusion (Left) | - |  | Double | VitalSigns | IntelliVue X3 | Perf l
4512 | Relative Perfusion (Right) | - |  | Double | VitalSigns | IntelliVue X3 | Perf r
4513 | Ao Systolic | - |  | Double | VitalSigns | Common | IBP: Ao
4514 | Ao Diastolic | - |  | Double | VitalSigns | Common | IBP: Ao
4515 | Ao Mean | - |  | Double | VitalSigns | Common | IBP: Ao
4516 | RAP Systolic | - |  | Double | VitalSigns | Common | IBP: RAP
4517 | RAP Diastolic | - |  | Double | VitalSigns | Common | IBP: RAP
4518 | RAP Mean | - |  | Double | VitalSigns | Common | IBP: RAP
4519 | Intracranial Pressure ICP Systolic | - |  | Double | VitalSigns | Common | IBP: ICP
4520 | Intracranial Pressure ICP Diastolic | - |  | Double | VitalSigns | Common | IBP: ICP
4521 | ICP Mean | - |  | Double | VitalSigns | Common | IBP: ICP
4522 | UAP Systolic | - |  | Double | VitalSigns | Common | IBP: UAP
4523 | UAP Diastolic | - |  | Double | VitalSigns | Common | IBP: UAP
4524 | UAP Mean | - |  | Double | VitalSigns | Common | IBP: UAP
4525 | UVP Systolic | - |  | Double | VitalSigns | Common | IBP: UVP
4526 | UVP Diastolic | - |  | Double | VitalSigns | Common | IBP: UVP
4527 | UVP Mean | - |  | Double | VitalSigns | Common | IBP: UVP
4528 | Unspecific Pressure Systolic | - |  | Double | VitalSigns | Common | IBP: P
4529 | Unspecific Pressure Diastolic | - |  | Double | VitalSigns | Common | IBP: P
4530 | Unspecific Pressure Mean | - |  | Double | VitalSigns | Common | IBP: P
4531 | Cerebral Perfusion Pressure (CPP) | - |  | Double | VitalSigns | IntelliVue X3 | IBP: CPP
4532 | Stroke Volume Variation | - |  | Double | VitalSigns | IntelliVue X3 | SVV
4533 | Index of Left Ventricular Contractility | - | This variable contains measurement from label "NLS_NOM_GRAD_PRESS_BLD_AORT_POS_MAX" and observed value "NOM_GRAD_PRESS_BLD_AORT_POS_MAX". | Double | VitalSigns | IntelliVue X3 | dPmax
4556 | Aortic Blood Pressure (Ao) Wave | - |  | Double | Waveforms | IntelliVue X3 | Waves: Ao
4560 | Auxiliary Waveform | - |  | Double | Waveforms | IntelliVue X3 | Waves: AUX
4561 | Intermittent Mandatory Ventilation Status | - |  | Double | VitalSigns | IntelliVue X3 | VENT: IMV
4562 | Left Cardiac Work | - |  | Double | VitalSigns | IntelliVue X3 | LCW
4564 | Right Cardiac Work | - |  | Double | VitalSigns | IntelliVue X3 | RCW
4565 | Serum Creatinine | - |  | Double | VitalSigns | IntelliVue X3 | SCrea
4566 | Urine Plasma Osmolarity Ratio | - |  | Double | VitalSigns | IntelliVue X3 | U/POsm
4567 | Urine Sodium Potassium Ratio | - |  | Double | VitalSigns | IntelliVue X3 | UrNa/K
4568 | Urine Volume | - |  | Double | VitalSigns | IntelliVue X3 | UrVol
4574 | Left Cardiac Work Index | - |  | Double | VitalSigns | IntelliVue X3 | LCWI
4576 | Right Cardiac Work Index | - |  | Double | VitalSigns | IntelliVue X3 | RCWI
4577 | Light Intensity (SvO2 Module) | - |  | Double | VitalSigns | IntelliVue X3 | LI
4578 | SpO2 SvO2 Difference | - |  | Double | VitalSigns | IntelliVue X3 | Sp-vO2
4581 | Normalized Post Replacement Volume | - |  | Double | VitalSigns | IntelliVue X3 | VdpoNh
4582 | Normalized Effluent Volume | - | This variable is associated with ChannelID identifier. | Double | VitalSigns | IntelliVue X3 | Veff/EffDn
4584 | Normalized Total Replacement Volume (PBP) | - | This variable is associated with ChannelID identifier. | Double | VitalSigns | IntelliVue X3 | VrePN/VrePNh
4585 | Normalized Replacement Volume (Pre) | - |  | Double | VitalSigns | IntelliVue X3 | VdprNh
4586 | Normalized Ultrafiltration Volume | - | This variable is associated with ChannelID identifier. | Double | VitalSigns | IntelliVue X3 | VufrN/VufrNh
4587 | Patient Fluid Balance Limit Setting | - |  | Double | Settings | IntelliVue X3 | sEPFL
4588 | Total Replacement Volume (PBP) | - | This variable is associated with ChannelID identifier. | Double | VitalSigns | IntelliVue X3 | VreP/VreP/h
4589 | Cumulated Liquid Volume (Pre) | - | This variable is associated with ChannelID identifier. | Double | VitalSigns | IntelliVue X3 | Vdpr/Vdpr/h
4593 | Filtration Fraction Setting | - |  | Double | Settings | IntelliVue X3 | sFufrP
4616 | Diff.-Temperature between Tblood and Tskin | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
4617 | Diff.-Temperature between Tblood and Trect | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
4618 | Diff.-Temperature between Tblood and Tnaso | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
4619 | Diff.-Temperature between Tblood and Tesop | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
4620 | Diff.-Temperature between Tblood and Tcore | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
4621 | Diff.-Temperature between Tblood and Tart | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
4622 | Diff.-Temperature between Tblood and Tven | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
4623 | Diff.-Temperature between Tskin and Trect | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
4624 | Diff.-Temperature between Tskin and Tnaso | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
4625 | Diff.-Temperature between Tskin and Tesop | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
4626 | Diff.-Temperature between Tskin and Tcore | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
4627 | Diff.-Temperature between Tskin and Tart | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
4628 | Diff.-Temperature between Tskin and Tven | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
4629 | Diff.-Temperature between Trect and Tnaso | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
4630 | Diff.-Temperature between Trect and Tesop | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
4631 | Diff.-Temperature between Trect and Tcore | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
4632 | Diff.-Temperature between Trect and Tart | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
4633 | Diff.-Temperature between Trect and Tven | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
4634 | Diff.-Temperature between Tnaso and Tesop | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
4635 | Diff.-Temperature between Tnaso and Tcore | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
4636 | Diff.-Temperature between Tnaso and Tart | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
4637 | Diff.-Temperature between Tnaso and Tven | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
4638 | Diff.-Temperature between Tesop and Tcore | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
4639 | Diff.-Temperature between Tesop and Tart | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
4640 | Diff.-Temperature between Tesop and Tven | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
4641 | Diff.-Temperature between Tcore and Tart | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
4642 | Diff.-Temperature between Tcore and Tven | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
4643 | Diff.-Temperature between Tart and Tven | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
4644 | Newborn Body Length | - |  | Double | VitalSigns | IntelliVue X3 | Length
4646 | Newborn Gestational Age | - |  | Double | PatientInfo | IntelliVue X3 | G.Age
4647 | Calories | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
4648 | Input/Output Totals | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
4649 | Newborn Apgar Scores | - |  | Double | PatientInfo | IntelliVue X3 | Available only in revF or earlier
4650 | BSA Formula Boyd | - |  | Double | PatientInfo | IntelliVue X3 | BSA(B)
4651 | BSA Formula Dubois | - |  | Double | PatientInfo | IntelliVue X3 | BSA(D)
4653 | Pupil Reaction (Left Eye) | - | Light reaction of left eye's pupil. | Double | VitalSigns | IntelliVue X3 | PRL
4654 | Pupil Reaction (Right Eye) | - | Light reaction of right eye's pupil. | Double | VitalSigns | IntelliVue X3 | PRR
4661 | Pleth Gain | - |  | Double | DeviceInfo | IntelliVue X3 | PlGain
4662 | Airway Temperature | - |  | Double | VitalSigns | IntelliVue X3 | Tairwy
4663 | Ventilation Perfusion Index | - |  | Double | VitalSigns | IntelliVue X3 | VENT: VQI
4664 | O2 Extraction Index | - |  | Double | VitalSigns | IntelliVue X3 | O2EI
4665 | Duration Above Base Pressure | - |  | Double | VitalSigns | IntelliVue X3 | VENT: DPosP
4666 | Inspiration Rise Time | - |  | Double | VitalSigns | IntelliVue X3 | VENT: RiseTi
4667 | Continuous Positive Airway Pressure (CPAP) | - |  | Double | VitalSigns | IntelliVue X3 | CPAP
4668 | HFV Amplitude | - |  | Double | VitalSigns | IntelliVue X3 | VENT: HFVAmp
4669 | Urine Urea | - |  | Double | VitalSigns | IntelliVue X3 | UrUrea
4670 | Urine pH Value | - |  | Double | VitalSigns | IntelliVue X3 | UrpH
4671 | Total of CO2 (tCO2) | - | Result of Blood gas Analysis. | Double | VitalSigns | IntelliVue X3 | tCO2
4672 | Total Bilirubin | - |  | Double | VitalSigns | IntelliVue X3 | tBili
4673 | Serum Glucose | - |  | Double | VitalSigns | IntelliVue X3 | SerGlu
4674 | Urine Glucose | - |  | Double | VitalSigns | IntelliVue X3 | UrGlu
4675 | Direct Bilirubin | - |  | Double | VitalSigns | IntelliVue X3 | dBili
4676 | Serum Calcium | - |  | Double | VitalSigns | IntelliVue X3 | SerCa
4677 | Total Serum Calcium | - |  | Double | VitalSigns | IntelliVue X3 | tSerCa
4678 | Serum Magnesium | - |  | Double | VitalSigns | IntelliVue X3 | SerMg
4679 | Serum Phosphat | - |  | Double | VitalSigns | IntelliVue X3 | SerPho
4680 | Serum Kalium (Potassium) | - |  | Double | VitalSigns | IntelliVue X3 | SerK
4681 | Serum Clorid | - |  | Double | VitalSigns | IntelliVue X3 | SerCl
4682 | Serum Albumine | - |  | Double | VitalSigns | IntelliVue X3 | SerAlb
4683 | Urine Clorid | - |  | Double | VitalSigns | IntelliVue X3 | UrCl
4684 | Serum Globulin | - |  | Double | VitalSigns | IntelliVue X3 | SerGlo
4685 | Serum Protein | - |  | Double | VitalSigns | IntelliVue X3 | SerPro
4686 | Serum Urea | - |  | Double | VitalSigns | IntelliVue X3 | SrUrea
4687 | Partial Thromboplastin Time | - |  | Double | VitalSigns | IntelliVue X3 | PTT
4688 | Prothrombin Time | - |  | Double | VitalSigns | IntelliVue X3 | PT
4689 | Thrombin Time | - |  | Double | VitalSigns | IntelliVue X3 | TT
4690 | Thromboplastine Time (Quick Time) | - |  | Double | VitalSigns | IntelliVue X3 | Quick
4691 | Alkalische Phosphatase | - |  | Double | VitalSigns | IntelliVue X3 | AP
4692 | Alpha Amylase | - |  | Double | VitalSigns | IntelliVue X3 | alphaA
4693 | Cholesterinesterase | - |  | Double | VitalSigns | IntelliVue X3 | CHE
4694 | Creatinin Kinase | - |  | Double | VitalSigns | IntelliVue X3 | SerCK
4695 | Creatine Kinase of Type Muscle-Brain | - |  | Double | VitalSigns | IntelliVue X3 | CK-MB
4696 | Creatine Kinase of Type Muscle | - |  | Double | VitalSigns | IntelliVue X3 | CK-MM
4697 | Gamma Glutamyl Transpeptidase | - |  | Double | VitalSigns | IntelliVue X3 | GGT
4698 | Glutamic Oxaloacetic Transaminase | - |  | Double | VitalSigns | IntelliVue X3 | GOT
4699 | Glutamic Pyruvic Transaminase | - |  | Double | VitalSigns | IntelliVue X3 | GPT
4700 | Cholesterin | - |  | Double | VitalSigns | IntelliVue X3 | Chol
4701 | Triglyzeride | - |  | Double | VitalSigns | IntelliVue X3 | TGL
4702 | Urine Protein | - | (Total) Protein in Urine | Double | VitalSigns | IntelliVue X3 | UrPro
4703 | Urine Calcium | - |  | Double | VitalSigns | IntelliVue X3 | UrCa
4704 | Carboxy Hemoglobin | - |  | Double | VitalSigns | IntelliVue X3 | CO-Hb
4705 | Total Protein | - |  | Double | VitalSigns | IntelliVue X3 | tPro
4706 | Lactate Dehydrogenase | - |  | Double | VitalSigns | IntelliVue X3 | LDH
4707 | Aspartin Aminotransferase | - |  | Double | VitalSigns | IntelliVue X3 | AST
4708 | Alveolarproteinose (Alkaline Phosphatase) | - |  | Double | VitalSigns | IntelliVue X3 | ALP
4709 | Reticulocyte Count | - |  | Double | VitalSigns | IntelliVue X3 | RC
4710 | Coagulation Time | - |  | Double | VitalSigns | IntelliVue X3 | CT
4711 | Erithrocyte Sedimentation Rate | - |  | Double | VitalSigns | IntelliVue X3 | ESR
4712 | Packed Cell Volume | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
4713 | Kaolin Cephalin Time | - |  | Double | VitalSigns | IntelliVue X3 | KCT
4715 | Minimum Pressure During Breathing Cycle | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
4716 | Infusion Pump Accumulated Volume | - |  | Double | OtherMeasuredData | IntelliVue X3 | AccVol
4717 | Transthoracic Impedance | - |  | Double | VitalSigns | Common | T.I.
4718 | Inspired Expired N2O Difference | - |  | Double | VitalSigns | IntelliVue X3 | VENT: i-eN2O
4719 | Inspired Expired Halothane Difference | - |  | Double | VitalSigns | IntelliVue X3 | VENT: i-eHAL
4720 | Inspired Expired Enfluran Difference | - |  | Double | VitalSigns | IntelliVue X3 | VENT: i-eENF
4721 | Inspired Expired Isofluran Difference | - |  | Double | VitalSigns | IntelliVue X3 | VENT: i-eISO
4722 | Inspired Expired Sevofluran Difference | - |  | Double | VitalSigns | IntelliVue X3 | VENT: i-eSEV
4723 | Inspired Expired Desfluran Difference | - |  | Double | VitalSigns | IntelliVue X3 | VENT: i-eDES
4724 | Inspired Expired Agent Difference | - |  | Double | VitalSigns | IntelliVue X3 | VENT: i-eAGT
4725 | Inspired Expired O2 Difference | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
4726 | Patient Circuit O2 | - | O2 measured in the Patient Circuit | Double | VitalSigns | IntelliVue X3 | VENT: cktO2
4727 | Barometric Pressure in Sample Chamber | - | This variable is obsolete that may be removed in a future release. It is duplicated with variable 4082. | Double | OtherMeasuredData | IntelliVue X3 | Obsolete Variables
4728 | Minute Volume | - |  | Double | VitalSigns | IntelliVue X3 | MINVOL
4729 | Mean Airway Pressure | - |  | Double | VitalSigns | IntelliVue X3 | MnAwP
4730 | Airway Respiration Rate | - |  | Double | VitalSigns | Common | awRR
4731 | Inspiratory Airway Pressure Waveform | - |  | Double | Waveforms | IntelliVue X3 | Waves: AWPin
4732 | Inspiratory Flow Waveform | - |  | Double | Waveforms | IntelliVue X3 | Waves: AWFin
4733 | HFV Inspired Mandatory Minute Volume | - |  | Double | VitalSigns | IntelliVue X3 | HFMVin
4734 | HFV Gas Transport Coefficient Value | - |  | Double | VitalSigns | IntelliVue X3 | DCO2
4735 | HFV Inspired Tidal Volume | - |  | Double | VitalSigns | IntelliVue X3 | VENT: HFTVin
4736 | HFV Tidal Volume | - |  | Double | VitalSigns | IntelliVue X3 | VENT: HFVTV
4738 | Generic Heart Rate | - | Heart Rate received from an external device | Double | VitalSigns | IntelliVue X3 | extHR
4739 | PERF from the 2nd SpO2/PLETH module | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
4741 | Percent Alpha (Left) | - |  | Double | VitalSigns | IntelliVue X3 | EEG: LT %AL
4742 | Percent Beta (Left) | - |  | Double | VitalSigns | IntelliVue X3 | EEG: LT %BE
4745 | Absolute Alpha (Left) | - |  | Double | VitalSigns | IntelliVue X3 | EEG: LT AL
4746 | Absolute Beta (Left) | - |  | Double | VitalSigns | IntelliVue X3 | EEG: LT BE
4747 | Absolute Delta (Left) | - |  | Double | VitalSigns | IntelliVue X3 | EEG: LT DL
4748 | Absolute Theta (Left) | - |  | Double | VitalSigns | IntelliVue X3 | EEG: LT TH
4749 | Mean Dominant Frequency (Left) | - |  | Double | VitalSigns | IntelliVue X3 | EEG: LT MDF
4750 | Median Power Frequency (Left) | - |  | Double | VitalSigns | IntelliVue X3 | EEG: LT MPF
4751 | Peak Power Frequency (Left) | - |  | Double | VitalSigns | IntelliVue X3 | EEG: LT PPF
4752 | Spectral Edge Frequency (Left) | - |  | Double | VitalSigns | IntelliVue X3 | EEG: LT SEF
4753 | Total Power (Left) | - |  | Double | VitalSigns | IntelliVue X3 | EEG: LT TP
4754 | Scale of the Left Channel EEG Wave | - |  | Double | VitalSigns | IntelliVue X3 | EEG: LSCALE
4755 | Percent Alpha (Right) | - |  | Double | VitalSigns | IntelliVue X3 | EEG: RT %AL
4756 | Percent Beta (Right) | - |  | Double | VitalSigns | IntelliVue X3 | EEG: RT %BE
4759 | Absolute Alpha (Right) | - |  | Double | VitalSigns | IntelliVue X3 | EEG: RT AL
4760 | Absolute Beta (Right) | - |  | Double | VitalSigns | IntelliVue X3 | EEG: RT BE
4761 | Absolute Delta (Right) | - |  | Double | VitalSigns | IntelliVue X3 | EEG: RT DL
4762 | Absolute Theta (Right) | - |  | Double | VitalSigns | IntelliVue X3 | EEG: RT TH
4763 | Mean Dominant Frequency (Right) | - |  | Double | VitalSigns | IntelliVue X3 | EEG: RT MDF
4764 | Median Power Frequency (Right) | - |  | Double | VitalSigns | IntelliVue X3 | EEG: RT MPF
4765 | Peak Power Frequency (Right) | - |  | Double | VitalSigns | IntelliVue X3 | EEG: RT PPF
4766 | Spectral Edge Frequency (Right) | - |  | Double | VitalSigns | IntelliVue X3 | EEG: RT SEF
4767 | Total Power (Right) | - |  | Double | VitalSigns | IntelliVue X3 | EEG: RT TP
4768 | Scale of the Right Channel EEG Wave | - |  | Double | VitalSigns | IntelliVue X3 | EEG: RSCALE
4769 | Duration of Positive Pressure | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
4772 | Fresh Gas N2O Concentration | - |  | Double | VitalSigns | IntelliVue X3 | fgN2O
4774 | Liquid Level (Anesthetic Agent Bottle) | - |  | Double | OtherMeasuredData | IntelliVue X3 | AGTLev
4775 | Liquid Level (ISOflurane Bottle) | - |  | Double | OtherMeasuredData | IntelliVue X3 | ISOLev
4776 | Liquid Level (ENFlurane Bottle) | - |  | Double | OtherMeasuredData | IntelliVue X3 | ENFLev
4777 | Liquid Level (HALothane Bottle) | - |  | Double | OtherMeasuredData | IntelliVue X3 | HALLev
4778 | Liquid Level (DESflurane Bottle) | - |  | Double | OtherMeasuredData | IntelliVue X3 | DESLev
4779 | Liquid Level (SEVoflurane Bottle) | - |  | Double | OtherMeasuredData | IntelliVue X3 | SEVLev
4780 | Invasive Blood Pressure Wave | - |  | Double | Waveforms | IntelliVue X3 | Waves: P
4781 | Urine Shift Volume | - |  | Double | VitalSigns | IntelliVue X3 | UrVSht
4782 | Urine Flow Rate | - |  | Double | VitalSigns | IntelliVue X3 | UrFl
4786 | Mean Dominant Frequency (CH1) | - |  | Double | VitalSigns | IntelliVue X3 | EEG: Frequ1
4787 | Mean Dominant Frequency (CH2) | - |  | Double | VitalSigns | IntelliVue X3 | EEG: Frequ2
4788 | Peak Power Frequency (CH1) | - |  | Double | VitalSigns | IntelliVue X3 | EEG: Frequ1
4789 | Peak Power Frequency (CH2) | - |  | Double | VitalSigns | IntelliVue X3 | EEG: Frequ2
4794 | % of Power in the Alpha Band (CH1) | - |  | Double | VitalSigns | IntelliVue X3 | EEG: Prcnt1
4795 | % of Power in the Alpha Band (CH2) | - |  | Double | VitalSigns | IntelliVue X3 | EEG: Prcnt2
4796 | % of Power in the Beta Band (1st EEG) | - |  | Double | VitalSigns | IntelliVue X3 | EEG: Prcnt1
4797 | % of Power in the Beta Band (2nd EEG) | - |  | Double | VitalSigns | IntelliVue X3 | EEG: Prcnt2
4800 | Hemoglobin | - |  | Double | VitalSigns | IntelliVue X3 | Hb
4801 | Patient Temperature | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
4802 | Capillaries pH Value | - |  | Double | VitalSigns | IntelliVue X3 | pHc
4803 | Adjusted pH (at Patient Temperature) | - | Adjusted pH at Patient Temperature | Double | VitalSigns | IntelliVue X3 | &pH
4804 | Adjusted Capillaries pH Value | - |  | Double | VitalSigns | IntelliVue X3 | &pHc
4805 | Partial O2 in the Capillaries | - |  | Double | VitalSigns | IntelliVue X3 | PcO2
4806 | Adjusted PO2 (at Patient Temperature) | - | Adjusted PO2 at Patient Temperature | Double | VitalSigns | IntelliVue X3 | &PO2
4807 | Adjusted PcO2 (at Patient Temperature) | - |  | Double | VitalSigns | IntelliVue X3 | &PcO2
4808 | Venous CO2 Partial Pressure (PvCO2) | - |  | Double | VitalSigns | IntelliVue X3 | PvCO2
4809 | Partial CO2 in the Capillaries | - |  | Double | VitalSigns | IntelliVue X3 | PcCO2
4811 | Adjusted PcCO2 (at Patient Temperature) | - |  | Double | VitalSigns | IntelliVue X3 | &PcCO2
4812 | Calculated Total CO2 (tCO2) | - |  | Double | VitalSigns | IntelliVue X3 | CALC: 'tCO2
4813 | Calculated SO2 | - | Calculated SaO2 | Double | VitalSigns | IntelliVue X3 | CALC: 'SaO2
4815 | Calculated SvO2 | - |  | Double | VitalSigns | IntelliVue X3 | CALC: 'SvO2
4816 | Calculated ScO2 | - |  | Double | VitalSigns | IntelliVue X3 | CALC: 'ScO2
4817 | Calculated HCO3 | - |  | Double | VitalSigns | IntelliVue X3 | CALC: HCO3
4818 | Base Excess of Extra Cellular Fluid | - |  | Double | VitalSigns | IntelliVue X3 | CALC: Beecf
4819 | Calculated Anion Gap | - |  | Double | VitalSigns | IntelliVue X3 | CALC: AnGap
4821 | Urea | - | Measured by the i-Stat module. | Double | VitalSigns | IntelliVue X3 | Urea
4822 | Manually Entered FiO2 | - | In the Blood Review window. | Double | Settings | IntelliVue X3 | Available only in revF or earlier
4823 | Calculated Arterial Blood Base Excess | - |  | Double | VitalSigns | IntelliVue X3 | BE,B
4824 | Ionized Magnesium | - |  | Double | VitalSigns | IntelliVue X3 | iMg
4826 | Desired Module Temperature | - | It's the selected/adjusted Temp in i-Stat Window. | Double | Settings | IntelliVue X3 | Available only in revF or earlier
4827 | Calculated BUN/Creatinine Ratio | - | Calculated by the i-Stat module. | Double | VitalSigns | IntelliVue X3 | B/Cre
4828 | Calculated Urea/Creatinine Ratio | - | Calculated by the i-Stat module. | Double | VitalSigns | IntelliVue X3 | U/Cre
4829 | Lactate | - | Measured value by the i-Stat module. | Double | VitalSigns | IntelliVue X3 | Lact
4830 | Device Simulator Status | - | Device data with the values: Passed and Failed. | Double | DeviceInfo | IntelliVue X3 | Available only in revF or earlier
4835 | Pressure of Sensor 1 | - | Pressure value of sensor 1. Device specific data. | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
4836 | Pressure of Sensor 2 | - | Pressure value of sensor 2. Device specific data. | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
4837 | Module Temperature | - |  | Double | OtherMeasuredData | IntelliVue X3 | Available only in revF or earlier
4843 | Air Temperature | - |  | Double | OtherMeasuredData | IntelliVue X3 | Available only in revF or earlier
4844 | Micro Environment Air Temperature Setting | - |  | Double | Settings | IntelliVue X3 | sTemp
4847 | Weight of the Urine Disposable Bag | - |  | Double | VitalSigns | IntelliVue X3 | BagWgt
4848 | Total Urine Volume of the Current Measurement Period | - |  | Double | VitalSigns | IntelliVue X3 | tUrVol
4849 | Urine Fluid Density | - |  | Double | VitalSigns | IntelliVue X3 | UrDens
4850 | Urine Fluid Temperature | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
4854 | Ventilation Activated Status | - | Parameter to Informs whether the Patient is Ventilated | Double | VitalSigns | IntelliVue X3 | VENT: PtVent
4856 | Simplified Acuity Physiologic Score | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
4857 | Value Normal | - | Missing parameters in SAPS parameter set which are considered in normal range. | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
4858 | Aspartate Aminotransferase | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
4859 | Spontaneous Respiration Rate | - |  | Double | VitalSigns | IntelliVue X3 | SpRR
4864 | Current Fluid in Urine Bag | - |  | Double | VitalSigns | IntelliVue X3 | BagVol
4865 | Cardiac Function Index (Thermodilution) | - |  | Double | VitalSigns | IntelliVue X3 | iCFI
4866 | Intrathoracic Blood Volume (Thermodilution) | - |  | Double | VitalSigns | IntelliVue X3 | iITVB
4867 | Intrathoracic Blood Volume Index (Thermodilution) | - |  | Double | VitalSigns | IntelliVue X3 | iITBVI
4868 | Extravascular Lung Water (Thermodilution) | - |  | Double | VitalSigns | IntelliVue X3 | iEVLW
4869 | Extravascular Lung Water Index (Thermodilution) | - |  | Double | VitalSigns | IntelliVue X3 | iEVLWI
4873 | NMT TOF Count | - |  | Double | VitalSigns | IntelliVue X3 | NMT: TOFcnt
4874 | NMT TOF Ratio | - |  | Double | VitalSigns | IntelliVue X3 | NMT: TOFrat
4875 | Twitch Height of the 1Hz/0.1Hz Stimulation Response | - |  | Double | VitalSigns | IntelliVue X3 | NMT: Twitch
4876 | Post Tetatic Count | - |  | Double | VitalSigns | IntelliVue X3 | NMT: PTC
4877 | Remaining Time to Next Stimulation | - |  | Double | VitalSigns | IntelliVue X3 | NMT: RemTi
4878 | TOF Response TOF1 | - |  | Double | VitalSigns | IntelliVue X3 | NMT: TOF1
4879 | TOF Response TOF2 | - |  | Double | VitalSigns | IntelliVue X3 | NMT: TOF2
4880 | TOF Response TOF3 | - |  | Double | VitalSigns | IntelliVue X3 | NMT: TOF3
4881 | TOF Response TOF4 | - |  | Double | VitalSigns | IntelliVue X3 | NMT: TOF4
4883 | Activated Partial Thromboplastin Time (aPTT Whole Blood Time) | - |  | Double | VitalSigns | IntelliVue X3 | aPTTWB
4884 | Activated Partial Thromboplastin Time (aPTT Plasma Equivalent Time) | - |  | Double | VitalSigns | IntelliVue X3 | aPTTPE
4886 | Prothrombin Time (Whole Blood Time) | - |  | Double | VitalSigns | IntelliVue X3 | PT WB
4887 | Prothrombin Time (Plasma Equivalent Time) | - |  | Double | VitalSigns | IntelliVue X3 | PT PE
4888 | Prothrombin Time Ratio | - |  | Double | VitalSigns | IntelliVue X3 | PTrat
4889 | Spontaneous Inspiration Time | - |  | Double | VitalSigns | IntelliVue X3 | InsTi
4890 | Overdistension Index | - |  | Double | VitalSigns | IntelliVue X3 | C20/C
4891 | Lung Time Constant | - |  | Double | VitalSigns | IntelliVue X3 | TC
4892 | Airway Correlation Coefficient | - |  | Double | VitalSigns | IntelliVue X3 | r
4893 | Rate Volume Ratio | - |  | Double | VitalSigns | IntelliVue X3 | RVrat
4894 | Normalized Ionized Calcium (Ca) | - |  | Double | VitalSigns | IntelliVue X3 | iCa(N)
4956 | Peak Inspiratory Flow Setting | - |  | Double | Settings | IntelliVue X3 | s%PIF
4957 | Trigger Sensitivity Setting | - |  | Double | Settings | IntelliVue X3 | Available only in revF or earlier
4958 | Plateau Time Setting | - |  | Double | Settings | IntelliVue X3 | sPltTi
4959 | Multiple Sigh Number Setting | - |  | Double | Settings | IntelliVue X3 | VENT: sSghNr
4960 | Apnea Tidal Volume Setting | - |  | Double | Settings | IntelliVue X3 | sATV
4961 | Apnea Respiration Rate Setting | - |  | Double | Settings | IntelliVue X3 | sARR
4962 | Apnea Peak Flow Setting | - |  | Double | Settings | IntelliVue X3 | sAPkFl
4963 | Apnea Inspired O2 Setting | - |  | Double | Settings | IntelliVue X3 | VENT: sAFIO2
4964 | PSV Pressure Setting | - | Pressure Support Ventilation Setting | Double | Settings | IntelliVue X3 | VENT: sPSV
4966 | Sigh Activation Status | - |  | Double | DeviceInfo | IntelliVue X3 | VENT: sEnSgh
4967 | Suction O2 Concentration Setting | - |  | Double | Settings | IntelliVue X3 | VENT: sO2Suc
4968 | PCV Control Pressure Setting | - |  | Double | Settings | IntelliVue X3 | VENT: sPin
4969 | PCV Inspiratory Time Setting | - |  | Double | Settings | IntelliVue X3 | VENT: sPVinT
4970 | PCV Apnea Respiration Rate Setting | - |  | Double | Settings | IntelliVue X3 | VENT: sAPVRR
4971 | PCV Apnea Inspiratory Time Setting | - |  | Double | Settings | IntelliVue X3 | VENT: sAPVTi
4972 | PCV Apnea O2 Setting | - |  | Double | Settings | IntelliVue X3 | VENT: sAPVO2
4973 | PCV Apnea High Airway Pressure Setting | - |  | Double | Settings | IntelliVue X3 | VENT: sAPVhP
4975 | Inspiratory Component of I:E Ratio Setting | - | Pressure Ventilation I component of I:E Ratio setting | Double | Settings | IntelliVue X3 | sPVI
4976 | Expiratory Component of I:E Ratio Setting | - | Pressure Ventilation E component of I:E Ratio setting | Double | Settings | IntelliVue X3 | sPVE
4977 | Inspiratory Component of Apnea Ventilation I:E Ratio Setting | - |  | Double | Settings | IntelliVue X3 | sAPVI
4978 | Expiratory Component of Apnea Ventilation I:E Ratio Setting | - |  | Double | Settings | IntelliVue X3 | sAPVE
4979 | Setting: Pressure Ventilation Control | - |  | Double | Settings | IntelliVue X3 | Available only in revF or earlier
4981 | Cycle Time Setting | - |  | Double | Settings | IntelliVue X3 | sCycTi
4982 | IPPV Ventilation Respiration Rate Setting | - |  | Double | Settings | IntelliVue X3 | VENT: sIPPV
4983 | IMV Ventilation Respiration Rate Setting | - |  | Double | Settings | IntelliVue X3 | VENT: sIMV
4984 | Pressure Support PEEP Setting | - |  | Double | Settings | IntelliVue X3 | VENT: sSPEEP
4987 | Apnea Ventilation Delay | - |  | Double | Settings | IntelliVue X3 | Available only in revF or earlier
4989 | O2 Monitoring Mode | - |  | Double | DeviceInfo | IntelliVue X3 | VENT: sO2Mon
4990 | O2 Calibration Mode | - |  | Double | DeviceInfo | IntelliVue X3 | VENT: sO2Cal
4993 | Inspiratory Time Setting | - |  | Double | Settings | IntelliVue X3 | VENT: sInsTi
4994 | Expiratory Time Setting | - |  | Double | Settings | IntelliVue X3 | VENT: sExpTi
4996 | Alarm Percentage on Rise Time Setting | - |  | Double | Settings | IntelliVue X3 | VENT: sALMRT
4997 | Continuous Positive Airway Pressure (CPAP) Setting | - | This is the device's measurement with label="NLS_NOM_SETT_CONC_AWAY_O2" and physio_id="NOM_CONC_AWAY_O2". | Double | Settings | IntelliVue X3 | sCPAP
4998 | Ventilator Flow Setting | - |  | Double | Settings | IntelliVue X3 | VENT: sFlow
5000 | Low Inspiratory Pressure Setting | - |  | Double | Settings | IntelliVue X3 | sPmin
5001 | HFV Flow Setting | - |  | Double | Settings | IntelliVue X3 | sHFVFl
5002 | HFV Respiration Rate Setting | - |  | Double | Settings | IntelliVue X3 | sHFVRR
5005 | Mechanical Ventilation Activated Status | - | Controlled Mechanical Ventilation Setting | Double | DeviceInfo | IntelliVue X3 | VENT: sCMV
5006 | SIMV Mode | - | Synchronized Intermittent Mandatory Ventilation Setting | Double | DeviceInfo | IntelliVue X3 | VENT: sSIMV
5007 | Mandatory Inspired Minute Volume Setting | - |  | Double | Settings | IntelliVue X3 | VENT: sMMV
5009 | Infusion Pump Delivery Rate Setting | - |  | Double | Settings | IntelliVue X3 | sDRate
5011 | Airway Respiration Rate Setting | - |  | Double | Settings | IntelliVue X3 | sAWRR
5012 | Trigger Flow/Volume Setting | - | This is the device's measurement with label="NLS_NOM_SETT_VENT_VOL_LUNG_TRAPD" and physio_id="NOM_SETT_VENT_VOL_LUNG_TRAPD". | Double | Settings | IntelliVue X3 | VENT: sTrVol
5013 | Apnea Alarm Delay | - |  | Double | DeviceInfo | IntelliVue X3 | sAADel
5014 | HFV Amplitude Setting | - |  | Double | Settings | IntelliVue X3 | sHFVAm
5015 | Volume Alarm Delay | - |  | Double | Settings | IntelliVue X3 | sMVDel
5016 | Pressure Increase Rate Setting | - |  | Double | Settings | IntelliVue X3 | VENT: sPincR
5017 | Volume Warning Setting | - | Delivered by the Evita 2 VueLink Driver. | Double | Settings | IntelliVue X3 | VENT: sVmax
5026 | Airway Inspiratory Pressure Lower Alarm Limit | - |  | Double | Settings | IntelliVue X3 | AlarmLimit: loPmax
5028 | Apnea Alarm Delay Setting | - |  | Double | Settings | IntelliVue X3 | Available only in revF or earlier
5029 | Pressure Ramp Setting (ASB Ramp) | - |  | Double | Settings | IntelliVue X3 | Available only in revF or earlier
5030 | Applied Tidal Volume Setting | - |  | Double | Settings | IntelliVue X3 | sTVap
5031 | Assist Pressure Sensitivity Setting | - | Assist Sensitivity. Used by the Bear 1000 ventilator. | Double | Settings | IntelliVue X3 | sSens
5032 | High Inspiratory Pressure Setting | - |  | Double | Settings | IntelliVue X3 | Available only in revF or earlier
5033 | Background Flow Setting | - |  | Double | Settings | IntelliVue X3 | VENT: sBkgFl
5035 | Vaporizer Agent Concentration Setting | - |  | Double | Settings | IntelliVue X3 | sAGT
5036 | Vaporizer Isoflurane Concentration Setting | - |  | Double | Settings | IntelliVue X3 | sISO
5037 | Vaporizer Enflurane Concentration Setting | - |  | Double | Settings | IntelliVue X3 | sENF
5038 | Vaporizer Halothane Concentration Setting | - |  | Double | Settings | IntelliVue X3 | sHAL
5039 | Vaporizer Desflurane Concentration Setting | - |  | Double | Settings | IntelliVue X3 | sDES
5040 | Vaporizer Sevoflurane Concentration Setting | - |  | Double | Settings | IntelliVue X3 | sSEV
5041 | Gas Measurement Sensor Position | - |  | Double | DeviceInfo | IntelliVue X3 | VENT: sGasPr
5042 | O2 Measurement Sensor Position | - |  | Double | DeviceInfo | IntelliVue X3 | VENT: sO2Pr
5044 | Inspired Tidal Volume Setting | - |  | Double | Settings | IntelliVue X3 | sTVin
5045 | Incubator/Warmer Temperature Setting | - |  | Double | Settings | IntelliVue X3
5046 | Preset Period of UrVol Setting | - |  | Double | Settings | IntelliVue X3 | sUrTi
5050 | Low Time Setting | - | Setting variable for Evita 4 Airway Pressure Release Ventilation Mode. | Double | Settings | IntelliVue X3 | VENT: sTlow
5051 | High Time Setting | - | Setting variable for Evita 4 Airway Pressure Release Ventilation Mode. | Double | Settings | IntelliVue X3 | VENT: sThigh
5052 | Low Pressure Setting | - | Setting variable for Evita 4 Airway Pressure Release Ventilation Mode. | Double | Settings | IntelliVue X3 | VENT: sPlow
5053 | High Pressure Setting | - | Setting variable for Evita 4 Airway Pressure Release Ventilation Mode. | Double | Settings | IntelliVue X3 | VENT: sPhigh
5054 | Volume Assist Level Setting | - |  | Double | Settings | IntelliVue X3 | VENT: sVolas
5055 | CPAP Flow Assist Level Setting | - |  | Double | Settings | IntelliVue X3 | VENT: sFlas
5056 | Preset Stimulation Current Setting | - |  | Double | Settings | IntelliVue X3 | sCurnt
5057 | Preset Stimulation Charge Setting | - |  | Double | Settings | IntelliVue X3 | sChrge
5058 | Preset Stimulation Impulse Duration Setting | - |  | Double | Settings | IntelliVue X3 | sPulsD
5059 | NMT TOF Repetition Time | - | Preset Train of Four (Slow TOF) Repetition Time Setting | Double | DeviceInfo | IntelliVue X3 | NMT: sRepTi
5060 | Panting Limit Setting | - |  | Double | Settings | IntelliVue X3 | VENT: sfmax
5065 | Airway Inspiratory Pressure Upper Alarm Limit | - | Alarm Limit: High Pressure | Double | Settings | IntelliVue X3 | AlarmLimit: highP
5066 | Sigh Volume | - |  | Double | VitalSigns | IntelliVue X3 | Vsigh
5069 | Sustained Pressure Alarm Limit | - |  | Double | Settings | IntelliVue X3 | AlarmLimit: sustP
5072 | Intrathoracic Blood Volume (ITBV) | - |  | Double | VitalSigns | IntelliVue X3 | ITBV
5073 | Intrathoracic Blood Volume Index (ITBI) | - |  | Double | VitalSigns | IntelliVue X3 | ITBVI
5074 | Extravascular Lung Water | - |  | Double | VitalSigns | IntelliVue X3 | EVLW
5075 | Extravascular Lung Water Index | - |  | Double | VitalSigns | IntelliVue X3 | EVLWI
5076 | Global End Diastolic Volume | - |  | Double | VitalSigns | IntelliVue X3 | GEDV
5077 | Global End Diastolic Volume Index | - |  | Double | VitalSigns | IntelliVue X3 | GEDVI
5078 | Cardiac Function Index | - |  | Double | VitalSigns | IntelliVue X3 | CFI
5079 | Mixed Venous O2 Saturation | - |  | Double | VitalSigns | IntelliVue X3 | SvO2
5080 | Generic Term for the Transcutaneous Gases | - |  | Double | VitalSigns | IntelliVue X3 | TCUT: tcGas
5083 | Intrinsic PEEP | - |  | Double | VitalSigns | IntelliVue X3 | iPEEP
5086 | Positive Peak Inspiratory Pressure | - |  | Double | VitalSigns | IntelliVue X3 | PIP
5087 | Dead Space to Tidal Volume Ratio | - |  | Double | VitalSigns | IntelliVue X3 | Vd/Vt
5088 | Lung Volume Trapped | - |  | Double | VitalSigns | IntelliVue X3 | VENT: TrpVol
5089 | Lung Vital Capacity | - |  | Double | VitalSigns | IntelliVue X3 | VC
5090 | Generic Lung Compliance | - |  | Double | VitalSigns | IntelliVue X3 | COMP
5092 | EEG Electromyography | - |  | Double | VitalSigns | IntelliVue X3 | EEG: EMG
5093 | Suppression Ratio | - |  | Double | VitalSigns | IntelliVue X3 | EEG: SR
5095 | EEG Peak Power Frequency | - |  | Double | VitalSigns | IntelliVue X3 | EEG: PPF
5101 | End Tidal N2 | - |  | Double | VitalSigns | IntelliVue X3 | N2
5102 | Inspired N2 | - |  | Double | VitalSigns | IntelliVue X3 | N2
5109 | Generic Agent Wave | - |  | Double | Waveforms | IntelliVue X3 | Waves: AGT
5110 | End Tidal MAC Concentration | - |  | Double | VitalSigns | IntelliVue X3 | MAC
5111 | Left Ventricular Stroke Volume Index | - |  | Double | VitalSigns | IntelliVue X3 | LVSWI
5112 | Left Ventricular Stroke Volume | - |  | Double | VitalSigns | IntelliVue X3 | LVSW
5113 | Right Ventricular Stroke Volume | - |  | Double | VitalSigns | IntelliVue X3 | RVSW
5114 | Glasgow Coma Score | - |  | Double | VitalSigns | IntelliVue X3 | GCS
5115 | SubScore of the GCS - Eye Response | - |  | Double | VitalSigns | IntelliVue X3 | EyeRsp
5116 | SubScore of the GCS - Motoric Response | - |  | Double | VitalSigns | IntelliVue X3 | MotRsp
5117 | SubScore of the GCS - Verbal Response | - |  | Double | VitalSigns | IntelliVue X3 | VblRsp
5118 | Head Circumference | - |  | Double | VitalSigns | IntelliVue X3 | HC
5119 | pH in Arterial Blood | - |  | Double | VitalSigns | IntelliVue X3 | Ph
5120 | Arterial Blood Hemoglobin (HB) | - |  | Double | VitalSigns | IntelliVue X3 | Hb
5121 | pH in Venous Blood | - |  | Double | VitalSigns | IntelliVue X3 | pHv
5122 | Urine Natrium | - |  | Double | VitalSigns | IntelliVue X3 | UrNa
5123 | Natrium in Serum | - |  | Double | VitalSigns | IntelliVue X3 | SerNa
5124 | pH in the Blood Plasma | - |  | Double | VitalSigns | IntelliVue X3 | pH
5125 | Blood Natrium (Sodium) | - |  | Double | VitalSigns | IntelliVue X3 | Na
5126 | Kalium (Potassium) | - |  | Double | VitalSigns | IntelliVue X3 | K
5127 | Generic Glucose Parameter | - |  | Double | VitalSigns | IntelliVue X3 | Glu
5128 | Partial CO2 | - |  | Double | VitalSigns | IntelliVue X3 | PCO2
5129 | Partial O2 | - |  | Double | VitalSigns | IntelliVue X3 | PO2
5130 | Base Excess of Blood | - |  | Double | VitalSigns | IntelliVue X3 | BE
5131 | Percent Alveolarvenous Shunt Qs/Qt | - |  | Double | VitalSigns | IntelliVue X3 | Qs/Qt
5132 | Extravascular Thermal Volume Index | - |  | Double | VitalSigns | IntelliVue X3 | ETVI
5133 | VPB Rate | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
5134 | ECG Lead V1 Wave | - |  | Double | Waveforms | IntelliVue X3 | Waves: V1
5135 | ECG Lead V2 Wave | - |  | Double | Waveforms | IntelliVue X3 | Waves: V2
5136 | ECG Lead V3 Wave | - |  | Double | Waveforms | IntelliVue X3 | Waves: V3
5137 | ECG Lead V4 Wave | - |  | Double | Waveforms | IntelliVue X3 | Waves: V4
5138 | ECG Lead V5 Wave | - |  | Double | Waveforms | IntelliVue X3 | Waves: V5
5139 | ECG Lead V6 Wave | - |  | Double | Waveforms | IntelliVue X3 | Waves: V6
5140 | ECG Lead I Wave | - |  | Double | Waveforms | IntelliVue X3 | Waves: I
5141 | ECG Lead II Wave | - |  | Double | Waveforms | IntelliVue X3 | Waves: II
5142 | ECG Lead III Wave | - |  | Double | Waveforms | IntelliVue X3 | Waves: III
5143 | ECG Lead aVR Wave | - |  | Double | Waveforms | IntelliVue X3 | Waves: aVR
5144 | ECG Lead aVL Wave | - |  | Double | Waveforms | IntelliVue X3 | Waves: aVL
5145 | ECG Lead aVF Wave | - |  | Double | Waveforms | IntelliVue X3 | Waves: aVF
5185 | Mean Dominant Frequency (CH3) | - |  | Double | VitalSigns | IntelliVue X3 | MDF3
5186 | Mean Dominant Frequency (CH4) | - |  | Double | VitalSigns | IntelliVue X3 | MDF4
5189 | NMT DBS Count | - |  | Double | VitalSigns | IntelliVue X3 | DBScnt
5191 | Peak Power Frequency (CH3) | - |  | Double | VitalSigns | IntelliVue X3 | PPF3
5192 | Peak Power Frequency (CH4) | - |  | Double | VitalSigns | IntelliVue X3 | PPF4
5497 | Pump Infusion Mode | - |  | Double | DeviceInfo | IntelliVue X3 | sPuMod
5507 | Remaining Time Counter | - | Time to Elapse Counter used by the i-Stat module. | Double | OtherMeasuredData | IntelliVue X3 | Elapse
5598 | APRV Low Pressure Synchronisation Time Setting | - |  | Double | Settings | IntelliVue X3 | sTloS
5603 | APRV High Pressure Synchronisation Time Setting | - |  | Double | Settings | IntelliVue X3 | sThiS
5614 | Acoustic Respiration Rate | - |  | Double | VitalSigns | IntelliVue X3 | RRac
5616 | Regional O2 Saturation (rSO2 2) | - |  | Double | VitalSigns | Common
5617 | Regional O2 Saturation (rSO2 3) | - |  | Double | VitalSigns | Common
5620 | Regional O2 Saturation (rSO2 4) | - |  | Double | VitalSigns | Common
5632 | Regional O2 Saturation (rSO2 5) | - |  | Double | VitalSigns | IntelliVue X3 | SO2 5
5633 | Regional O2 Saturation (rSO2 6) | - |  | Double | VitalSigns | IntelliVue X3 | SO2 6
5637 | O2 Delivery Index (Thermodilution) | - |  | Double | VitalSigns | IntelliVue X3 | iDO2I
5659 | Noninvasive Venous Hemoglobin (SpHbv) | - |  | Double | VitalSigns | IntelliVue X3 | SpHbv
5660 | Noninvasive Arterial Hemoglobin (SpHba) | - |  | Double | VitalSigns | IntelliVue X3 | SpHba
5661 | Generic Pulse Rate | - |  | Double | VitalSigns | IntelliVue X3 | iAvgPR
5680 | Corrected Total Hemoglobin | - |  | Double | VitalSigns | IntelliVue X3 | cSpHb
5681 | Corrected Arterial Total Hemoglobin | - |  | Double | VitalSigns | IntelliVue X3 | cSpHb
5682 | Corrected Venous Total Hemoglobin | - |  | Double | VitalSigns | IntelliVue X3 | cSpHbv
5694 | Systemic Vascular Resistance (Thermodilution) | - |  | Double | VitalSigns | IntelliVue X3 | iSVR
5695 | Systemic Vascular Resistance Index (Thermodilution) | - |  | Double | VitalSigns | IntelliVue X3 | iSVRI
5707 | Pulse Rate (Abdominal) | - |  | Double | VitalSigns | IntelliVue X3 | Pulse
5719 | ECG Lead V3R | - |  | Double | VitalSigns | IntelliVue X3 | ECG: ST
5720 | ECG Lead V4R | - |  | Double | VitalSigns | IntelliVue X3 | ECG: ST
5721 | ECG Lead V5R | - |  | Double | VitalSigns | IntelliVue X3 | ECG: ST
5763 | Inspired NO | - |  | Double | VitalSigns | IntelliVue X3 | NO
5767 | ECG Lead V9 | - |  | Double | VitalSigns | IntelliVue X3 | ECG: ST
5774 | ECG Lead V7 | - |  | Double | VitalSigns | IntelliVue X3 | ECG: ST
5775 | ECG Lead V8 | - |  | Double | VitalSigns | IntelliVue X3 | ECG: ST
5807 | Apnea Ventilation I:E Ratio Setting | unitless | This variable is associated with ChannelID identifier. This variable is calculated from variable 4977 and 4978, which is sent as a string in the format "1:xx.xx". | String | Settings | IntelliVue X3 | Calculated Variables
5817 | Assisted Respiration Rate | - |  | Double | VitalSigns | IntelliVue X3 | Rrasst
5818 | Correlation Coefficient Squared | - |  | Double | VitalSigns | IntelliVue X3 | r2
5837 | ECG Lead V6R | - |  | Double | VitalSigns | IntelliVue X3 | ECG: ST
5873 | Citrate Solution Concentration | - |  | Double | Settings | IntelliVue X3 | sCitCo
5874 | Calcium Solution Concentration | - |  | Double | Settings | IntelliVue X3 | sCConc
5886 | Citrate Acid Concentration | - |  | Double | Settings | IntelliVue X3 | sCSCAC
5887 | Citrate Load Setting | - |  | Double | Settings | IntelliVue X3 | sCitLo
5949 | Backup Respiration Rate Setting | - |  | Double | Settings | IntelliVue X3 | sRRbak
5958 | Inspired Expired Pressure Difference | - |  | Double | VitalSigns | IntelliVue X3 | sPi-e
5959 | Backup Inspired Expired Pressure Difference Setting | - |  | Double | Settings | IntelliVue X3 | sPi-eB
6124 | Femoral Arterial Pressure Mean | - |  | Double | VitalSigns | IntelliVue X3 | IBP: FAP
6125 | Femoral Arterial Pressure Systolic | - |  | Double | VitalSigns | IntelliVue X3 | IBP: FAP
6126 | Femoral Arterial Pressure Diastolic | - |  | Double | VitalSigns | IntelliVue X3 | IBP: FAP
6281 | Tidal Volume Lower Alarm Limit | - |  | Double | Settings | IntelliVue X3 | AlarmLimit: lowTV
6283 | Minute Volume Lower Alarm Limit | - |  | Double | Settings | IntelliVue X3 | AlarmLimit: lowMV
6322 | ECG Lead V7 Wave | - |  | Double | Waveforms | IntelliVue X3 | Waves: V7
6323 | ECG Lead V8 Wave | - |  | Double | Waveforms | IntelliVue X3 | Waves: V8
6324 | ECG Lead V9 Wave | - |  | Double | Waveforms | IntelliVue X3 | Waves: V9
6326 | Airway Aux Wave | - |  | Double | Waveforms | IntelliVue X3 | Waves: AWPx
6331 | Fetal DECG Wave | - |  | Double | Waveforms | IntelliVue X3 | Waves: DECG
6339 | Patient BSA Formula | unitless | Consult detailed help for more information on this variable | Double | PatientInfo | Common | Patient Demographics
6343 | Platelets (Thrombocyte Count) | - |  | Double | VitalSigns | IntelliVue X3 | Plts
6388 | Endotracheal Tube Type | - |  | Double | DeviceInfo | IntelliVue X3 | sTube
6390 | Minute Volume Upper Alarm Limit | - |  | Double | Settings | IntelliVue X3 | AlarmLimit: highMV
6394 | Pulse Rate from Plethysmogram | - |  | Double | VitalSigns | Common | Pulse
6396 | ECG Ectopic Status | - | Consult detailed help for more information on this variable | Double | VitalSigns | Common | Ectopic Status (on IntelliVue)
6399 | ST Lead ES | - |  | Double | VitalSigns | IntelliVue X3 | ECG: ST
6400 | ST Lead AS | - |  | Double | VitalSigns | IntelliVue X3 | ECG: ST
6401 | ST Lead AI | - |  | Double | VitalSigns | IntelliVue X3 | ECG: ST
6402 | ST Index | - |  | Double | VitalSigns | IntelliVue X3 | ECG: Stindx
6409 | Pulse Pressure | - |  | Double | VitalSigns | IntelliVue X3 | PP
6410 | Mean Transit Time (MTt) | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
6450 | Alveolar Dead Space | - |  | Double | VitalSigns | IntelliVue X3 | Vdalv
6451 | Tympanic Temperature | - |  | Double | VitalSigns | IntelliVue X3 | Ttymp
6452 | Injectate Volume (CO) | - |  | Double | VitalSigns | IntelliVue X3 | InjVol
6453 | SpO2 (Telemetry) | - |  | Double | VitalSigns | IntelliVue X3 | %SpO2T
6454 | Pulse Rate (Telemetry) | - |  | Double | VitalSigns | IntelliVue X3 | PulseT
6455 | Intracranial Pressure ICP1 (Systolic) | - |  | Double | VitalSigns | IntelliVue X3 | IBP: IC2
6456 | Intracranial Pressure ICP1 (Diastolic) | - |  | Double | VitalSigns | IntelliVue X3 | IBP: IC1
6457 | Intracranial Pressure ICP1 (Mean) | - |  | Double | VitalSigns | IntelliVue X3 | IBP: IC3
6458 | Intracranial Pressure ICP2 Systolic | - |  | Double | VitalSigns | IntelliVue X3 | IBP: IC2
6459 | Intracranial Pressure ICP2 Diastolic | - |  | Double | VitalSigns | IntelliVue X3 | IBP: IC2
6460 | Intracranial Pressure ICP2 Mean | - |  | Double | VitalSigns | IntelliVue X3 | IBP: IC2
6461 | Brachial Arterial Pressure (BA) Systolic | - |  | Double | VitalSigns | IntelliVue X3 | IBP: BAP
6462 | Brachial Arterial Pressure (BA) Diastolic | - |  | Double | VitalSigns | IntelliVue X3 | IBP: BAP
6463 | Brachial Arterial Pressure (BA) Mean | - |  | Double | VitalSigns | IntelliVue X3 | IBP: BAP
6464 | Vesical Temperature (Bladder) | - |  | Double | VitalSigns | IntelliVue X3 | Tvesic
6465 | Cerebral Temperature | - |  | Double | VitalSigns | IntelliVue X3 | Tcereb
6466 | ST Level for Lead I | - |  | Double | VitalSigns | IntelliVue X3 | ECG: Rf-I
6467 | ST Level for Lead II | - |  | Double | VitalSigns | IntelliVue X3 | ECG: Rf-II
6468 | ST Level for Lead III | - |  | Double | VitalSigns | IntelliVue X3 | ECG: Rf-III
6469 | ST Level for Lead aVR | - |  | Double | VitalSigns | IntelliVue X3 | ECG: Rf-aVR
6470 | ST Level for Lead aVL | - |  | Double | VitalSigns | IntelliVue X3 | ECG: Rf-aVL
6471 | ST Level for Lead aVF | - |  | Double | VitalSigns | IntelliVue X3 | ECG: Rf-aVF
6472 | ST Level for Lead V1 | - |  | Double | VitalSigns | IntelliVue X3 | ECG: Rf-V1
6473 | ST Level for Lead V2 | - |  | Double | VitalSigns | IntelliVue X3 | ECG: Rf-V2
6474 | ST Level for Lead V3 | - |  | Double | VitalSigns | IntelliVue X3 | ECG: Rf-V3
6475 | ST Level for Lead V4 | - |  | Double | VitalSigns | IntelliVue X3 | ECG: Rf-V4
6476 | ST Level for Lead V5 | - |  | Double | VitalSigns | IntelliVue X3 | ECG: Rf-V5
6477 | ST Level for Lead V6 | - |  | Double | VitalSigns | IntelliVue X3 | ECG: Rf-V6
6478 | Prothrombin Time (International Normalized Ratio) | - |  | Double | VitalSigns | IntelliVue X3 | PT INR
6479 | PSV Tidal Volume | - | Tidal Volume in Pressure Support Ventilation Mode | Double | VitalSigns | IntelliVue X3 | TVPSV
6480 | Negative Inspiratory Force Index | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
6481 | A-Line ARX Index | - |  | Double | VitalSigns | IntelliVue X3 | EEG: AAI
6484 | Impedance Cardiography (ICG) Wave | - |  | Double | Waveforms | IntelliVue X3 | Waves: ICG
6485 | Thoracic Fluid Content | - |  | Double | VitalSigns | IntelliVue X3 | TFC
6486 | Thoracic Fluid Content Index | - |  | Double | VitalSigns | IntelliVue X3 | TFI
6487 | Accelerated Cardiac Index | - |  | Double | VitalSigns | IntelliVue X3 | ACI
6488 | Heart Contractility Index | - |  | Double | VitalSigns | IntelliVue X3 | HI
6489 | O2 Saturation (Left) | - |  | Double | VitalSigns | IntelliVue X3 | SO2 l
6490 | O2 Saturation (Right) | - |  | Double | VitalSigns | IntelliVue X3 | SO2 r
6491 | Pulse Pressure Variation | - |  | Double | VitalSigns | IntelliVue X3 | PPV
6492 | Downslope Time | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
6493 | Pulmonary Vascular Permeability Index | - | This is the device's measurement with label="NLS_NOM_PERM_VASC_PULM_INDEX" and physio_id="NOM_PERM_VASC_PULM_INDEX". See also variable 8791. | Double | VitalSigns | IntelliVue X3 | PVPI
6502 | ART Pulse Rate | - | ART pulse rate | Double | VitalSigns | IntelliVue X3 | Pulse
6543 | Umbilical Arterial Pressure Pulse Rate | - |  | Double | VitalSigns | IntelliVue X3 | Pulse
6544 | Patient Visit Number | unitless | Patient encounter ID contains the ID of the current visit of the patient | String | PatientInfo | IntelliVue X3 | Patient Demographics
6555 | Patient Height at Entered Time | - |  | Double | PatientInfo | Common | Patient Demographics
6557 | Patient Weight at Entered Time | - |  | Double | PatientInfo | Common | Patient Demographics
6567 | Patient Body Surface Area (BSA) | - |  | Double | PatientInfo | Common | Patient Demographics
6685 | Inspired O2 Setting (FiO2) | - |  | Double | Settings | IntelliVue X3 | VENT: sFIO2
6702 | Inspiratory Component of I:E Ratio Setting | - |  | Double | Settings | IntelliVue X3 | sI
6703 | Expiratory Component of I:E Ratio Setting | - |  | Double | Settings | IntelliVue X3 | sE
6735 | Lung Elastance | - |  | Double | VitalSigns | IntelliVue X3 | E
6739 | Backup Inspiratory Time Setting | - |  | Double | Settings | IntelliVue X3 | sInTiB
6744 | Pressure Support Level Above Phigh | - |  | Double | Settings | IntelliVue X3 | sdPin
6746 | Total Respiration Duration Setting | - |  | Double | Settings | IntelliVue X3 | sBreaT
6772 | Heater Power | - |  | Double | DeviceInfo | IntelliVue X3 | sPower
6861 | Capillary Refill Time | - |  | Double | VitalSigns | IntelliVue X3 | CRT
6935 | Estimated Patient Circuit Leakage | - |  | Double | OtherMeasuredData | IntelliVue X3 | LkPat
6936 | Estimated Ventilation Total Leakage | - |  | Double | OtherMeasuredData | IntelliVue X3 | LkTot
6967 | Ventilation Mode Additional Information | - |  | Double | DeviceInfo | IntelliVue X3 | sVSupp
6995 | Backup Plateau Time Setting | - |  | Double | Settings | IntelliVue X3 | sPltTA
7000 | ECG Lead V3R Wave | - |  | Double | Waveforms | IntelliVue X3 | Waves: V3R
7001 | ECG Lead V4R Wave | - |  | Double | Waveforms | IntelliVue X3 | Waves: V4R
7002 | ECG Lead V5R Wave | - |  | Double | Waveforms | IntelliVue X3 | Waves: V5R
7003 | ECG Lead V6R Wave | - |  | Double | Waveforms | IntelliVue X3 | Waves: V6R
7019 | Signal Quality Index (BIS) | - | This variable is associated with ChannelID identifier. | Double | OtherMeasuredData | IntelliVue X3 | SQI L/SQI R
7106 | Continuous NIBP Systolic | - |  | Double | VitalSigns | IntelliVue X3 | cNBP SYS
7107 | Continuous NIBP Diastolic | - |  | Double | VitalSigns | IntelliVue X3 | cNBP DIA
7108 | Continuous NIBP Mean | - |  | Double | VitalSigns | IntelliVue X3 | cNBP MEAN
7118 | Next Intervention Time | - |  | Double | OtherMeasuredData | IntelliVue X3 | InTtvTi
7315 | Total Power (CH3) | - |  | Double | VitalSigns | IntelliVue X3 | TP3
7316 | Total Power (CH4) | - |  | Double | VitalSigns | IntelliVue X3 | TP4
7335 | Total Power (Left) | - |  | Double | VitalSigns | IntelliVue X3 | TP L
7336 | Total Power (Right) | - |  | Double | VitalSigns | IntelliVue X3 | TP R
7414 | PCV I:E Ratio Setting | unitless | This variable is associated with ChannelID identifier. This variable is calculated from variable 4975 and 4976, which is sent as a string in the format "1:xx.xx". | String | Settings | IntelliVue X3 | Calculated Variables
7434 | Inspiratory CO2 (FiCO2) | - |  | Double | VitalSigns | IntelliVue X3 | VENT: FICO2
7532 | End Tidal NO2 | - |  | Double | VitalSigns | IntelliVue X3 | NO2
7535 | Fresh Gas Total Flow | - |  | Double | OtherMeasuredData | IntelliVue X3 | fgFlow
7545 | Bolus Volume Setting | - |  | Double | Settings | IntelliVue X3 | sVBol
7557 | Generic NIBP Pressure | - | Available on IntelliVue devices with firmware revision G or later. | Double | VitalSigns | Common | NBP
7587 | Endotracheal Tube Diameter | - |  | Double | DeviceInfo | IntelliVue X3 | sTubeD
7640 | Mandatory Respiration Rate | - |  | Double | VitalSigns | IntelliVue X3 | VENT: MRR
7641 | Mandatory Minute Volume | - |  | Double | VitalSigns | IntelliVue X3 | VENT: MMV
7642 | Mandatory Tidal Volume | - |  | Double | VitalSigns | IntelliVue X3 | VENT: MTV
7644 | N2O Consumption | - |  | Double | OtherMeasuredData | IntelliVue X3 | AccN2O
7765 | Device Power Up Time | s | Number of seconds since the Monitor device's last power-on. | Double | DeviceInfo | Common | System Information
7768 | O2 Consumption | - |  | Double | VitalSigns | IntelliVue X3 | AccO2
7770 | Air Consumption | - |  | Double | VitalSigns | IntelliVue X3 | AccAir
7773 | Arterial Blood Pressure (ABP) | - | Available on IntelliVue devices with firmware revision G or later. Arterial Blood Pressure (NOM_PRESS_BLD_ART_ABP). | Double | VitalSigns | Common | IBP: ABP
7777 | ABP Pulse Rate | - | ABP pulse rate | Double | VitalSigns | IntelliVue X3 | Pulse
7797 | Signal Noise Ratio (CCO) | - |  | Double | VitalSigns | IntelliVue X3 | SNR
7803 | Spontaneous Airway Respiration Rate | - |  | Double | VitalSigns | IntelliVue X3 | SpAWRR
7804 | Synchronized Respiration Rate | - |  | Double | VitalSigns | IntelliVue X3 | RRsync
7805 | Mechanical Respiration Rate | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
7806 | Daily Urine Output | - |  | Double | VitalSigns | IntelliVue X3 | U/O
7809 | Activated Partial Thromboplastin Time Ratio | - |  | Double | VitalSigns | IntelliVue X3 | PTTrat
7810 | Global Ejection Fraction | - | This is the device's measurement with label="NLS_NOM_FRACT_EJECT" and physio_id="NOM_FRACT_EJECT". See also 8790. | Double | VitalSigns | IntelliVue X3 | GEF
7811 | Generic CO2 Parameter | - | This variable contains "PECO2, Partial expired CO2", which is measured from label="NLS_NOM_VENT_AWAY_CO2_EXP", physio_id="NOM_VENT_AWAY_CO2_EXP". | Double | VitalSigns | IntelliVue X3 | VENT: PECO2
7823 | Central Venous O2 Saturation (ScvO2) | - |  | Double | VitalSigns | IntelliVue X3 | ScvO2
7858 | Suppression Ratio (BIS) | - | This variable is associated with ChannelID identifier. | Double | VitalSigns | IntelliVue X3 | SR L/SR R
7865 | Right Ventricular Stroke Work Index | - |  | Double | VitalSigns | IntelliVue X3 | RVSWI
7870 | Base Excess of Extracellular Fluid | - |  | Double | VitalSigns | IntelliVue X3 | BEecf
7878 | Rise Time Setting | - |  | Double | Settings | IntelliVue X3 | sRmpTi
7879 | Time Duration at Low Pressure (Tlow) | - |  | Double | VitalSigns | IntelliVue X3 | Th/TL
7880 | Time Duration at High Pressure (Thigh) | - |  | Double | VitalSigns | IntelliVue X3 | Th/TL
7904 | Spontaneous Tidal Volume | - |  | Double | VitalSigns | IntelliVue X3 | VENT: SpTV
7905 | Cardiac Troponin l | - |  | Double | VitalSigns | IntelliVue X3 | cTnI
7906 | Cardiopulmonary Bypass Status | - |  | Double | VitalSigns | IntelliVue X3 | CPB
7907 | Cardiac Brain Natriuretic Peptide | - |  | Double | VitalSigns | IntelliVue X3 | BNP
7908 | Mean Dominant Frequency | - |  | Double | VitalSigns | IntelliVue X3 | EEG: MDF
7983 | Filtration Current Volume | - |  | Double | VitalSigns | IntelliVue X3 | Vufr/h
7986 | Calcium Flow Rate Setting | - |  | Double | Settings | IntelliVue X3 | sCSR
7987 | Calcium Compensation Setting | - |  | Double | Settings | IntelliVue X3 | sCComp
8189 | Cerebral State Index (CSI) | - |  | Double | VitalSigns | IntelliVue X3 | CSI
8201 | EEG Amplitude | - | This variable is associated with ChannelID identifier. | Double | VitalSigns | IntelliVue X3 | aEEG1/aEEG2/aEEG3/aEEG4
8339 | Patient Middle Names | unitless | The Middle Name attribute contains the middle name of the patient. | String | PatientInfo | IntelliVue X3 | Patient Demographics
8350 | Ao Pulse Rate | - |  | Double | VitalSigns | IntelliVue X3 | Pulse
8351 | PAP Pulse Rate | - |  | Double | VitalSigns | IntelliVue X3 | Pulse
8352 | CVP Pulse Rate | - |  | Double | VitalSigns | IntelliVue X3 | Pulse
8353 | RAP Pulse Rate | - |  | Double | VitalSigns | IntelliVue X3 | Pulse
8354 | LAP Pulse Rate | - |  | Double | VitalSigns | IntelliVue X3 | Pulse
8355 | ICP Pulse Rate | - |  | Double | VitalSigns | IntelliVue X3 | Pulse
8356 | Umbilical Venous Pressure Pulse Rate | - |  | Double | VitalSigns | IntelliVue X3 | Pulse
8357 | FAP Pulse Rate | - |  | Double | VitalSigns | IntelliVue X3 | Pulse
8358 | BAP Pulse Rate | - |  | Double | VitalSigns | IntelliVue X3 | Pulse
8359 | IC1 Pulse Rate | - |  | Double | VitalSigns | IntelliVue X3 | Pulse
8360 | Intracranial Pressure ICP2 Pulse Rate | - |  | Double | VitalSigns | IntelliVue X3 | Pulse
8361 | Pulse Rate (P1) | - |  | Double | VitalSigns | IntelliVue X3 | Pulse
8362 | Pulse Rate (P2) | - |  | Double | VitalSigns | IntelliVue X3 | Pulse
8363 | Pulse Rate (P3) | - |  | Double | VitalSigns | IntelliVue X3 | Pulse
8364 | Pulse Rate (P4) | - |  | Double | VitalSigns | IntelliVue X3 | Pulse
8365 | Pulse Rate from Plethysmogram (Left) | - |  | Double | VitalSigns | IntelliVue X3 | Pulse
8367 | Pleth Wave (Right) | - |  | Double | Waveforms | IntelliVue X3 | Waves: PLETHr
8368 | Pleth Wave (Left) | - |  | Double | Waveforms | IntelliVue X3 | Waves: PLETHl
8369 | Calculated SO2 | - |  | Double | VitalSigns | IntelliVue X3 | CALC: 'SO2
8370 | Airway Respiration Rate (Ohmeda) | - |  | Double | VitalSigns | IntelliVue X3 | VENT: RRaw
8371 | Airway Respiration Rate Setting (Ohmeda) | - |  | Double | Settings | IntelliVue X3 | VENT: sRRaw
8372 | Fresh Gas O2 Concentration | - |  | Double | OtherMeasuredData | IntelliVue X3 | fgO2
8373 | Manually entered FIO2 or Flow | - | This variable contains measurement from label "NLS_NOM_CONC_AWAY_O2" and observed value "NOM_CONC_AWAY_O2". | Double | VitalSigns | IntelliVue X3 | O2
8374 | Generic O2 Parameter | - | This is the device's measurement with label="NLS_NOM_SETT_CONC_AWAY_O2" and physio_id="NOM_CONC_AWAY_O2". | Double |  | IntelliVue X3 | sO2
8375 | Fresh Gas O2 Flow Setting | - | Mixer parameter. | Double | Settings | IntelliVue X3 | sfgO2
8376 | Airway Pressure Numeric Value (Ohmeda) | - |  | Double | VitalSigns | IntelliVue X3 | AWP
8377 | Inspiratory Peak Flow | - |  | Double | VitalSigns | IntelliVue X3 | PIF
8378 | Maximum Pressure During Breathing Cycle | - |  | Double | VitalSigns | IntelliVue X3 | VENT: Pmax
8379 | Peak Pressure | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
8380 | Peak Inspiratory Flow Setting | - |  | Double | Settings | IntelliVue X3 | sPIF
8381 | Airway Maximum Pressure Setting | - |  | Double | Settings | IntelliVue X3 | VENT: sPmax
8382 | Peak Inspiratory Pressure Setting | - |  | Double | Settings | IntelliVue X3 | sPIP
8383 | Setting: Inspiratory Pressure (Ohmeda) | - |  | Double | Settings | IntelliVue X3 | Available only in revF or earlier
8384 | Mean Airway Pressure (Ohmeda) | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
8385 | EEG Wave (Left) | - |  | Double | Waveforms | IntelliVue X3 | Waves: LT EEG
8386 | EEG Wave (Right) | - |  | Double | Waveforms | IntelliVue X3 | Waves: RT EEG
8387 | % of Power in the Delta Band - 1st EEG | - |  | Double | VitalSigns | IntelliVue X3 | EEG: Prcnt1
8388 | % of Power in the Delta Band - 2nd EEG | - |  | Double | VitalSigns | IntelliVue X3 | EEG: Prcnt2
8389 | Percent Delta - Left Side | - |  | Double | VitalSigns | IntelliVue X3 | EEG: LT %DL
8390 | Percent Delta - Right Side | - |  | Double | VitalSigns | IntelliVue X3 | EEG: RT %DL
8391 | % of Power in the Theta Band (1st EEG) | - |  | Double | VitalSigns | IntelliVue X3 | EEG: Prcnt1
8392 | % of Power in the Theta Band (2nd EEG) | - |  | Double | VitalSigns | IntelliVue X3 | EEG: Prcnt2
8393 | Percent Theta - Left Side | - |  | Double | VitalSigns | IntelliVue X3 | EEG: LT %TH
8394 | Percent Theta - Right Side | - |  | Double | VitalSigns | IntelliVue X3 | EEG: RT %TH
8395 | Generic Calculation Constant | - | This variable contains measurement from Philips label "NLS_NOM_NUM_CALC_CONST" and observed value "NOM_METRIC_NOS". See also variable 8493. | Double | VitalSigns | IntelliVue X3 | CompCt
8396 | Pressure Ventilation Control Pressure | - |  | Double | VitalSigns | IntelliVue X3 | VENT: PVcP
8397 | Right Heart Ejection Fraction | - |  | Double | VitalSigns | IntelliVue X3 | REF
8398 | White Blood Count (Leucocyte Count) | - |  | Double | VitalSigns | IntelliVue X3 | WBC
8399 | Red Blood Count (Erithrocyte Count) | - |  | Double | VitalSigns | IntelliVue X3 | RBC
8400 | Mean Corpuscular Volume | - |  | Double | VitalSigns | IntelliVue X3 | MCV
8401 | Corpuscular Mean Hemoglobin Content | - |  | Double | VitalSigns | IntelliVue X3 | MCH
8402 | Ferrum | - |  | Double | VitalSigns | IntelliVue X3 | Fe
8403 | Fetal Hemoglobin | - |  | Double | VitalSigns | IntelliVue X3 | HbF
8404 | MetHemoglobin | - |  | Double | VitalSigns | IntelliVue X3 | Met-Hb
8405 | Airway Volume Wave | - |  | Double | Waveforms | IntelliVue X3 | Waves: AWV
8406 | Expiratory Peak Flow | - |  | Double | VitalSigns | IntelliVue X3 | PEF
8407 | Expiratory Peak Flow | - |  | Double | VitalSigns | IntelliVue X3 | eeFlow
8408 | The V stands for Volume | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
8409 | The R stands for Rate | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
8410 | Adjusted PaCO2 (at Patient Temperature) | - |  | Double | VitalSigns | IntelliVue X3 | &PaCO2
8411 | Incubator Heating Power | - | Power Required to set Air&Pat Temp in the Incubator | Double | OtherMeasuredData | IntelliVue X3 | INCUBATOR: Power
8412 | Actual Patient Age | - | Actual patient age, provided by a module, measured in years. Please note that this Patient Age will be available only when the corresponding module is connected to the Monitor. Variable 169 contains the patient age as provided by the Monitor itself (calculated from the birth date entered on the Monitor device.) | Double | PatientInfo | IntelliVue X3 | Age
8413 | The P stands for Pressure | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
8414 | PCV Apnea Control Pressure Setting | - |  | Double | Settings | IntelliVue X3 | VENT: sAPVcP
8415 | Fresh Gas Air Flow Setting | - | Mixer parameter. | Double | Settings | IntelliVue X3 | sfgAir
8416 | PEEP/CPAP Lower Alarm Limit | - |  | Double | Settings | IntelliVue X3 | AlarmLimit: loPEEP
8417 | Tidal Volume Upper Alarm Limit | - |  | Double | Settings | IntelliVue X3 | AlarmLimit: highTV
8418 | T-Tre Temperature difference between Temp and Trect | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
8419 | T-Tbl Temperature difference between Temp and Tblood | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
8420 | T-Tco Temperature difference between Temp and Tcore | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
8421 | T-Tsk Temperature difference between Temp and Tskin | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
8422 | T-Tes Temperature difference between Temp and Tesoph | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
8423 | T-Tna Temperature difference between Temp and Tnaso | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
8424 | T-Tar Temperature difference between Temp and Tart | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
8425 | T-Tve Temperature difference between Temp and Tven | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
8426 | T-Taw Temperature difference between Temp and Tairwy | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
8427 | Tre-T Temperature difference between Trect and Temp | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
8428 | Tre-bl Temperature difference between Trect and Tblood | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
8429 | Tre-sk Temperature difference between Trect and Tskin | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
8430 | Tre-aw Temperature difference between Trect and Tairwy | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
8431 | Temperature difference between Tblood and Temp | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
8433 | Tbl-aw Temperature difference between Tblood and Tairwy | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
8434 | Tco-T Temperature difference between Tcore and Temp | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
8435 | Tco-re Temperature difference between Tcore and Trect | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
8436 | Tco-bl Temperature difference between Tcore and Tblood | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
8437 | Tco-sk Temperature difference between Tcore and Tskin | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
8438 | Tco-es Temperature difference between Tcore and Tesoph | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
8439 | Tco-na Temperature difference between Tcore and Tnaso | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
8440 | Tco-aw Temperature difference between Tcore and Tairwy | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
8441 | Tsk-T Temperature difference between Tskin and Temp | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
8442 | Tsk-bl Temperature difference between Tskin and Tblood | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
8443 | Tsk-aw Temperature difference between Tskin and Tairwy | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
8444 | Tes-T Temperature difference between Tesoph and Temp | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
8445 | Tes-re Temperature difference between Tesoph and Trect | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
8446 | Tes-bl Temperature difference between Tesoph and Tblood | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
8447 | Tes-sk Temperature difference between Tesoph and Tskin | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
8448 | Tes-na Temperature difference between Tesoph and Tnaso | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
8449 | Tes-aw Temperature difference between Tesoph and Tairwy | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
8450 | Tna-T Temperature difference between Tnaso and Temp | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
8451 | Tna-re Temperature difference between Tnaso and Trect | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
8452 | Tna-bl Temperature difference between Tnaso and Tblood | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
8453 | Tna-sk Temperature difference between Tnaso and Tskin | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
8454 | Tna-aw Temperature difference between Tnaso and Tairwy | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
8455 | Tar-T Temperature difference between Tart and Temp | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
8456 | Tar-re Temperature difference between Tart and Trect | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
8457 | Tar-bl Temperature difference between Tart and Tblood | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
8458 | Tar-co Temperature difference between Tart and Tcore | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
8459 | Tar-sk Temperature difference between Tart and Tskin | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
8460 | Tar-es Temperature difference between Tart and Tesoph | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
8461 | Tar-na Temperature difference between Tart and Tnaso | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
8462 | Tar-aw Temperature difference between Tart and Tairwy | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
8463 | Tve-T Temperature difference between Tven and Temp | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
8464 | Tve-re Temperature difference between Tven and Trect | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
8465 | Tve-bl Temperature difference between Tven and Tblood | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
8466 | Tve-co Temperature difference between Tven and Tcore | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
8467 | Tve-sk Temperature difference between Tven and Tskin | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
8468 | Tve-es Temperature difference between Tven and Tesoph | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
8469 | Tve-na Temperature difference between Tven and Tnaso | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
8470 | MAC Concentration (Awake) | - |  | Double | VitalSigns | IntelliVue X3 | MAC
8471 | Tve-ar Temperature difference between Tven and Tart | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
8472 | Tve-aw Temperature difference between Tven and Tairwy | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
8473 | Taw-T Temperature difference between Tairwy and Temp | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
8474 | Taw-re Temperature difference between Tairwy and Trect | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
8475 | Taw-bl Temperature difference between Tairwy and Tblood | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
8476 | Taw-co Temperature difference between Tairwy and Tcore | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
8477 | Taw-sk Temperature difference between Tairwy and Tskin | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
8478 | Taw-es Temperature difference between Tairwy and Tesoph | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
8479 | Taw-na Temperature difference between Tairwy and Tnaso | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
8480 | Taw-ar Temperature difference between Tairwy and Tart | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
8481 | Taw-ve Temperature difference between Tairwy and Tven | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
8483 | Generic Agent 1 Wave | - |  | Double | Waveforms | IntelliVue X3 | Waves: AGT1
8484 | End Tidal Generic Agent 1 | - |  | Double | VitalSigns | IntelliVue X3 | AGT1
8485 | Inspired Generic Agent 1 | - |  | Double | VitalSigns | IntelliVue X3 | AGT1
8486 | Generic Agent 2 Wave | - |  | Double | Waveforms | IntelliVue X3 | Waves: AGT2
8487 | End Tidal Generic Agent 2 | - |  | Double | VitalSigns | IntelliVue X3 | AGT2
8488 | Inspired Generic Agent 2 | - |  | Double | VitalSigns | IntelliVue X3 | AGT2
8489 | Adjusted Arterial Blood pH Value | - |  | Double | VitalSigns | IntelliVue X3 | &pHa
8490 | Adjusted Venous Blood pH Value | - |  | Double | VitalSigns | IntelliVue X3 | &pHv
8491 | Adjusted PvCO2 (at Patient Temperature) | - |  | Double | VitalSigns | IntelliVue X3 | &PvCO2
8492 | Adjusted PvO2 (at Patient Temperature) | - |  | Double | VitalSigns | IntelliVue X3 | &PvO2
8493 | Generic Calculation Constant | - | This variable contains parameter from Philips label "NLS_NOM_NUM_CATHETER_CONST" and observed value "NOM_NUM_CATHETER_CONST". See also variable 8395. | Double | VitalSigns | IntelliVue X3 | CathCt
8501 | ECG Lead V Wave | - |  | Double | Waveforms | IntelliVue X3 | Waves: V
8502 | ECG Lead MCL Wave | - |  | Double | Waveforms | IntelliVue X3 | Waves: MCL
8503 | ECG Lead MCL1 Wave | - |  | Double | Waveforms | IntelliVue X3 | Waves: MCL1
8510 | MAC Concentration | - |  | Double | VitalSigns | IntelliVue X3 | MAC
8511 | Pleth Wave (Telemetry) | - |  | Double | Waveforms | IntelliVue X3 | Waves: PlethT
8513 | NIBP Systolic (Telemetry) | - |  | Double | VitalSigns | IntelliVue X3 | NBP_T SYS
8514 | NIBP Diastolic (Telemetry) | - |  | Double | VitalSigns | IntelliVue X3 | NBP_T DIA
8515 | NIBP Mean (Telemetry) | - |  | Double | VitalSigns | IntelliVue X3 | NBP_T MEAN
8517 | Generic Pressure 1 (P1) Systolic | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
8518 | Generic Pressure 1 (P1) Diastolic | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
8519 | Generic Pressure 1 (P1) Mean | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
8521 | Generic Pressure 2 (P2) Systolic | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
8522 | Generic Pressure 2 (P2) Diastolic | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
8523 | Generic Pressure 2 (P2) Mean | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
8525 | Generic Pressure 3 (P3) Systolic | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
8526 | Generic Pressure 3 (P3) Diastolic | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
8527 | Generic Pressure 3 (P3) Mean | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
8529 | Generic Pressure 4 (P4) Systolic | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
8530 | Generic Pressure 4 (P4) Diastolic | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
8531 | Generic Pressure 4 (P4) Mean | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
8535 | Incubator Humidity | - |  | Double | OtherMeasuredData | IntelliVue X3 | INCUBATOR: Hum
8536 | Predictive Oral Temperature | - |  | Double | VitalSigns | IntelliVue X3 | pToral
8537 | Predictive Rectal Temperature | - |  | Double | VitalSigns | IntelliVue X3 | pTrect
8538 | Predictive Axillary Temperature | - |  | Double | VitalSigns | IntelliVue X3 | pTaxil
8539 | Perfusion (Telemetry) | - |  | Double | VitalSigns | IntelliVue X3 | Perf T
8541 | Right to Left Heart Shunt | - |  | Double | VitalSigns | IntelliVue X3 | RLShnt
8542 | Adjusted PcO2 (at Patient Temperature) | - |  | Double | VitalSigns | IntelliVue X3 | &PCO2
8543 | Adjusted Arterial PaO2 (at Patient Temperature) | - |  | Double | VitalSigns | IntelliVue X3 | &PaO2
8544 | Pulse Rate from Plethysmogram (Right) | - |  | Double | VitalSigns | IntelliVue X3 | Pulse
8549 | Patient O2 Consumption per Body Surface | - |  | Double | VitalSigns | IntelliVue X3 | VO2m²
8550 | Patient O2 Consumption per Body Weight | - |  | Double | VitalSigns | IntelliVue X3 | VO2kg
8560 | Ideal Patient Weight | - |  | Double | PatientInfo | IntelliVue X3 | IBW
8562 | End Flow Setting | - |  | Double | Settings | IntelliVue X3 | sEndFl
8564 | PSV Backup Delay Setting | - |  | Double | Settings | IntelliVue X3 | sPSVbd
8640 | Technical Alarm Internal Source ID | unitless | Consult detailed help for more information on this variable | Double (Multiset) | Alarms | Common | Alarms
8641 | Technical Alarm Source Type | unitless |  | Double (Multiset) | Alarms | Common | Alarms
8642 | Device Technical Alarm | unitless | Consult detailed help for more information on this variable | Double (Multiset) | Alarms | Common | Alarms
8643 | Patient Alarm Internal Source ID | unitless | Consult detailed help for more information on this variable | Double (Multiset) | Alarms | Common | Alarms
8644 | Patient Alarm Source Type | unitless |  | Double (Multiset) | Alarms | Common | Alarms
8645 | Patient Alarm | unitless | Consult detailed help for more information on this variable | Double (Multiset) | Alarms | Common | Alarms
8656 | Accoustic Respiration Rate Wave | - |  | Double | Waveforms | IntelliVue X3 | Waves: acRESP
8703 | Blood Temperature Wave | - |  | Double | Waveforms | IntelliVue X3 | Waves: Tblood
8714 | Non-specific Pressure 1 Wave | - |  | Double | Waveforms | IntelliVue X3 | Waves: P1
8715 | Generic Invasive Blood Pressure P1 | - | Available on IntelliVue devices. | Double | VitalSigns | IntelliVue X3 | IBP: P1
8716 | Generic Invasive Blood Pressure P1 Systolic | - |  | Double | VitalSigns | IntelliVue X3 | IBP: P1
8717 | Generic Invasive Blood Pressure P1 Diastolic | - |  | Double | VitalSigns | IntelliVue X3 | IBP: P1
8718 | Generic Invasive Blood Pressure P1 Mean | - |  | Double | VitalSigns | IntelliVue X3 | IBP: P1
8719 | Non-specific Pressure 2 Wave | - |  | Double | Waveforms | IntelliVue X3 | Waves: P2
8720 | Generic Invasive Blood Pressure P2 | - | Available on IntelliVue devices. | Double | VitalSigns | IntelliVue X3 | IBP: P2
8721 | Generic Invasive Blood Pressure P2 Systolic | - |  | Double | VitalSigns | IntelliVue X3 | IBP: P2
8722 | Generic Invasive Blood Pressure P2 Diastolic | - |  | Double | VitalSigns | IntelliVue X3 | IBP: P2
8723 | Generic Invasive Blood Pressure P2 Mean | - |  | Double | VitalSigns | IntelliVue X3 | IBP: P2
8724 | Non-specific Pressure 3 Wave | - |  | Double | Waveforms | IntelliVue X3 | Waves: P3
8725 | Generic Invasive Blood Pressure P3 | - | Available on IntelliVue devices. | Double | VitalSigns | IntelliVue X3 | IBP: P3
8726 | Generic Invasive Blood Pressure P3 Systolic | - |  | Double | VitalSigns | IntelliVue X3 | IBP: P3
8727 | Generic Invasive Blood Pressure P3 Diastolic | - |  | Double | VitalSigns | IntelliVue X3 | IBP: P3
8728 | Generic Invasive Blood Pressure P3 Mean | - |  | Double | VitalSigns | IntelliVue X3 | IBP: P3
8729 | Non-specific Pressure 4 Wave | - |  | Double | Waveforms | IntelliVue X3 | Waves: P4
8730 | Generic Invasive Blood Pressure P4 | - | Available on IntelliVue devices. | Double | VitalSigns | IntelliVue X3 | IBP: P4
8731 | Generic Invasive Blood Pressure P4 Systolic | - |  | Double | VitalSigns | IntelliVue X3 | IBP: P4
8732 | Generic Invasive Blood Pressure P4 Diastolic | - |  | Double | VitalSigns | IntelliVue X3 | IBP: P4
8733 | Generic Invasive Blood Pressure P4 Mean | - |  | Double | VitalSigns | IntelliVue X3 | IBP: P4
8734 | Non-specific Pressure 5 Wave | - |  | Double | Waveforms | IntelliVue X3 | Waves: P_5
8735 | Generic Invasive Blood Pressure P5 | - | Available on IntelliVue devices. | Double | VitalSigns | IntelliVue X3 | IBP: P5
8736 | Generic Invasive Blood Pressure P5 Systolic | - |  | Double | VitalSigns | IntelliVue X3 | IBP: P5
8737 | Generic Invasive Blood Pressure P5 Diastolic | - |  | Double | VitalSigns | IntelliVue X3 | IBP: P5
8738 | Generic Invasive Blood Pressure P5 Mean | - |  | Double | VitalSigns | IntelliVue X3 | IBP: P5
8739 | Non-specific Pressure 6 Wave | - |  | Double | Waveforms | IntelliVue X3 | Waves: P_6
8740 | Generic Invasive Blood Pressure P6 | - | Available on IntelliVue devices. | Double | VitalSigns | IntelliVue X3 | IBP: P6
8741 | Generic Invasive Blood Pressure P6 Systolic | - |  | Double | VitalSigns | IntelliVue X3 | IBP: P6
8742 | Generic Invasive Blood Pressure P6 Diastolic | - |  | Double | VitalSigns | IntelliVue X3 | IBP: P6
8743 | Generic Invasive Blood Pressure P6 Mean | - |  | Double | VitalSigns | IntelliVue X3 | IBP: P6
8744 | Non-specific Pressure 7 Wave | - |  | Double | Waveforms | IntelliVue X3 | Waves: P_7
8745 | Generic Invasive Blood Pressure P7 | - | Available on IntelliVue devices. | Double | VitalSigns | IntelliVue X3 | IBP: P7
8746 | Generic Invasive Blood Pressure P7 Systolic | - |  | Double | VitalSigns | IntelliVue X3 | IBP: P7
8747 | Generic Invasive Blood Pressure P7 Diastolic | - |  | Double | VitalSigns | IntelliVue X3 | IBP: P7
8748 | Generic Invasive Blood Pressure P7 Mean | - |  | Double | VitalSigns | IntelliVue X3 | IBP: P7
8749 | Non-specific Pressure 8 Wave | - |  | Double | Waveforms | IntelliVue X3 | Waves: P_8
8750 | Generic Invasive Blood Pressure P8 | - | Available on IntelliVue devices. | Double | VitalSigns | IntelliVue X3 | IBP: P8
8751 | Generic Invasive Blood Pressure P8 Systolic | - |  | Double | VitalSigns | IntelliVue X3 | IBP: P8
8752 | Generic Invasive Blood Pressure P8 Diastolic | - |  | Double | VitalSigns | IntelliVue X3 | IBP: P8
8753 | Generic Invasive Blood Pressure P8 Mean | - |  | Double | VitalSigns | IntelliVue X3 | IBP: P8
8754 | Intra-Uterine Pressure (IUP) Wave | - |  | Double | Waveforms | IntelliVue X3 | Waves: IUP
8755 | Intra-Uterine Pressure | - |  | Double | VitalSigns | IntelliVue X3 | IBP: IUP
8756 | Intra-Uterine Pressure Systolic | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
8757 | Intra-Uterine Pressure Diastolic | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
8758 | Intra-Uterine Pressure Mean | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
8760 | Unspecified Blood Pressure | - |  | Double | VitalSigns | IntelliVue X3 | IBP: P
8761 | Unspecified Blood Pressure Systolic | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
8762 | Unspecified Blood Pressure Diastolic | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
8763 | Unspecified Blood Pressure Mean | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
8788 | Airway Flow Wave | - |  | Double | Waveforms | IntelliVue X3 | Waves: AWF
8789 | Auxiliary Wave/Parameter | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
8790 | Global Ejection Fraction (2) | - | This is the device's measurement with label="NLS_NOM_EMFC_GEF" and physio_id="NOM_METRIC_NOS". See also 7810. | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
8791 | Pulmonary Vascular Permeability Index (2) | - | This is the device's measurement with label="NLS_NOM_EMFC_PVPI" and physio_id="NOM_METRIC_NOS". See also variable 6493. | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
8793 | Pressure Drop | - |  | Double | OtherMeasuredData | IntelliVue X3 | DP
8803 | Replacement Pre-delivered Rate Setting | - |  | Double | Settings | IntelliVue X3 | sRPrFl
8804 | Replacement Post-delivered Rate Setting | - |  | Double | Settings | IntelliVue X3 | sRPoFl
8805 | Pre-blood Flow Rate Setting | - |  | Double | Settings | IntelliVue X3 | sPBPFl
8806 | Patient Weight Balance Setting | - |  | Double | Settings | IntelliVue X3 | sWTbal
8808 | Anticoagulant Bolus Interval Setting | - |  | Double | Settings | IntelliVue X3 | sBolIn
8809 | Anticoagulant Infusion Rate Setting | - |  | Double | Settings | IntelliVue X3 | sHepRt
8810 | Ultrafiltration Ratio Setting | - |  | Double | Settings | IntelliVue X3 | sUFR
8811 | Therapy Elapsed Time | - |  | Double | OtherMeasuredData | IntelliVue X3 | RunTi
8812 | Airway Volume | - |  | Double | VitalSigns | IntelliVue X3 | Available only in revF or earlier
8813 | Processed Cumulated Blood Volume | - |  | Double | VitalSigns | IntelliVue X3 | BVolPr
8814 | Replacement Pre-delivered Cumulated Volume | - |  | Double | VitalSigns | IntelliVue X3 | RpreC
8815 | Replacement Post-delivered Cumulated Volume | - | This variable is associated with ChannelID identifier. | Double | VitalSigns | IntelliVue X3 | Vdpo/h / RpostC
8816 | Pre-blood Cumulated Volume | - | This variable is associated with ChannelID identifier. | Double | VitalSigns | IntelliVue X3 | VPBP/h / PreBC
8835 | Syringe Speed | - |  | Double | OtherMeasuredData | IntelliVue X3 | Vsyr/h
8851 | Maximum Volume Gain Setting | - |  | Double | Settings | IntelliVue X3 | sMaxG
8853 | Patient Fluid Balance | - |  | Double | VitalSigns | IntelliVue X3 | Wtcum
8857 | Arterial Blood Pressure (ABP) Wave | - |  | Double | Waveforms | IntelliVue X3 | Waves: ABP
8858 | Arterial Blood Pressure (ART) Wave | - |  | Double | Waveforms | IntelliVue X3 | Waves: ART
8859 | Pulmonary Arterial Pressure (PAP) Wave | - |  | Double | Waveforms | IntelliVue X3 | Waves: PAP
8860 | Central Venous Pressure (CVP) Wave | - |  | Double | Waveforms | IntelliVue X3 | Waves: CVP
8861 | Right Atrial Pressure (RAP) Wave | - |  | Double | Waveforms | IntelliVue X3 | Waves: RAP
8862 | Left Atrial Pressure (LAP) Wave | - |  | Double | Waveforms | IntelliVue X3 | Waves: LAP
8863 | Intra-cranial Pressure (ICP) Wave | - |  | Double | Waveforms | IntelliVue X3 | Waves: ICP
8864 | Umbilical Arterial Pressure (UAP) Wave | - |  | Double | Waveforms | IntelliVue X3 | Waves: UAP
8865 | Umbilical Venous Pressure (UVP) Wave | - |  | Double | Waveforms | IntelliVue X3 | Waves: UVP
8866 | Airway CO2 Level Waveform | - |  | Double | Waveforms | IntelliVue X3 | Waves: CO2
8867 | Airway O2 Wave | - |  | Double | Waveforms | IntelliVue X3 | Waves: O2
8868 | Airway N2 Wave | - |  | Double | Waveforms | IntelliVue X3 | Waves: N2
8869 | Generic Agent Wave (2nd) | - |  | Double | Waveforms | IntelliVue X3 | Waves: AGTs
8870 | Impedance Respiration Wave | - |  | Double | Waveforms | IntelliVue X3 | Waves: Resp
8871 | Brachial Arterial Pressure (BAP) Wave | - |  | Double | Waveforms | IntelliVue X3 | Waves: BAP
8872 | Intracranial Pressure 1 (IC1) Wave | - |  | Double | Waveforms | IntelliVue X3 | Waves: IC1
8873 | Intracranial Pressure 2 (IC2) Wave | - |  | Double | Waveforms | IntelliVue X3 | Waves: IC2
8874 | Femoral Arterial Pressure (FAP) Wave | - |  | Double | Waveforms | IntelliVue X3 | Waves: FAP
8875 | Vector ECG Wave | - |  | Double | Waveforms | IntelliVue X3 | Waves: vECG
8889 | Transcutaneous Heater Power | - |  | Double | OtherMeasuredData | IntelliVue X3 | HeatPa
8939 | BIS Standard Deviation | - | This variable is associated with ChannelID identifier. | Double | VitalSigns | IntelliVue X3 | sdBISl/sdBISr
8940 | EMG Standard Deviation (BIS) | - | This variable is associated with ChannelID identifier. | Double | VitalSigns | IntelliVue X3 | sdEMGl/sdEMGr
8941 | Asymmetric Index | - | This variable is associated with ChannelID identifier. | Double | VitalSigns | IntelliVue X3 | ASYM/ASYM L/ASYM R
9037 | O2 Consumption Index Estimate | - |  | Double | VitalSigns | IntelliVue X3 | iVO2Ie
9110 | Aortic Blood Pressure (Ao) | - | Available on IntelliVue devices with firmware revision G or later. Arterial Blood Pressure in the Aorta (NOM_PRESS_BLD_AORT). | Double | VitalSigns | Common | IBP: Ao
9111 | Umbilical Arterial Pressure (UAP) | - | Available on IntelliVue devices with firmware revision G or later. | Double | VitalSigns | Common | IBP: UAP
9112 | Umbilical Venous Pressure (UVP) | - | Available on IntelliVue devices with firmware revision G or later. | Double | VitalSigns | Common | IBP: UVP
9154 | Humidifier Activation Status | - |  | Double | DeviceInfo | IntelliVue X3 | sHmCtl
9155 | Nebulizer Time Setting | - |  | Double | Settings | IntelliVue X3 | sNebTi
9165 | Edi Diaphragm Electrical Activity Waveform | - |  | Double | Waveforms | IntelliVue X3 | Waves: EDI
9166 | Peak Diaphragm Electrical Activity (Edi) | - |  | Double | VitalSigns | IntelliVue X3 | Edimax
9167 | Minimum Diaphragm Electrical Activity (Edi) | - |  | Double | VitalSigns | IntelliVue X3 | Edimin
9168 | % Edi Trigger In NAVA | - |  | Double | VitalSigns | IntelliVue X3 | %EdiTr
9169 | % Edi Cycle Off In NAVA | - |  | Double | VitalSigns | IntelliVue X3 | %EdiCy
9170 | NAVA Level Setting | - |  | Double | Settings | IntelliVue X3 | sNAVA
9171 | Edi Trigger Setting | - |  | Double | Settings | IntelliVue X3 | sEdiTr
9196 | Patient Body Mass Index (BMI) | - |  | Double | PatientInfo | IntelliVue X3 | BMI
9270 | Maximum Positive Auxiliary Pressure | - |  | Double | VitalSigns | IntelliVue X3 | Pxpeak
9271 | Mean Positive Auxiliary Pressure | - |  | Double | VitalSigns | IntelliVue X3 | Pxmean
9272 | Minimum Positive Auxiliary Pressure | - |  | Double | VitalSigns | IntelliVue X3 | Pxmin
9286 | Patient CO2 Production Indexed By Body Surface | - |  | Double | VitalSigns | IntelliVue X3 | VCO2m²
9287 | Patient CO2 Production Indexed By Body Weight | - |  | Double | VitalSigns | IntelliVue X3 | VCO2kg
9288 | Tidal Volume Indexed By Body Weight | - |  | Double | VitalSigns | IntelliVue X3 | TV/kg
9314 | Functional Residual Capacity | - |  | Double | VitalSigns | IntelliVue X3 | FRC
9315 | Corrected QT Interval (QTc) | - |  | Double | VitalSigns | IntelliVue X3 | ECG: QTc
9317 | Airway N2 Concentration | - | Measurement from VueLink | Double | VitalSigns | IntelliVue X3 | N2
9318 | Airway Agent Concentration | - | Measurement from VueLink. This variable is associated with a SourceID identifier. | Double | VitalSigns | IntelliVue X3 | AGT
9321 | Respiration Rate (Spirometry) | - |  | Double | VitalSigns | IntelliVue X3 | RRspir
9323 | QT Heart Rate (QT-HR) | - | QT heart rate | Double | VitalSigns | IntelliVue X3 | ECG: QT-HR
9324 | Delta QTc | - |  | Double | VitalSigns | IntelliVue X3 | ECG: DeltaQTc
9329 | Humidification Type Setting | - |  | Double | Settings | IntelliVue X3 | sHumid
9336 | Normalized Rapid Shallow Breathing Index | - |  | Double | VitalSigns | IntelliVue X3 | RSBInm
9337 | Percent of Spontaneous Inspiratory Time | - |  | Double | VitalSigns | IntelliVue X3 | Tin/Tt
9408 | Energy Balance Calculated by Metabolics | - |  | Double | VitalSigns | IntelliVue X3 | dE
9437 | Alveolar Volume | - |  | Double | VitalSigns | IntelliVue X3 | Valv
9438 | Apparatus Dead Space | - |  | Double | VitalSigns | IntelliVue X3 | appVd
9439 | Pleth (Pre-Ductal) | - | Pulse Rate from Plethysmogram (pre ductal) | Double | VitalSigns | IntelliVue X3 | Pulse
9440 | Pleth (Post-Ductal) | - | Pulse Rate from Plethysmogram (post ductal) | Double | VitalSigns | IntelliVue X3 | Pulse
9441 | Relative Perfusion (Pre-Ductal) | - |  | Double | VitalSigns | IntelliVue X3 | PerfPr
9442 | Relative Perfusion (Post-Ductal) | - |  | Double | VitalSigns | IntelliVue X3 | PerfPo
9444 | Cardioplegia Main Pump Time of Current Plegia Session | - |  | Double | VitalSigns | IntelliVue X3 | CpOnTi
9445 | Cardioplegia Main Pump Time Since Stop of Plegia Session | - |  | Double | VitalSigns | IntelliVue X3 | CpOffT
9446 | Cardioplegia Main Pump Total Plegia Time | - |  | Double | VitalSigns | IntelliVue X3 | CpPlTi
9447 | Cardioplegia Slave Pump Time Since Start of Current Plegia Session | - |  | Double | VitalSigns | IntelliVue X3 | CsOnTi
9448 | Cardioplegia Slave Pump Time Since Stop of Plegia Session | - |  | Double | VitalSigns | IntelliVue X3 | CsOffT
9449 | Cardioplegia Slave Pump Total Plegia Time | - |  | Double | VitalSigns | IntelliVue X3 | CsPlTi
9453 | PAV-based Lung Compliance | - |  | Double | VitalSigns | IntelliVue X3 | Cpav
9454 | PAV-based Patient Resistance | - |  | Double | VitalSigns | IntelliVue X3 | Rpav
9455 | PAV-based Total Resistance Estimate | - |  | Double | VitalSigns | IntelliVue X3 | Rtot
9456 | PAV-based Lung Elastance | - |  | Double | VitalSigns | IntelliVue X3 | Epav
9467 | Backup Inspiratory Pressure Setting | - |  | Double | Settings | IntelliVue X3 | sPIbak
9476 | Percent of Leakage | - |  | Double | OtherMeasuredData | IntelliVue X3 | %Leak
9485 | Maximum Inspiratory Time Setting | - |  | Double | Settings | IntelliVue X3 | sTimax
9585 | Plateau Time | - |  | Double | VitalSigns | IntelliVue X3 | PlatTi
9586 | QT Baseline | - |  | Double | VitalSigns | IntelliVue X3 | ECG: QT Bsl
9587 | QT Heart Rate Baseline | - |  | Double | VitalSigns | IntelliVue X3 | ECG: QTHRBl
9588 | High Density Lipoprotein | - |  | Double | VitalSigns | IntelliVue X3 | HDL
9589 | Low Density Lipoprotein | - |  | Double | VitalSigns | IntelliVue X3 | LDL
9590 | C-reactive Protein | - |  | Double | VitalSigns | IntelliVue X3 | CRP
9591 | Urine Hemoglobin | - |  | Double | VitalSigns | IntelliVue X3 | UrHb
9595 | Noninvasive Total Hemoglobin (SpHb) | - |  | Double | VitalSigns | IntelliVue X3 | SpHb
9596 | Noninvasive Oxygen Content (SpOC) | - |  | Double | VitalSigns | IntelliVue X3 | SpOC
9597 | Noninvasive Carboxyhemoglobin (SpCO) | - |  | Double | VitalSigns | IntelliVue X3 | SpCO
9598 | Noninvasive Methemoglobin (SpMet) | - |  | Double | VitalSigns | IntelliVue X3 | SpMet
9600 | Pleth Variability Index (PVI) | - |  | Double | VitalSigns | IntelliVue X3 | PVI
9603 | Regional O2 Saturation (rSO2) | - |  | Double | VitalSigns | Common
9605 | rSO2 Area Under Curve (AUC) | - | This variable is associated with ChannelID identifier. | Double | VitalSigns | IntelliVue X3 | AUC_1/AUC_2/AUC_3/AUC_4/AUC_5/AUC_6
9610 | Generic O2 Saturation (3) | - |  | Double | VitalSigns | IntelliVue X3 | SO2 3
9611 | Generic O2 Saturation (4) | - |  | Double | VitalSigns | IntelliVue X3 | SO2 4
9612 | Tissue O2 Saturation (StO2) | - |  | Double | VitalSigns | IntelliVue X3 | StO2
9613 | Nitrogen Balance (NsLoss) | - |  | Double | VitalSigns | IntelliVue X3 | NsLoss
52523 | Inspired O2 Concentration | % | See variable 52527 for partial pressure | Double | VitalSigns | IntelliVue X3 | O2
52527 | Inspired O2 Partial Pressure | mmHg | This variable is calculated from variable 52523 and will output inspired O2 partial pressure. | Double | VitalSigns | IntelliVue X3 | O2

## MP20
ID | Label | Unit | Description
2 | NIBP Systolic | -
3 | NIBP Diastolic | -
4 | NIBP Mean | -
7 | Patient Delta Temperature | -
14 | SpO2 | -
19 | ECG Heart Rate | -
22 | Respiration Rate | -
26 | Device Inactivation Status | unitless | Possible values: "enabled" and "audio-paused". See "Multiset Alarms" on help file.
28 | Active Alarms | unitless | See "Multiset Alarms" on help file.
Consult detailed help for more information on this variable
29 | Latched Alarms | unitless | See "Multiset Alarms" on help file.
Consult detailed help for more information on this variable
33 | Inactive Alarms | unitless | See "Multiset Alarms" on help file.
Consult detailed help for more information on this variable
34 | General ST Level 1 | - | Physiological identifier: NOM_ECG_AMPL_ST_I
35 | General ST Level 2 | - | Physiological identifier: NOM_ECG_AMPL_ST_II
36 | General ST Level 3 | - | Physiological identifier: NOM_ECG_AMPL_ST_III
127 | ECG Measurement Status | - | ECG Rhythm Status
Consult detailed help for more information on this variable
169 | Patient Age | - | Patient age. The Monitor device calculates the age of the patient from the birth data manually entered on the Monitor.
170 | Patient Sex | unitless | Possible values: "unknown", "male", "female", and "unspecified".
497 | ECG Lead Name | - | Physiological identifier: NOM_ECG_AMPL_ST
501 | Arterial Pressure (ART) Systolic | -
502 | Arterial Pressure (ART) Diastolic | -
503 | Arterial Pressure (ART) Mean | -
512 | Left Atrial Mean Pressure (LAP) | -
519 | Central Venous Pressure (CVP) | - | Available on IntelliVue devices with firmware revision G or later.
538 | Device Language | unitless | Consult detailed help for more information on this variable
608 | End Tidal CO2 | -
630 | Cardiac Beat-to-Beat Rate | - | Cardiac Beat-to-Beat Rate (NOM_ECG_CARD_BEAT_RATE_BTB).
764 | Device Specification ID | unitless | Variables 764, 765 and 4201 must be interpreted conjointly.
Consult detailed help for more information on this variable
765 | Device Specification Type | unitless | Variables 764, 765 and 4201 must be interpreted conjointly.
Consult detailed help for more information on this variable
1220 | Patient Birth Date | unitless | Format: 14 ascii digits string of "YYYYMMDDhhmmss"
1255 | ECG Lead MCL1 | - | Physiological identifier: NOM_ECG_AMPL_ST_MCL
1257 | ECG Lead V1 | - | Physiological identifier: NOM_ECG_AMPL_ST_V1
1258 | ECG Lead V6 | - | Physiological identifier: NOM_ECG_AMPL_ST_V6
1259 | ECG Lead aVL | - | Physiological identifier: NOM_ECG_AMPL_ST_AVL
1260 | ECG Lead aVF | - | Physiological identifier: NOM_ECG_AMPL_ST_AVF
1261 | ECG Lead aVR | - | Physiological identifier: NOM_ECG_AMPL_ST_AVR
1262 | ECG Lead V2 | - | Physiological identifier: NOM_ECG_AMPL_ST_V2
1263 | ECG Lead V3 | - | Physiological identifier: NOM_ECG_AMPL_ST_V3
1264 | ECG Lead V4 | - | Physiological identifier: NOM_ECG_AMPL_ST_V4
1265 | ECG Lead V5 | - | Physiological identifier: NOM_ECG_AMPL_ST_V5
1523 | Premature Ventricular Contractions | -
1533 | LAP Systolic | -
1534 | LAP Diastolic | -
1911 | Device Unique ID | unitless
1914 | Patient Bed ID | unitless | This variable contains a printable string identifying the system location.
1929 | Patient ID | unitless
1930 | Patient Name | unitless
1942 | Patient Comment Text | unitless
2200 | Device Hardware Revision | unitless | This variable is obsolete because it may contains incorrect information. Please use variables 764, 765 and 4201 must be used instead.
2201 | Serial Number | unitless | This variable is obsolete because it may contains incorrect information. Please use variables 764, 765 and 4201 must be used instead.
2255 | Device Date/Time | unitless | Format: 16 ascii digits string of "YYYYMMDDhhmmssff", where the "ff" is second_fractions that is not used by IntelliVue monitors.
2259 | Arterial BP Systolic | -
2260 | Arterial BP Diastolic | -
2261 | Arterial BP Mean | -
2262 | Pulmonary Artery BP - Systolic | -
2263 | Pulmonary Artery BP - Diastolic | -
2264 | Pulmonary Artery BP Mean | -
2271 | Right Atrial Pressure (RAP) | - | Available on IntelliVue devices with firmware revision G or later.
2272 | Left Atrial Pressure (LAP) | - | Available on IntelliVue devices with firmware revision G or later.
2273 | Intra Cranial Pressure (ICP) | - | Available on IntelliVue devices with firmware revision G or later.
2321 | Generic NIBP Systolic Pressure | - | This variable is same as variable 2, except that the early estimated values are marked as NonSignificant in this variable.
2322 | Generic NIBP Diastolic Pressure | - | This variable is same as variable 3, except that the early estimated values are marked as NonSignificant in this variable.
2323 | Generic NIBP Mean Pressure | - | This variable is same as variable 4, except that the early estimated values are marked as NonSignificant in this variable.
2598 | Patient Admission Status | unitless | Possible values: "empty", "pre-admitted", "admitted", and "discharged".
2602 | ECG Lead V | - | Physiological identifier: NOM_ECG_AMPL_ST_V
2821 | System General Status | unitless | This variable identifies the current operating mode of the device. Possible values: "not specified", "monitoring", "demo", and "service". Moreover, "- standby" or "- config" can be appended
2845 | Frequency | Hz | This variable describes the frequency of the main power supply. Possible values: "unspecified", "50", and "60".
2855 | Device Status | unitless | Possible values: "disconnected", "unassociated", "operating".
2879 | Device Software Revision | unitless | This variable is obsolete because it may contains incorrect information. Please use variables 764, 765 and 4201 must be used instead.
2901 | Patient First Name | unitless
2907 | Patient Temperature | -
2995 | Device Altitude | m
3506 | Inspiratory Minimum CO2 | -
3570 | Skin Temperature | -
3571 | Rectal Temperature | -
3572 | Nasopharyngeal Temperature | -
3573 | Esophageal Temperature | -
3574 | Core Temperature | -
3575 | Arterial Temperature | -
3576 | Venous Temperature | -
3591 | Device Number | unitless | Device Model Number (example: "M8000")
3639 | Manufacturers Name | unitless | "Philips"
3646 | Arterial Blood Pressure (ART) | - | Available on IntelliVue devices with firmware revision G or later. Arterial Blood Pressure (NOM_PRESS_BLD_ART).
3647 | Pulmonary Artery Blood Pressure (PA) | - | Available on IntelliVue devices with firmware revision G or later.
3789 | Arterial O2 Saturation (SaO2) | - | Obsolete variable that may be removed in a future release. The label of this variable is not correct. Please use variable 14 instead.
3853 | Patient Type | unitless | Possible values: "unspecified", "adult", "pediatric", and "neonatal".
3854 | Patient Pace Maker Status | unitless | Possible values: "not paced", "paced gen".
4084 | Pulse Rate ( Unspecified Source ) | -
4089 | Central Venous Pressure Diastolic | -
4093 | Central Venous Mean Pressure | -
4094 | Central Venous Pressure Systolic | -
4201 | Product Description | unitless | Variables 764, 765 and 4201 must be interpreted conjointly.
Consult detailed help for more information on this variable
4457 | System Alarm Status | unitless | Overall alert state of the system.
Consult detailed help for more information on this variable
4458 | Patient Alarm Priority | unitless | Overall Patient alarm priority. This is the current highest priority value of currently active Patient Alarms.
Consult detailed help for more information on this variable
4459 | Technical Alarm Priority | unitless | Overall Technical alarm priority. This is the current highest priority value of currently active Technical Alarms.
Consult detailed help for more information on this variable
4460 | Audio Alarm Priority | unitless | Overall audible alarm priority. This is the current highest priority value of currently active audible alarms.
Consult detailed help for more information on this variable
4461 | Technical Alarm Source ID | unitless
4462 | Technical Alarm Global State | unitless | Consult detailed help for more information on this variable
4463 | Technical Alarm Priority | unitless | Consult detailed help for more information on this variable
4464 | Technical Alarm Flag | unitless | Consult detailed help for more information on this variable
4465 | Technical Alarm Priority | unitless
4466 | Technical Alarm Description | unitless
4467 | Patient Alarm Source ID | unitless
4468 | Patient Alarm Global State | unitless | Consult detailed help for more information on this variable
4469 | Patient Alarm Priority | unitless | Consult detailed help for more information on this variable
4470 | Patient Alarm Flag | unitless | Consult detailed help for more information on this variable
4471 | Patient Alarm Priority | unitless
4472 | Patient Alarm Description | unitless
4473 | Device Nomenclature Revision | unitless | Format: major.minor (such as 1.0)
4474 | Device Text Catalog Revision | unitless | Format: "MM.NN mm.nn", where "MM.NN" is text catalog revision and "mm.nn" is language revision.
4475 | Device Data Format | unitless | Always 11 (Unicode).
4476 | Device Operating Area | unitless | Possible values: "unspecified", "operating room", "intensive care unit", "neonatal intensive care unit", and "cardiology care unit".
4513 | Ao Systolic | -
4514 | Ao Diastolic | -
4515 | Ao Mean | -
4516 | RAP Systolic | -
4517 | RAP Diastolic | -
4518 | RAP Mean | -
4519 | Intracranial Pressure ICP Systolic | -
4520 | Intracranial Pressure ICP Diastolic | -
4521 | ICP Mean | -
4522 | UAP Systolic | -
4523 | UAP Diastolic | -
4524 | UAP Mean | -
4525 | UVP Systolic | -
4526 | UVP Diastolic | -
4527 | UVP Mean | -
4528 | Unspecific Pressure Systolic | -
4529 | Unspecific Pressure Diastolic | -
4530 | Unspecific Pressure Mean | -
4717 | Transthoracic Impedance | -
4730 | Airway Respiration Rate | -
5616 | Regional O2 Saturation (rSO2 2) | -
5617 | Regional O2 Saturation (rSO2 3) | -
5620 | Regional O2 Saturation (rSO2 4) | -
6339 | Patient BSA Formula | unitless | Consult detailed help for more information on this variable
6394 | Pulse Rate from Plethysmogram | -
6396 | ECG Ectopic Status | - | Consult detailed help for more information on this variable
6555 | Patient Height at Entered Time | -
6557 | Patient Weight at Entered Time | -
6567 | Patient Body Surface Area (BSA) | -
7557 | Generic NIBP Pressure | - | Available on IntelliVue devices with firmware revision G or later.
7765 | Device Power Up Time | s | Number of seconds since the Monitor device's last power-on.
7773 | Arterial Blood Pressure (ABP) | - | Available on IntelliVue devices with firmware revision G or later. Arterial Blood Pressure (NOM_PRESS_BLD_ART_ABP).
8640 | Technical Alarm Internal Source ID | unitless | Consult detailed help for more information on this variable
8641 | Technical Alarm Source Type | unitless
8642 | Device Technical Alarm | unitless | Consult detailed help for more information on this variable
8643 | Patient Alarm Internal Source ID | unitless | Consult detailed help for more information on this variable
8644 | Patient Alarm Source Type | unitless
8645 | Patient Alarm | unitless | Consult detailed help for more information on this variable
9110 | Aortic Blood Pressure (Ao) | - | Available on IntelliVue devices with firmware revision G or later. Arterial Blood Pressure in the Aorta (NOM_PRESS_BLD_AORT).
9111 | Umbilical Arterial Pressure (UAP) | - | Available on IntelliVue devices with firmware revision G or later.
9112 | Umbilical Venous Pressure (UVP) | - | Available on IntelliVue devices with firmware revision G or later.
9603 | Regional O2 Saturation (rSO2) | -

## MP70
ID | Label | Unit | Description | Type | Group | Device | Source
2 | NIBP Systolic | - |  | Double | VitalSigns | Common | NBP
3 | NIBP Diastolic | - |  | Double | VitalSigns | Common | NBP
4 | NIBP Mean | - |  | Double | VitalSigns | Common | NBP
7 | Patient Delta Temperature | - |  | Double | VitalSigns | Common | DeltaTemp
14 | SpO2 | - |  | Double | VitalSigns | Common | SpO2
16 | NIBP Pulse Rate | - |  | Double | VitalSigns | IntelliVue MP70 | Pulse
19 | ECG Heart Rate | - |  | Double | VitalSigns | Common | ECG: HR
22 | Respiration Rate | - |  | Double | VitalSigns | Common | RR
26 | Device Inactivation Status | unitless | Possible values: "enabled" and "audio-paused". See "Multiset Alarms" on help file. | String | DeviceInfo | Common | Alarms
28 | Active Alarms | unitless | See "Multiset Alarms" on help file. | Double (Multiset) | Alarms | Common | Alarms
Consult detailed help for more information on this variable
29 | Latched Alarms | unitless | See "Multiset Alarms" on help file. | Double (Multiset) | Alarms | Common | Alarms
Consult detailed help for more information on this variable
33 | Inactive Alarms | unitless | See "Multiset Alarms" on help file. | Double (Multiset) | Alarms | Common | Alarms
Consult detailed help for more information on this variable
34 | General ST Level 1 | - | Physiological identifier: NOM_ECG_AMPL_ST_I | Double | VitalSigns | Common | ECG: ST
35 | General ST Level 2 | - | Physiological identifier: NOM_ECG_AMPL_ST_II | Double | VitalSigns | Common | ECG: ST
36 | General ST Level 3 | - | Physiological identifier: NOM_ECG_AMPL_ST_III | Double | VitalSigns | Common | ECG: ST
45 | Temperature 3 | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
46 | Temperature 4 | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
48 | O2 Saturation (SO2) | - |  | Double | VitalSigns | IntelliVue MP70 | SO2
113 | Pleth Waveform | - |  | Double | Waveforms | IntelliVue MP70 | Waves: Pleth
127 | ECG Measurement Status | - | ECG Rhythm Status | Double | OtherMeasuredData | Common | ECG Status (on IntelliVue)
Consult detailed help for more information on this variable
169 | Patient Age | - | Patient age. The Monitor device calculates the age of the patient from the birth data manually entered on the Monitor. | Double | PatientInfo | Common | Patient Demographics
170 | Patient Sex | unitless | Possible values: "unknown", "male", "female", and "unspecified". | String | PatientInfo | Common | Patient Demographics
405 | Generic O2 Saturation (1) | - |  | Double | VitalSigns | IntelliVue MP70 | SO2 1
406 | Generic O2 Saturation (2) | - |  | Double | VitalSigns | IntelliVue MP70 | SO2 2
414 | Cardiac Output | - |  | Double | VitalSigns | IntelliVue MP70 | C.O.
465 | Fetal Heart Rate (FHR) | - | This variable is associated with ChannelID identifier. | Double | VitalSigns | IntelliVue MP70 | DFHR/DFHR1/DFHR2/DFHR3
467 | Fetal Beat-to-Beat Rate (FHR) | - | This variable is associated with ChannelID identifier. | Double | VitalSigns | IntelliVue MP70 | bDFHR/bDFHR1/bDFHR2/bDFHR3
469 | TOCO | - |  | Double | VitalSigns | IntelliVue MP70 | Toco
497 | ECG Lead Name | - | Physiological identifier: NOM_ECG_AMPL_ST | Double | OtherMeasuredData | Common | Available only in revF or earlier
501 | Arterial Pressure (ART) Systolic | - |  | Double | VitalSigns | Common | IBP: ART
502 | Arterial Pressure (ART) Diastolic | - |  | Double | VitalSigns | Common | IBP: ART
503 | Arterial Pressure (ART) Mean | - |  | Double | VitalSigns | Common | IBP: ART
512 | Left Atrial Mean Pressure (LAP) | - |  | Double | VitalSigns | Common | IBP: LAP
519 | Central Venous Pressure (CVP) | - | Available on IntelliVue devices with firmware revision G or later. | Double | VitalSigns | Common | IBP: CVP
538 | Device Language | unitless | Consult detailed help for more information on this variable | String | DeviceInfo | Common | System Information
591 | ECG Lead II | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
592 | Airway N2O Concentration | - | Measurement from VueLink | Double | VitalSigns | IntelliVue MP70 | N2O
593 | Inspired Desflurane | - |  | Double | VitalSigns | IntelliVue MP70 | DES
595 | Inspired Sevoflurane | - |  | Double | VitalSigns | IntelliVue MP70 | SEV
597 | Fresh Gas Agent Value | - |  | Double | VitalSigns | IntelliVue MP70 | fgAGT
598 | Fresh Gas Halothane | - |  | Double | VitalSigns | IntelliVue MP70 | fgHAL
599 | Fresh Gas Enflurane | - |  | Double | VitalSigns | IntelliVue MP70 | fgENF
600 | Fresh Gas Isoflurane | - |  | Double | VitalSigns | IntelliVue MP70 | fgISO
601 | Inspired Halothane | - |  | Double | VitalSigns | IntelliVue MP70 | HAL
603 | Inspired Enflurane | - |  | Double | VitalSigns | IntelliVue MP70 | ENF
605 | Inspired Isoflurane | - |  | Double | VitalSigns | IntelliVue MP70 | ISO
608 | End Tidal CO2 | - |  | Double | VitalSigns | Common | CO2
616 | Airway Pressure | - | Measured in the inspiratory path. | Double | VitalSigns | IntelliVue MP70 | AWPin
618 | Airway N2O Wave | - |  | Double | Waveforms | IntelliVue MP70 | Waves: N2O
620 | Halothane Concentration Wave | - |  | Double | Waveforms | IntelliVue MP70 | Waves: HAL
621 | Enflurane Concentration Wave | - |  | Double | Waveforms | IntelliVue MP70 | Waves: ENF
622 | Isoflurane Concentration Wave | - |  | Double | Waveforms | IntelliVue MP70 | Waves: ISO
623 | Desflurane Concentration Wave | - |  | Double | Waveforms | IntelliVue MP70 | Waves: DES
624 | Sevoflurane Concentration Wave | - |  | Double | Waveforms | IntelliVue MP70 | Waves: SEV
625 | Apnea Duration | - | This is the device's measurement with label="NLS_NOM_TIME_PD_APNEA" and physio_id="NOM_TIME_PD_APNEA". | Double | VitalSigns | IntelliVue MP70 | ApneaD
626 | End Tidal CO2 | - |  | Double | VitalSigns | IntelliVue MP70 | CRT
630 | Cardiac Beat-to-Beat Rate | - | Cardiac Beat-to-Beat Rate (NOM_ECG_CARD_BEAT_RATE_BTB). | Double | VitalSigns | Common | ECG: btbHR
633 | Airway Delta O2 Concentration | - |  | Double | VitalSigns | IntelliVue MP70 | VENT: DeltaO2
649 | Cardiac Output Index | - |  | Double | VitalSigns | IntelliVue MP70 | C.I.
650 | Stroke Volume | - |  | Double | VitalSigns | IntelliVue MP70 | SV
651 | Stroke Volume Index | - |  | Double | VitalSigns | IntelliVue MP70 | SI
654 | Systemic Vascular Resistance | - |  | Double | VitalSigns | IntelliVue MP70 | SVR
669 | Normalized Ionized Calcium Concentration (Ca) | - |  | Double | VitalSigns | IntelliVue MP70 | iCa-N
678 | Potassium Ion Concentration (K+) | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
696 | QT Interval (QT) | - |  | Double | VitalSigns | IntelliVue MP70 | ECG: QT
764 | Device Specification ID | unitless | Variables 764, 765 and 4201 must be interpreted conjointly. | String (Multiset) | DeviceInfo | Common | System Information
Consult detailed help for more information on this variable
765 | Device Specification Type | unitless | Variables 764, 765 and 4201 must be interpreted conjointly. | String (Multiset) | DeviceInfo | Common | System Information
Consult detailed help for more information on this variable
794 | Pleth Wave (2) | - |  | Double | Waveforms | IntelliVue MP70 | Waves: PLETH2
854 | Total Number of Used Disposable | - |  | Double | DeviceInfo | IntelliVue MP70 | nFilt
946 | Cumulated Syringe Infused Volume | - |  | Double | OtherMeasuredData | IntelliVue MP70 | SyrC
966 | Current Therapy Time | - |  | Double | OtherMeasuredData | IntelliVue MP70 | FiltTi
1043 | Manually Entered Hematocrit | - |  | Double | Settings | IntelliVue MP70 | scHCTi
1055 | Infusion Drug Name | - |  | Double | PatientInfo | IntelliVue MP70 | sDrug
1072 | Infusion Time Remaining | - |  | Double | VitalSigns | IntelliVue MP70 | InfTim
1095 | Infusion Dose Rate Setting | - |  | Double | Settings | IntelliVue MP70 | sDosRt
1096 | Volume To Be Infused Original Setting | - |  | Double | Settings | IntelliVue MP70 | sVtot
1110 | KVO Rate Setting | - |  | Double | Settings | IntelliVue MP70 | sKVO
1186 | Inspired Minute Volume | - |  | Double | VitalSigns | IntelliVue MP70 | MINVOL
1190 | I:E Ratio | - | This variable contains expiratory component of I:E Ratio assuming that the inspiratory component is "1". | Double | VitalSigns | IntelliVue MP70 | I:E 1:
1208 | SpO2 | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
1220 | Patient Birth Date | unitless | Format: 14 ascii digits string of "YYYYMMDDhhmmss" | String | PatientInfo | Common | Patient Demographics
1234 | SpO2 (Pre-Ductal) | - |  | Double | VitalSigns | IntelliVue MP70 | SpO2pr
1236 | SpO2 (Post-Ductal) | - |  | Double | VitalSigns | IntelliVue MP70 | SpO2po
1244 | ECG Lead III | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
1246 | Hematocrit | - |  | Double | VitalSigns | IntelliVue MP70 | Hct
1250 | Pulmonary Vascular Resistance | - |  | Double | VitalSigns | IntelliVue MP70 | PVR
1251 | Body Surface Area | - |  | Double | PatientInfo | IntelliVue MP70 | BSA
1255 | ECG Lead MCL1 | - | Physiological identifier: NOM_ECG_AMPL_ST_MCL | Double | VitalSigns | Common | ECG: ST
1257 | ECG Lead V1 | - | Physiological identifier: NOM_ECG_AMPL_ST_V1 | Double | VitalSigns | Common | ECG: ST
1258 | ECG Lead V6 | - | Physiological identifier: NOM_ECG_AMPL_ST_V6 | Double | VitalSigns | Common | ECG: ST
1259 | ECG Lead aVL | - | Physiological identifier: NOM_ECG_AMPL_ST_AVL | Double | VitalSigns | Common | ECG: ST
1260 | ECG Lead aVF | - | Physiological identifier: NOM_ECG_AMPL_ST_AVF | Double | VitalSigns | Common | ECG: ST
1261 | ECG Lead aVR | - | Physiological identifier: NOM_ECG_AMPL_ST_AVR | Double | VitalSigns | Common | ECG: ST
1262 | ECG Lead V2 | - | Physiological identifier: NOM_ECG_AMPL_ST_V2 | Double | VitalSigns | Common | ECG: ST
1263 | ECG Lead V3 | - | Physiological identifier: NOM_ECG_AMPL_ST_V3 | Double | VitalSigns | Common | ECG: ST
1264 | ECG Lead V4 | - | Physiological identifier: NOM_ECG_AMPL_ST_V4 | Double | VitalSigns | Common | ECG: ST
1265 | ECG Lead V5 | - | Physiological identifier: NOM_ECG_AMPL_ST_V5 | Double | VitalSigns | Common | ECG: ST
1283 | Patient Temperature | - | This variable is associated with a SourceID Identifier. | Double | VitalSigns | IntelliVue MP70 | Infrared Temp
1286 | Micro Environment Air Temperature | - |  | Double | OtherMeasuredData | IntelliVue MP70 | INCUBATOR: Air T
1306 | Tidal Volume | - |  | Double | VitalSigns | IntelliVue MP70 | TV
1319 | Static Resistance | - |  | Double | VitalSigns | IntelliVue MP70 | Rstat
1321 | Tidal Volume Setting | - |  | Double | Settings | IntelliVue MP70 | sTV
1323 | Pressure Sensitivity Setting | - |  | Double | Settings | IntelliVue MP70 | sSenPr
1324 | Sigh Interval Setting | - |  | Double | Settings | IntelliVue MP70 | sTsigh
1325 | Sigh Volume Setting | - |  | Double | Settings | IntelliVue MP70 | sVsigh
1332 | Pressure Support Level Above PEEP Setting | - |  | Double | Settings | IntelliVue MP70 | sdPEEP
1333 | Spontaneous Minute Volume | - |  | Double | VitalSigns | IntelliVue MP70 | VENT: SpMV
1340 | Negative Inspiratory Pressure | - |  | Double | VitalSigns | IntelliVue MP70 | NgInsP
1342 | Peak Spontaneous Flow | - |  | Double | VitalSigns | IntelliVue MP70 | SpPkFl
1343 | Flow-by Base Flow Setting | - |  | Double | Settings | IntelliVue MP70 | VENT: sBasFl
1344 | Flow Sensitivity Setting | - |  | Double | Settings | IntelliVue MP70 | VENT: sSenFl
1348 | Expiratory Time | - |  | Double | VitalSigns | IntelliVue MP70 | ExpTi
1381 | Inspiratory Flow Setting | - |  | Double | Settings | IntelliVue MP70 | sInsFl
1382 | Expiratory Flow Setting | - |  | Double | Settings | IntelliVue MP70 | sExpFl
1383 | Inspiratory Flow Trigger Setting | - | Delivered by the Evita 2 VueLink Driver. | Double | Settings | IntelliVue MP70 | VENT: sTrgFl
1387 | APRV Low Pressure Setting | - |  | Double | Settings | IntelliVue MP70 | sPSVl
1389 | APRV High Pressure Setting | - |  | Double | Settings | IntelliVue MP70 | sPSVh
1391 | Apnea Time Setting | - |  | Double | Settings | IntelliVue MP70 | sAlnTi
1398 | Minute Volume Setting | - |  | Double | Settings | IntelliVue MP70 | sMV
1418 | Positive End Expiratory Pressure (PEEP) | - |  | Double | VitalSigns | IntelliVue MP70 | VENT: PEEP
1419 | Ventilator Work of Breathing | - |  | Double | VitalSigns | IntelliVue MP70 | WOBv
1422 | Rapid Shallow Breathing Index | - |  | Double | VitalSigns | IntelliVue MP70 | RSBI
1424 | Patient Work of Breathing | - |  | Double | VitalSigns | IntelliVue MP70 | WOBp
1435 | Systemic Vascular Resistance Index | - |  | Double | VitalSigns | IntelliVue MP70 | SVRI
1436 | Pulmonary Vascular Resistance Index | - |  | Double | VitalSigns | IntelliVue MP70 | PVRI
1442 | Continuous Cardiac Output | - |  | Double | VitalSigns | IntelliVue MP70 | CCO
1443 | Continuous Cardiac Output Index | - |  | Double | VitalSigns | IntelliVue MP70 | CCI
1451 | Pump Infusion Status | - |  | Double | DeviceInfo | IntelliVue MP70 | PuStat
1463 | Hemoglobin (Hb) | - | This variable is associated with SourceID identifier. | Double | VitalSigns | IntelliVue MP70 | sHb
1464 | Pleth Wave (Pre-Ductal) | - |  | Double | Waveforms | IntelliVue MP70 | Waves: PLTHpr
1465 | Pleth Wave (Post-Ductal) | - |  | Double | Waveforms | IntelliVue MP70 | Waves: PLTHpo
1466 | Arterial Venous O2 Content Difference | - |  | Double | VitalSigns | IntelliVue MP70 | Ca-vO2
1467 | O2 Delivery (DO2) | - |  | Double | VitalSigns | IntelliVue MP70 | DO2
1469 | O2 Extraction Ratio (O2ER) | - |  | Double | VitalSigns | IntelliVue MP70 | O2ER
1471 | O2 Delivery Index (DO2I) | - |  | Double | VitalSigns | IntelliVue MP70 | DO2I
1472 | O2 Consumption Index (VO2I) | - |  | Double | VitalSigns | IntelliVue MP70 | VO2I
1473 | Oxygenation Ratio (PaO2/FiO2) | - | PaO2 to FIO2 ratio expressed in mmHg to % ratio. | Double | VitalSigns | IntelliVue MP70 | PaFIO2
1475 | Physiologic Dead Space | - |  | Double | VitalSigns | IntelliVue MP70 | VENT: Vd
1477 | Alveolar Ventilation Volume (VA) | - |  | Double | VitalSigns | IntelliVue MP70 | VENT: ALVENT
1478 | Urea Nitrogen Concentration (BUN) | - |  | Double | VitalSigns | IntelliVue MP70 | BUN
1479 | Creatinine Concentration (Crea) | - | Measured by the i-Stat module. | Double | VitalSigns | IntelliVue MP70 | Crea
1480 | Plasma Osmolarity (Osm) | - |  | Double | VitalSigns | IntelliVue MP70 | PlOsm
1482 | Urine Creatinine | - |  | Double | VitalSigns | IntelliVue MP70 | UrCrea
1484 | Urine Potassium (K+) | - |  | Double | VitalSigns | IntelliVue MP70 | UrK
1486 | Urine Osmolarity | - |  | Double | VitalSigns | IntelliVue MP70 | UrOsm
1487 | Blood Urea Nitrogen Creatinine Ratio (BUN/cr) | - |  | Double | VitalSigns | IntelliVue MP70 | BUN/cr
1488 | Osmolar Clearance (COSM) | - |  | Double | VitalSigns | IntelliVue MP70 | COsm
1489 | Creatinine Clearance (CreaCl) | - |  | Double | VitalSigns | IntelliVue MP70 | CreaCl
1490 | Fractional Excretion of Sodium (FeNa) | - |  | Double | VitalSigns | IntelliVue MP70 | FeNa
1493 | Urine Serum Creatinine Ratio | - |  | Double | VitalSigns | IntelliVue MP70 | U/SCr
1495 | Urine Potassium Excretion | - |  | Double | VitalSigns | IntelliVue MP70 | UrKEx
1496 | Urine Sodium Excretion | - |  | Double | VitalSigns | IntelliVue MP70 | UrNaEx
1497 | Free Water Clearance (CH2O) | - |  | Double | VitalSigns | IntelliVue MP70 | CH2O
1520 | Inspired N2O | - |  | Double | VitalSigns | IntelliVue MP70 | N2O
1523 | Premature Ventricular Contractions | - |  | Double | VitalSigns | Common | ECG: PVC
1533 | LAP Systolic | - |  | Double | VitalSigns | Common | IBP: LAP
1534 | LAP Diastolic | - |  | Double | VitalSigns | Common | IBP: LAP
1540 | TcGas Heating Power | - |  | Double | DeviceInfo | IntelliVue MP70 | TCUT: HeatPw
1556 | Abdominal Fetal Heart Rate (FHR) | - |  | Double | VitalSigns | IntelliVue MP70 | aFHR
1557 | Abdominal Fetal Beat-to-Beat Rate (FHR) | - |  | Double | VitalSigns | IntelliVue MP70 | baFHR
1558 | Patient Activity Level (Movement) | - |  | Double | VitalSigns | IntelliVue MP70 | mAct
1559 | Patient Activity Level (Wearable) | - |  | Double | VitalSigns | IntelliVue MP70 | wAct
1569 | Minimum Airway Pressure | - |  | Double | VitalSigns | IntelliVue MP70 | AWPmin
1570 | P0.1 Occlusion Pressure | - |  | Double | VitalSigns | IntelliVue MP70 | VENT: Poccl
1573 | Patient CO2 Production | - | Volume of CO2 produced by respiration. | Double | VitalSigns | IntelliVue MP70 | VCO2
1574 | Relative Dead Space | - | Relative Dead Space Volume. | Double | VitalSigns | IntelliVue MP70 | VENT: relVd
1590 | Procalcitonin (PCT) | - |  | Double | VitalSigns | IntelliVue MP70 | PCT
1607 | Fetal Heart Rate (USound) | - | This variable is associated with ChannelID identifier. | Double | VitalSigns | IntelliVue MP70 | FHR/FHR1/FHR2/FHR3
1609 | Fetal Beat-to-Beat Rate (USound) | - | This variable is associated with ChannelID identifier. | Double | VitalSigns | IntelliVue MP70 | btbFHR/bFHR1/bFHR2/bFHR3
1630 | HFV Mean Pressure Setting | - |  | Double | Settings | IntelliVue MP70 | sHFMAP
1631 | HFV Tidal Volume Setting | - |  | Double | Settings | IntelliVue MP70 | sHFVTV
1645 | PAV-based Resistance Setting | - |  | Double | Settings | IntelliVue MP70 | sRppv
1646 | PAV-based Elastance Setting | - |  | Double | Settings | IntelliVue MP70 | sEppv
1703 | Abdominal TOCO | - |  | Double | VitalSigns | IntelliVue MP70 | aToco
1704 | Chest Movement Respiration Rate | - |  | Double | VitalSigns | IntelliVue MP70 | cmRR
1705 | Patient Movement Respiration Rate | - |  | Double | VitalSigns | IntelliVue MP70 | mRR
1706 | Patient Movement Pulse Rate | - |  | Double | VitalSigns | IntelliVue MP70 | mPulse
1707 | Respiration Signal Quality Index (Movement) | - |  | Double | OtherMeasuredData | IntelliVue MP70 | mSQI
1710 | Cardiac Power Index (CCO) | - |  | Double | VitalSigns | IntelliVue MP70 | CPI
1711 | Cardiac Power Output (CCO) | - |  | Double | VitalSigns | IntelliVue MP70 | CPO
1715 | Central Venous Pressure (Thermodilution) | - |  | Double | VitalSigns | IntelliVue MP70 | iCVP
1716 | Mean Arterial Pressure (Thermodilution) | - |  | Double | VitalSigns | IntelliVue MP70 | iARTm
1718 | Average Blood Temperature (Thermodilution) | - |  | Double | VitalSigns | IntelliVue MP70 | iAvgBT
1723 | O2 Delivery DO2 (Thermodilution) | - |  | Double | VitalSigns | IntelliVue MP70 | iDO2
1724 | O2 Consumption (Thermodilution) | - |  | Double | VitalSigns | IntelliVue MP70 | iVO2
1725 | O2 Consumption Index (Thermodilution) | - |  | Double | VitalSigns | IntelliVue MP70 | iVO2I
1733 | Global End-Diastolic Volume (Thermodilution) | - |  | Double | VitalSigns | IntelliVue MP70 | iGEDV
1734 | Global End-Diastolic Volum Index (Thermodilution) | - |  | Double | VitalSigns | IntelliVue MP70 | iGEDVI
1735 | Pulmonary Vascular Permeability Index (Thermodilution) | - |  | Double | VitalSigns | IntelliVue MP70 | iPVPI
1737 | Stroke Volume (Thermodilution) | - |  | Double | VitalSigns | IntelliVue MP70 | iSV
1738 | Stroke Volume Index (Thermodilution) | - |  | Double | VitalSigns | IntelliVue MP70 | iSI
1739 | O2 Consumption Estimate | - |  | Double | VitalSigns | IntelliVue MP70 | iVO2e
1745 | EEG1 Wave | - |  | Double | Waveforms | IntelliVue MP70 | Waves: EEG1
1746 | EEG2 Wave | - |  | Double | Waveforms | IntelliVue MP70 | Waves: EEG2
1747 | EEG3 Wave | - |  | Double | Waveforms | IntelliVue MP70 | Waves: EEG3
1748 | EEG4 Wave | - |  | Double | Waveforms | IntelliVue MP70 | Waves: EEG4
1751 | Spectral Edge Frequency | - |  | Double | VitalSigns | IntelliVue MP70 | EEG: SEF
1757 | Burst Suppression Ratio | - | This variable is associated with ChannelID identifier. | Double | VitalSigns | IntelliVue MP70 | BSR1/BSR2/BSR3/BSR4
1759 | Patient Energy Expenditure | - |  | Double | VitalSigns | IntelliVue MP70 | EE
1760 | Patient Respiratory Quotient | - |  | Double | VitalSigns | IntelliVue MP70 | RQ
1782 | Spectral Edge Frequency (CH3) | - |  | Double | VitalSigns | IntelliVue MP70 | EEG: Frequ3
1784 | Relative Power In Delta Band (CH3) | - |  | Double | VitalSigns | IntelliVue MP70 | Delta3
1785 | Relative Power In Theta Band (CH3) | - |  | Double | VitalSigns | IntelliVue MP70 | Theta3
1786 | Relative Power In Alpha Band (CH3) | - |  | Double | VitalSigns | IntelliVue MP70 | Alpha3
1787 | Relative Power In Beta Band (CH3) | - |  | Double | VitalSigns | IntelliVue MP70 | Beta3
1790 | Spectral Edge Frequency (CH4) | - |  | Double | VitalSigns | IntelliVue MP70 | SEF4
1792 | Relative Power In Delta Band (CH4) | - |  | Double | VitalSigns | IntelliVue MP70 | Delta4
1793 | Relative Power In Theta Band (CH4) | - |  | Double | VitalSigns | IntelliVue MP70 | Theta4
1794 | Relative Power In Alpha (CH4) | - |  | Double | VitalSigns | IntelliVue MP70 | Alpha4
1795 | Relative Power In Beta Band (CH4) | - |  | Double | VitalSigns | IntelliVue MP70 | Beta4
1911 | Device Unique ID | unitless |  | String | DeviceInfo | Common | System Information
1914 | Patient Bed ID | unitless | This variable contains a printable string identifying the system location. | String | PatientInfo | Common | System Information
1929 | Patient ID | unitless |  | String | PatientInfo | Common | Patient Demographics
1930 | Patient Name | unitless |  | String | PatientInfo | Common | Patient Demographics
1942 | Patient Comment Text | unitless |  | String (Multiset) | PatientInfo | Common | Patient Demographics
1955 | Patient Stress Index | - |  | Double | PatientInfo | IntelliVue MP70 | StrssI
1982 | Generic Patient Temperature | - |  | Double | VitalSigns | IntelliVue MP70 | Tbody
2049 | Breath Cycle Type | - |  | Double | DeviceInfo | IntelliVue MP70 | sCycTy
2077 | Airway Gas Flow Rate | - | Measured in the inspiratory path. | Double | VitalSigns | IntelliVue MP70 | VENT: AWFin
2092 | Inspired Minute Volume | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
2104 | Positive End Expiratory Pressure (PEEP) Setting | - |  | Double | Settings | IntelliVue MP70 | sEPAP
2163 | Spectral Edge Frequency - 1st EEG | - |  | Double | VitalSigns | IntelliVue MP70 | EEG: Frequ1
2167 | Bispectral Index | - | This variable is associated with ChannelID identifier. | Double | VitalSigns | IntelliVue MP70 | BIS L/BIS R
2168 | Total Power (Ch1) | - |  | Double | VitalSigns | IntelliVue MP70 | EEG: TP1
2169 | Spike Rate | - | This variable is associated with ChannelID identifier. | Double | VitalSigns | IntelliVue MP70 | BurstL/BurstR
2173 | Spectral Edge Frequency - 2nd EEG | - |  | Double | VitalSigns | IntelliVue MP70 | EEG: Frequ2
2178 | Total Power (Ch2) | - |  | Double | VitalSigns | IntelliVue MP70 | EEG: TP2
2182 | Narcotrend Index | - |  | Double | VitalSigns | IntelliVue MP70 | NI
2183 | Sharp Transient Intensity | - |  | Double | VitalSigns | IntelliVue MP70 | STI
2188 | Total Power Relative | - |  | Double | VitalSigns | IntelliVue MP70 | TPrel
2200 | Device Hardware Revision | unitless | This variable is obsolete because it may contains incorrect information. Please use variables 764, 765 and 4201 must be used instead. | String | DeviceInfo | Common | Obsolete Variables
2201 | Serial Number | unitless | This variable is obsolete because it may contains incorrect information. Please use variables 764, 765 and 4201 must be used instead. | String | DeviceInfo | Common | Obsolete Variables
2246 | O2 Flush Percentage Setting | - |  | Double | Settings | IntelliVue MP70 | sO2max
2250 | CO2 Tidal Production | - |  | Double | VitalSigns | IntelliVue MP70 | VCO2ti
2254 | Inspiratory Resistance | - |  | Double | VitalSigns | IntelliVue MP70 | Rinsp
2255 | Device Date/Time | unitless | Format: 16 ascii digits string of "YYYYMMDDhhmmssff", where the "ff" is second_fractions that is not used by IntelliVue monitors. | String | DeviceInfo | Common | System Information
2259 | Arterial BP Systolic | - |  | Double | VitalSigns | Common | IBP: ABP
2260 | Arterial BP Diastolic | - |  | Double | VitalSigns | Common | IBP: ABP
2261 | Arterial BP Mean | - |  | Double | VitalSigns | Common | IBP: ABP
2262 | Pulmonary Artery BP - Systolic | - |  | Double | VitalSigns | Common | IBP: PAP
2263 | Pulmonary Artery BP - Diastolic | - |  | Double | VitalSigns | Common | IBP: PAP
2264 | Pulmonary Artery BP Mean | - |  | Double | VitalSigns | Common | IBP: PAP
2271 | Right Atrial Pressure (RAP) | - | Available on IntelliVue devices with firmware revision G or later. | Double | VitalSigns | Common | IBP: RAP
2272 | Left Atrial Pressure (LAP) | - | Available on IntelliVue devices with firmware revision G or later. | Double | VitalSigns | Common | IBP: LAP
2273 | Intra Cranial Pressure (ICP) | - | Available on IntelliVue devices with firmware revision G or later. | Double | VitalSigns | Common | IBP: ICP
2317 | Pulmonary Artery Wedge Pressure | - |  | Double | VitalSigns | IntelliVue MP70 | IBP: PAWP
2321 | Generic NIBP Systolic Pressure | - | This variable is same as variable 2, except that the early estimated values are marked as NonSignificant in this variable. | Double | VitalSigns | Common | NBP
2322 | Generic NIBP Diastolic Pressure | - | This variable is same as variable 3, except that the early estimated values are marked as NonSignificant in this variable. | Double | VitalSigns | Common | NBP
2323 | Generic NIBP Mean Pressure | - | This variable is same as variable 4, except that the early estimated values are marked as NonSignificant in this variable. | Double | VitalSigns | Common | NBP
2324 | Inspired Tidal Volume | - |  | Double | VitalSigns | IntelliVue MP70 | TVin
2325 | Expired Tidal Volume | - |  | Double | VitalSigns | IntelliVue MP70 | TVexp
2326 | Expired Minute Volume | - |  | Double | VitalSigns | IntelliVue MP70 | MINVOL
2341 | Patient Blood Temperature | - |  | Double | VitalSigns | IntelliVue MP70 | Tblood
2342 | Injectate Temperature | - |  | Double | OtherMeasuredData | IntelliVue MP70 | Tinj
2343 | Inspired O2 | - |  | Double | VitalSigns | IntelliVue MP70 | VENT: FIO2
2350 | Inspired O2 Pressure | - | This variable is marked as obsolete because the label is incorrect. Please see variable 52523 for concentration and 52527 for partial pressure. | Double | VitalSigns | IntelliVue MP70 | O2
2351 | End Tidal O2 | - |  | Double | VitalSigns | IntelliVue MP70 | O2
2352 | CO2 (Gas Analyzer) | - |  | Double | VitalSigns | IntelliVue MP70 | CO2
2357 | End Tidal N2O | - |  | Double | VitalSigns | IntelliVue MP70 | N2O
2360 | End Tidal Agent Concentration | - |  | Double | VitalSigns | IntelliVue MP70 | AGT
2362 | End Tidal Desflurane | - |  | Double | VitalSigns | IntelliVue MP70 | DES
2364 | End Tidal Sevoflurane | - |  | Double | VitalSigns | IntelliVue MP70 | SEV
2366 | End Tidal Halothane | - |  | Double | VitalSigns | IntelliVue MP70 | HAL
2368 | End Tidal Enflurane | - |  | Double | VitalSigns | IntelliVue MP70 | ENF
2370 | End Tidal Isoflurane | - |  | Double | VitalSigns | IntelliVue MP70 | ISO
2378 | Fresh Gas N2O Flow | - | Mixer parameter. Setting of fresh gas N2O flow. | Double | DeviceInfo | IntelliVue MP70 | sfgN2O
2379 | Fresh Gas Air Flow | - |  | Double | DeviceInfo | IntelliVue MP70 | fgAir
2391 | Transcutaneous O2 Pressure | - |  | Double | VitalSigns | IntelliVue MP70 | TCUT: tcpO2
2392 | Transcutaneous CO2 Pressure | - |  | Double | VitalSigns | IntelliVue MP70 | TCUT: tcpCO2
2399 | Arterial O2 Saturation (SaO2) | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
2408 | EEG Total Power | - |  | Double | VitalSigns | IntelliVue MP70 | EEG: TP
2409 | Bispectral Index (BIS) | - |  | Double | VitalSigns | IntelliVue MP70 | EEG: BIS
2411 | EEG Burst Suppression Ratio | - |  | Double | VitalSigns | IntelliVue MP70 | EEG: BSI
2420 | Spontaneous Alveolar Minute Volume | - |  | Double | VitalSigns | IntelliVue MP70 | SpMVal
2421 | Airway Dead Space | - |  | Double | VitalSigns | IntelliVue MP70 | awVd
2432 | Expired Spontaneous Tidal Volume | - |  | Double | VitalSigns | IntelliVue MP70 | SpTVex
2434 | Spontaneous Expired Minute Volume | - |  | Double | VitalSigns | IntelliVue MP70 | %SpMV
2440 | Total Alveolar Tidal Volume | - |  | Double | VitalSigns | IntelliVue MP70 | TValv
2447 | Respiration Waveform | - |  | Double | Waveforms | IntelliVue MP70 | Waves: cmResp
2456 | EEG Signal Quality Index | - |  | Double | VitalSigns | IntelliVue MP70 | EEG: SQI
2517 | Wearable Respiration Rate | - |  | Double | VitalSigns | IntelliVue MP70 | wRR
2518 | Wearable Pulse Rate | - |  | Double | VitalSigns | IntelliVue MP70 | wPulse
2519 | Wearable Patient Temperature | - |  | Double | VitalSigns | IntelliVue MP70 | wTemp
2520 | Respiration Signal Quality Index (Wearable) | - |  | Double | OtherMeasuredData | IntelliVue MP70 | wSQI
2521 | Pulse Rate Signal Quality Index (Wearable) | - |  | Double | OtherMeasuredData | IntelliVue MP70 | wSQI
2542 | Airway Pressure Wave | - |  | Double | Waveforms | IntelliVue MP70 | Waves: AWP
2543 | Citrate Blood Concentration Setting | - |  | Double | Settings | IntelliVue MP70 | sCitDo
2547 | Neutrophil Count | - |  | Double | VitalSigns | IntelliVue MP70 | Bands
2580 | Fresh Gas Total Flow Setting | - | Mixer parameter. | Double | Settings | IntelliVue MP70 | sfgFl
2583 | I:E Ratio Setting | unitless | This variable is associated with ChannelID identifier. This variable is calculated from variable 2773, which is sent as a string in the format "1:xx.xx". | String | Settings | IntelliVue MP70 | Calculated Variables
2598 | Patient Admission Status | unitless | Possible values: "empty", "pre-admitted", "admitted", and "discharged". | String | PatientInfo | Common | Patient Demographics
2602 | ECG Lead V | - | Physiological identifier: NOM_ECG_AMPL_ST_V | Double | VitalSigns | Common | ECG: ST
2738 | Plateau Pressure | - | This variable contains parameter from Philips label "NLS_NOM_VENT_PRESS_RESP_PLAT" and observed value "NLS_NOM_VENT_PRESS_RESP_PLAT". | Double | VitalSigns | IntelliVue MP70 | VENT: PEinsp
2773 | I:E Ratio Setting | - |  | Double | Settings | IntelliVue MP70 | sIE 1:
2776 | Total PEEP | - |  | Double | VitalSigns | IntelliVue MP70 | VENT: tPEEP
2787 | APRV Low Time (Tlow) | - |  | Double | VitalSigns | IntelliVue MP70 | Tlow
2821 | System General Status | unitless | This variable identifies the current operating mode of the device. Possible values: "not specified", "monitoring", "demo", and "service". Moreover, "- standby" or "- config" can be appended | String | DeviceInfo | Common | System Information
2845 | Frequency | Hz | This variable describes the frequency of the main power supply. Possible values: "unspecified", "50", and "60". | String | DeviceInfo | Common | System Information
2855 | Device Status | unitless | Possible values: "disconnected", "unassociated", "operating". | String | DeviceInfo | Common | System Information
2879 | Device Software Revision | unitless | This variable is obsolete because it may contains incorrect information. Please use variables 764, 765 and 4201 must be used instead. | String | DeviceInfo | Common | Obsolete Variables
2880 | Maximum Apnea Time Setting | - |  | Double | DeviceInfo | IntelliVue MP70 | sTAmax
2889 | Inspired Agent Concentration | - |  | Double | VitalSigns | IntelliVue MP70 | inAGT
2898 | Patient Weight | - |  | Double | PatientInfo | IntelliVue MP70 | Weight
2901 | Patient First Name | unitless |  | String | PatientInfo | Common | Patient Demographics
2907 | Patient Temperature | - |  | Double | VitalSigns | Common | Temp
2915 | Ejection Fraction | - |  | Double | VitalSigns | IntelliVue MP70 | REF
2916 | End Diastolic Volume | - |  | Double | VitalSigns | IntelliVue MP70 | EDV
2917 | End Diastolic Volume Index | - |  | Double | VitalSigns | IntelliVue MP70 | EDVI
2918 | End Systolic Volume | - |  | Double | VitalSigns | IntelliVue MP70 | ESV
2919 | End Systolic Volume Index | - |  | Double | VitalSigns | IntelliVue MP70 | ESVI
2922 | EEG Electromyography (EMG) | - | This variable is associated with ChannelID identifier. | Double | VitalSigns | IntelliVue MP70 | EMG L/EMG R
2923 | Generic EEG Wave | - |  | Double | Waveforms | IntelliVue MP70 | Waves: EEG
2924 | ECG Pad Wave | - |  | Double | Waveforms | IntelliVue MP70 | Waves: Pads
2925 | ECG Paddle Wave | - |  | Double | Waveforms | IntelliVue MP70 | Waves: Paddle
2930 | Unspecific Pressure Pulse Rate | - |  | Double | VitalSigns | IntelliVue MP70 | Pulse
2932 | Patient Temperature 1 | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
2933 | Patient Temperature 2 | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
2954 | APRV High Pressure (Phigh) | - |  | Double | VitalSigns | IntelliVue MP70 | Phigh
2959 | APRV Low Pressure (Plow) | - |  | Double | VitalSigns | IntelliVue MP70 | Plow
2964 | I:E Ratio | unitless | This variable is associated with ChannelID identifier This variable is calculated from variable 1190 and sent as a string in the format "1:xx.xx". | String | VitalSigns | IntelliVue MP70 | Calculated Variables
2968 | Plateau Airway Pressure | - |  | Double | VitalSigns | IntelliVue MP70 | Pplat
2984 | Plateau Pressure Setting | - |  | Double | Settings | IntelliVue MP70 | sPplat
2991 | FiO2 Upper Alarm Limit | - |  | Double | Settings | IntelliVue MP70 | AlarmLimit: highO2
2992 | FiO2 Lower Alarm Limit | - |  | Double | Settings | IntelliVue MP70 | AlarmLimit: lowO2
2995 | Device Altitude | m |  | Double | DeviceInfo | Common | System Information
3003 | Blood Flow Rate Setting | - |  | Double | Settings | IntelliVue MP70 | sBFl
3005 | Dialysate Flow Rate Setting | - |  | Double | Settings | IntelliVue MP70 | sDialF
3007 | Effluent Flow Rate Setting | - |  | Double | Settings | IntelliVue MP70 | sEfflF
3024 | Access Pod Pressure | - |  | Double | OtherMeasuredData | IntelliVue MP70 | AccP
3025 | Return Pod Pressure | - |  | Double | OtherMeasuredData | IntelliVue MP70 | RetP
3026 | Filter Pod Pressure | - |  | Double | OtherMeasuredData | IntelliVue MP70 | FiltP
3027 | Effluent Pod Pressure | - |  | Double | OtherMeasuredData | IntelliVue MP70 | EfflP
3029 | Transmembrane Pressure | - |  | Double | VitalSigns | IntelliVue MP70 | TMP
3062 | Hb Calculation Offset | - |  | Double | VitalSigns | IntelliVue MP70 | OffsHb
3091 | Patient Time Since Last Ward Admission | - |  | Double | PatientInfo | IntelliVue MP70 | TWrdAd
3092 | Patient Total Admission Time | - |  | Double | PatientInfo | IntelliVue MP70 | WardAd
3101 | Patient Drainage Volume | - |  | Double | VitalSigns | IntelliVue MP70 | DrnVol
3168 | Fluid Removal Actual Rate | - |  | Double | VitalSigns | IntelliVue MP70 | Vrem/h
3204 | Sum of Minimum Alveolar Concentration (MAC) | - |  | Double | VitalSigns | IntelliVue MP70 | MAC
3266 | HFV Frequency Setting | - |  | Double | Settings | IntelliVue MP70 | sHFFrq
3268 | Ventilator Base Flow Setting | - |  | Double | Settings | IntelliVue MP70 | sfgFmn
3271 | Blood Pressure Difference | - |  | Double | VitalSigns | IntelliVue MP70 | IBP: DeltaP
3272 | Generic Blood Pressure Difference 1 | - |  | Double | VitalSigns | IntelliVue MP70 | IBP: DeltaP1
3273 | Generic Blood Pressure Difference 2 | - |  | Double | VitalSigns | IntelliVue MP70 | IBP: DeltaP2
3278 | Patient Height | - |  | Double | PatientInfo | IntelliVue MP70 | Height
3291 | Patient Temperature Setting | - |  | Double | Settings | IntelliVue MP70 | sTskin
3385 | Arterial O2 Content (CaO2) | - |  | Double | VitalSigns | IntelliVue MP70 | CaO2
3391 | Alveolar Arterial O2 Difference (AaDO2) | - |  | Double | VitalSigns | IntelliVue MP70 | AaDO2
3400 | Venous O2 Content (CvO2) | - |  | Double | VitalSigns | IntelliVue MP70 | CvO2
3415 | Ionized Calcium Concentration (Ca) | - |  | Double | VitalSigns | IntelliVue MP70 | iCa
3421 | Chloride Concentration (Cl-) | - |  | Double | VitalSigns | IntelliVue MP70 | Cl
3435 | Activated Clotting Time | - | Measured by the i-Stat module. | Double | VitalSigns | IntelliVue MP70 | ACT
3446 | Generic Temperature | - |  | Double | VitalSigns | IntelliVue MP70 | T1
3447 | Generic Temperature 2 | - |  | Double | VitalSigns | IntelliVue MP70 | T2
3448 | Generic Temperature 3 | - |  | Double | VitalSigns | IntelliVue MP70 | T3
3449 | Generic Temperature 4 | - |  | Double | VitalSigns | IntelliVue MP70 | T4
3474 | Expiratory Resistance | - |  | Double | VitalSigns | IntelliVue MP70 | Rexp
3490 | TcGas Site Timer | - |  | Double | OtherMeasuredData | IntelliVue MP70 | TCUT: SitTim
3491 | O2 Saturation (SO2) | - |  | Double | VitalSigns | IntelliVue MP70 | SaO2
3496 | TcGas Transducer Temperature | - |  | Double | DeviceInfo | IntelliVue MP70 | TCUT: SensrT
3502 | Peak Expiratory Flow Setting | - |  | Double | Settings | IntelliVue MP70 | s%PEF
3503 | Auto Flow Active Status | - |  | Double | DeviceInfo | IntelliVue MP70 | sFlCtr
3506 | Inspiratory Minimum CO2 | - |  | Double | VitalSigns | Common | CO2
3507 | Expired Tidal Volume Indexed by Body Weight | - |  | Double | VitalSigns | IntelliVue MP70 | TVe/kg
3555 | ATC Activation Status | - |  | Double | DeviceInfo | IntelliVue MP70 | sTubeC
3570 | Skin Temperature | - |  | Double | VitalSigns | Common | Tskin
3571 | Rectal Temperature | - |  | Double | VitalSigns | Common | Trect
3572 | Nasopharyngeal Temperature | - |  | Double | VitalSigns | Common | Tnaso
3573 | Esophageal Temperature | - |  | Double | VitalSigns | Common | Tesoph
3574 | Core Temperature | - |  | Double | VitalSigns | Common | Tcore
3575 | Arterial Temperature | - |  | Double | VitalSigns | Common | Tart
3576 | Venous Temperature | - |  | Double | VitalSigns | Common | Tven
3583 | Generic Core Temperature 1 | - |  | Double | VitalSigns | IntelliVue MP70 | T1Core
3584 | Generic Core Temperature 2 | - |  | Double | VitalSigns | IntelliVue MP70 | T2Core
3591 | Device Number | unitless | Device Model Number (example: "M8000") | String | DeviceInfo | Common | System Information
3635 | Drug Delivery Rate | - |  | Double | DeviceInfo | IntelliVue MP70 | sDlvRt
3637 | Drug Concentration | - |  | Double | PatientInfo | IntelliVue MP70 | sDConc
3639 | Manufacturers Name | unitless | "Philips" | String | DeviceInfo | Common | System Information
3640 | Pressure Alarm Delay | - |  | Double | DeviceInfo | IntelliVue MP70 | sPADel
3646 | Arterial Blood Pressure (ART) | - | Available on IntelliVue devices with firmware revision G or later. Arterial Blood Pressure (NOM_PRESS_BLD_ART). | Double | VitalSigns | Common | IBP: ART
3647 | Pulmonary Artery Blood Pressure (PA) | - | Available on IntelliVue devices with firmware revision G or later. | Double | VitalSigns | Common | IBP: PAP
3704 | Arterial/Main Pump Flow | - |  | Double | VitalSigns | IntelliVue MP70 | HLMfl
3710 | Slave Pump Flow | - |  | Double | VitalSigns | IntelliVue MP70 | SlvPfl
3715 | Suction Pump Flow | - |  | Double | VitalSigns | IntelliVue MP70 | SucPfl
3720 | Auxiliary Pump Flow | - |  | Double | VitalSigns | IntelliVue MP70 | AuxPfl
3723 | Auxiliary Pump Time Since Start of Current Plegia Session | - |  | Double | VitalSigns | IntelliVue MP70 | AxOnTi
3724 | Auxiliary Pump Time Since Stop of Plegia Session | - |  | Double | VitalSigns | IntelliVue MP70 | AxOffT
3725 | Auxiliary Pump Delivered Volume in Current Plegia Session | - |  | Double | VitalSigns | IntelliVue MP70 | AxDVol
3726 | Auxiliary Pump Total Volume | - |  | Double | VitalSigns | IntelliVue MP70 | AxTVol
3727 | Auxiliary Pump Total Plegia Time | - |  | Double | VitalSigns | IntelliVue MP70 | AxPlTi
3730 | Cardioplegia Main Pump Flow | - |  | Double | VitalSigns | IntelliVue MP70 | PlePfl
3733 | Cardioplegia Main Pump Delivered Volume of Current Plegia Session | - |  | Double | VitalSigns | IntelliVue MP70 | CpDVol
3734 | Cardioplegia Main Pump Total Volume | - |  | Double | VitalSigns | IntelliVue MP70 | CpTVol
3737 | Cardioplegia Slave Pump Flow | - |  | Double | VitalSigns | IntelliVue MP70 | SplPfl
3740 | Cardioplegia Slave Pump Delivered Volume in Current Plegia Session | - |  | Double | VitalSigns | IntelliVue MP70 | CsDVol
3741 | Cardioplegia Slave Pump Total Volume | - |  | Double | VitalSigns | IntelliVue MP70 | CsTVol
3742 | Cycle Flow Sensitivity Setting | - |  | Double | Settings | IntelliVue MP70 | sFlSeC
3761 | Hydrocarbon Plasma Concentration (HCO3) | - |  | Double | VitalSigns | IntelliVue MP70 | HCO3
3781 | Corpuscular Mean Hemoglobin Concentration | - |  | Double | VitalSigns | IntelliVue MP70 | MCHC
3785 | Arterial O2 Partial Pressure | - |  | Double | VitalSigns | IntelliVue MP70 | PaO2
3786 | Venous O2 Partial Pressure | - |  | Double | VitalSigns | IntelliVue MP70 | PvO2
3787 | Arterial CO2 Partial Pressure | - |  | Double | VitalSigns | IntelliVue MP70 | PaCO2
3789 | Arterial O2 Saturation (SaO2) | - | Obsolete variable that may be removed in a future release. The label of this variable is not correct. Please use variable 14 instead. | Double | VitalSigns | Common | Obsolete Variables
3793 | Hematocrit (HCT) | - | Post-Dialysis Hematocrit | Double | VitalSigns | IntelliVue MP70 | cHCTo
3804 | Respiratory Rate (Flow) | - |  | Double | VitalSigns | IntelliVue MP70 | RRawf
3805 | Respiratory Rate (CO2) | - |  | Double | VitalSigns | IntelliVue MP70 | RRCO2
3853 | Patient Type | unitless | Possible values: "unspecified", "adult", "pediatric", and "neonatal". | String | PatientInfo | Common | Patient Demographics
3854 | Patient Pace Maker Status | unitless | Possible values: "not paced", "paced gen". | String | PatientInfo | Common | Patient Demographics
3869 | Heart Rate Signal Quality | - |  | Double | OtherMeasuredData | IntelliVue MP70 | mSQI
3917 | Tracheal Pressure | - |  | Double | VitalSigns | IntelliVue MP70 | Ptrach
3921 | Ventilator Leakage | - |  | Double | VitalSigns | IntelliVue MP70 | Lksys
3961 | Plateau Time Setting | - | This variable is obsolete that may be removed in a future release. The label of this variable is incorrect, please use variable 9585 instead. | Double | Settings | IntelliVue MP70 | Obsolete Variables
3963 | Warmer Mattress Temperature | - |  | Double | OtherMeasuredData | IntelliVue MP70 | Tmatt
3968 | Spontaneous Fraction of Minute Volume | - |  | Double | VitalSigns | IntelliVue MP70 | SpMVe
4004 | Breathing System Leakage | - |  | Double | OtherMeasuredData | IntelliVue MP70 | VENT: Leak
4006 | Inspired MAC Concentration | - |  | Double | VitalSigns | IntelliVue MP70 | MAC
4008 | Inspiratory 2nd Agent | - |  | Double | VitalSigns | IntelliVue MP70 | inAGTs
4009 | Expiratory 2nd Agent | - | This is the device's measurement with label="NLS_NOM_EMFC_etAGTs" and physio_id="NOM_CONC_AWAY_AGENT_ET". | Double | VitalSigns | IntelliVue MP70 | etAGTs
4012 | Volume Assist Level Setting | - |  | Double | Settings | IntelliVue MP70 | sWOBva
4014 | Backup Tidal Volume Setting | - |  | Double | Settings | IntelliVue MP70 | sTVbak
4016 | Flow Acceleration Setting | - |  | Double | Settings | IntelliVue MP70 | sFlacc
4017 | Cumulated Halothane Consumption | - |  | Double | OtherMeasuredData | IntelliVue MP70 | AccHAL
4018 | Cumulated Enflurane Consumption | - |  | Double | OtherMeasuredData | IntelliVue MP70 | AccENF
4019 | Cumulated Isoflurane Consumption | - |  | Double | OtherMeasuredData | IntelliVue MP70 | AccISO
4020 | Cumulated Desflurane Consumption | - |  | Double | OtherMeasuredData | IntelliVue MP70 | AccDes
4021 | Cumulated Sevoflurane Consumption | - |  | Double | OtherMeasuredData | IntelliVue MP70 | AccSEV
4026 | Air Humidity Setting | - |  | Double | Settings | IntelliVue MP70 | sHum
4028 | Mattress Temperature Setting | - |  | Double | Settings | IntelliVue MP70 | sTmatt
4029 | Inspired Agent Setting | - |  | Double | Settings | IntelliVue MP70 | sInAGT
4030 | Expired Agent Setting | - | End-Tidal agent concentration | Double | Settings | IntelliVue MP70 | sEtAGT
4037 | Inspired Desflurane Setting | - |  | Double | Settings | IntelliVue MP70 | sInDES
4038 | Expired Desflurane Setting | - | End-Tidal desflurane concentration | Double | Settings | IntelliVue MP70 | sEtDES
4039 | Inspired Sevoflurane Setting | - |  | Double | Settings | IntelliVue MP70 | sInSEV
4040 | Expired Sevoflurane Setting | - | End-Tidal sevoflurane concentration | Double | Settings | IntelliVue MP70 | sEtSEV
4079 | Fresh Gas Sevoflurane | - |  | Double | VitalSigns | IntelliVue MP70 | fgSEV
4080 | Fresh Gas Desflurane | - |  | Double | VitalSigns | IntelliVue MP70 | fgDES
4081 | Ambient Temperature | - |  | Double | OtherMeasuredData | IntelliVue MP70 | Tamb
4082 | Ambient Pressure | - |  | Double | OtherMeasuredData | IntelliVue MP70 | PB
4084 | Pulse Rate ( Unspecified Source ) | - |  | Double | VitalSigns | Common | Pulse
4089 | Central Venous Pressure Diastolic | - |  | Double | VitalSigns | Common | IBP: CVP
4093 | Central Venous Mean Pressure | - |  | Double | VitalSigns | Common | IBP: CVP
4094 | Central Venous Pressure Systolic | - |  | Double | VitalSigns | Common | IBP: CVP
4120 | Expiratory Volume Wave | - |  | Double | Waveforms | IntelliVue MP70 | Waves: AWVex
4143 | Apnea Ventilation I:E Ratio Setting (Number) | - |  | Double | Settings | IntelliVue MP70 | sBIE1:
4145 | Dynamic Plateau Airway Pressure | - |  | Double | VitalSigns | IntelliVue MP70 | PplDyn
4158 | Transcutaneous Relative Heater Power | - |  | Double | OtherMeasuredData | IntelliVue MP70 | HeatPr
4171 | Spontaneous Respiratory Rate Setting | - |  | Double | Settings | IntelliVue MP70 | sRRmax
4201 | Product Description | unitless | Variables 764, 765 and 4201 must be interpreted conjointly. | String (Multiset) | DeviceInfo | Common | System Information
Consult detailed help for more information on this variable
4226 | Replacement Cumulated Volume | - |  | Double | VitalSigns | IntelliVue MP70 | Vrep/h
4227 | Dialysate Cumulated Volume | - | This variable is associated with ChannelID identifier. | Double | VitalSigns | IntelliVue MP70 | DialC/Vdly/h
4228 | Effluent Cumulated Volume | - | This variable is associated with ChannelID identifier. | Double | VitalSigns | IntelliVue MP70 | Veff/h / EfflC
4229 | Fluid Removal Cumulated Volume | - |  | Double | VitalSigns | IntelliVue MP70 | Vufr
4352 | Base Pressure | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
4369 | PEEP/CPAP Setting | - |  | Double | Settings | IntelliVue MP70 | VENT: sPEEP
4370 | Sigh Rate Setting | - |  | Double | Settings | IntelliVue MP70 | VENT: sSghR
4371 | Sigh Tidal Volume Setting | - |  | Double | Settings | IntelliVue MP70 | VENT: sSghTV
4375 | Data Averaging Interval Setting | - |  | Double | Settings | IntelliVue MP70 | Available only in revF or earlier
4376 | Chart Interval Setting | - |  | Double | Settings | IntelliVue MP70 | Available only in revF or earlier
4383 | Static Lung Compliance | - |  | Double | VitalSigns | IntelliVue MP70 | Cstat
4384 | Static Lung Resistance | - |  | Double | VitalSigns | IntelliVue MP70 | Raw
4385 | Dynamic Lung Compliance | - |  | Double | VitalSigns | IntelliVue MP70 | Cdyn
4386 | Dynamic Lung Resistance | - | NLS_NOM_RES_AWAY_DYN | Double | VitalSigns | IntelliVue MP70 | Rdyn
4388 | Apnea Inspiratory Pressure Setting | - |  | Double | Settings | IntelliVue MP70 | sIPApn/sAInP
4393 | Patient Hose Compliance | - |  | Double | OtherMeasuredData | IntelliVue MP70 | Chos
4401 | Oximeter Averaging Interval Setting | - |  | Double | Settings | IntelliVue MP70 | Available only in revF or earlier
4457 | System Alarm Status | unitless | Overall alert state of the system. | Double | Alarms | Common | Alarms
Consult detailed help for more information on this variable
4458 | Patient Alarm Priority | unitless | Overall Patient alarm priority. This is the current highest priority value of currently active Patient Alarms. | Double | Alarms | Common | Alarms
Consult detailed help for more information on this variable
4459 | Technical Alarm Priority | unitless | Overall Technical alarm priority. This is the current highest priority value of currently active Technical Alarms. | Double | Alarms | Common | Alarms
Consult detailed help for more information on this variable
4460 | Audio Alarm Priority | unitless | Overall audible alarm priority. This is the current highest priority value of currently active audible alarms. | Double | Alarms | Common | Alarms
Consult detailed help for more information on this variable
4461 | Technical Alarm Source ID | unitless |  | Double (Multiset) | Alarms | Common | Alarms
4462 | Technical Alarm Global State | unitless | Consult detailed help for more information on this variable | Double (Multiset) | Alarms | Common | Alarms
4463 | Technical Alarm Priority | unitless | Consult detailed help for more information on this variable | Double (Multiset) | Alarms | Common | Alarms
4464 | Technical Alarm Flag | unitless | Consult detailed help for more information on this variable | Double (Multiset) | Alarms | Common | Alarms
4465 | Technical Alarm Priority | unitless |  | Double (Multiset) | Alarms | Common | Alarms
4466 | Technical Alarm Description | unitless |  | String (Multiset) | Alarms | Common | Alarms
4467 | Patient Alarm Source ID | unitless |  | Double (Multiset) | Alarms | Common | Alarms
4468 | Patient Alarm Global State | unitless | Consult detailed help for more information on this variable | Double (Multiset) | Alarms | Common | Alarms
4469 | Patient Alarm Priority | unitless | Consult detailed help for more information on this variable | Double (Multiset) | Alarms | Common | Alarms
4470 | Patient Alarm Flag | unitless | Consult detailed help for more information on this variable | Double (Multiset) | Alarms | Common | Alarms
4471 | Patient Alarm Priority | unitless |  | Double (Multiset) | Alarms | Common | Alarms
4472 | Patient Alarm Description | unitless |  | String (Multiset) | Alarms | Common | Alarms
4473 | Device Nomenclature Revision | unitless | Format: major.minor (such as 1.0) | String | DeviceInfo | Common | System Information
4474 | Device Text Catalog Revision | unitless | Format: "MM.NN mm.nn", where "MM.NN" is text catalog revision and "mm.nn" is language revision. | String | DeviceInfo | Common | System Information
4475 | Device Data Format | unitless | Always 11 (Unicode). | String | DeviceInfo | Common | System Information
4476 | Device Operating Area | unitless | Possible values: "unspecified", "operating room", "intensive care unit", "neonatal intensive care unit", and "cardiology care unit". | String | DeviceInfo | Common | System Information
4496 | Patient O2 Consumption (VO2) | - |  | Double | VitalSigns | IntelliVue MP70 | VO2
4497 | Patient CO2 Production (VCO2) | - |  | Double | VitalSigns | IntelliVue MP70 | bbvCO2
4507 | Arterial O2 Saturation (Left) | - |  | Double | VitalSigns | IntelliVue MP70 | SpO2 l
4508 | Arterial O2 Saturation (Right) | - |  | Double | VitalSigns | IntelliVue MP70 | SpO2 r
4509 | SpO2 Difference | - |  | Double | VitalSigns | IntelliVue MP70 | DeltaSpO2
4510 | Perfusion Indicator | - |  | Double | VitalSigns | IntelliVue MP70 | Perf
4511 | Relative Perfusion (Left) | - |  | Double | VitalSigns | IntelliVue MP70 | Perf l
4512 | Relative Perfusion (Right) | - |  | Double | VitalSigns | IntelliVue MP70 | Perf r
4513 | Ao Systolic | - |  | Double | VitalSigns | Common | IBP: Ao
4514 | Ao Diastolic | - |  | Double | VitalSigns | Common | IBP: Ao
4515 | Ao Mean | - |  | Double | VitalSigns | Common | IBP: Ao
4516 | RAP Systolic | - |  | Double | VitalSigns | Common | IBP: RAP
4517 | RAP Diastolic | - |  | Double | VitalSigns | Common | IBP: RAP
4518 | RAP Mean | - |  | Double | VitalSigns | Common | IBP: RAP
4519 | Intracranial Pressure ICP Systolic | - |  | Double | VitalSigns | Common | IBP: ICP
4520 | Intracranial Pressure ICP Diastolic | - |  | Double | VitalSigns | Common | IBP: ICP
4521 | ICP Mean | - |  | Double | VitalSigns | Common | IBP: ICP
4522 | UAP Systolic | - |  | Double | VitalSigns | Common | IBP: UAP
4523 | UAP Diastolic | - |  | Double | VitalSigns | Common | IBP: UAP
4524 | UAP Mean | - |  | Double | VitalSigns | Common | IBP: UAP
4525 | UVP Systolic | - |  | Double | VitalSigns | Common | IBP: UVP
4526 | UVP Diastolic | - |  | Double | VitalSigns | Common | IBP: UVP
4527 | UVP Mean | - |  | Double | VitalSigns | Common | IBP: UVP
4528 | Unspecific Pressure Systolic | - |  | Double | VitalSigns | Common | IBP: P
4529 | Unspecific Pressure Diastolic | - |  | Double | VitalSigns | Common | IBP: P
4530 | Unspecific Pressure Mean | - |  | Double | VitalSigns | Common | IBP: P
4531 | Cerebral Perfusion Pressure (CPP) | - |  | Double | VitalSigns | IntelliVue MP70 | IBP: CPP
4532 | Stroke Volume Variation | - |  | Double | VitalSigns | IntelliVue MP70 | SVV
4533 | Index of Left Ventricular Contractility | - | This variable contains measurement from label "NLS_NOM_GRAD_PRESS_BLD_AORT_POS_MAX" and observed value "NOM_GRAD_PRESS_BLD_AORT_POS_MAX". | Double | VitalSigns | IntelliVue MP70 | dPmax
4556 | Aortic Blood Pressure (Ao) Wave | - |  | Double | Waveforms | IntelliVue MP70 | Waves: Ao
4560 | Auxiliary Waveform | - |  | Double | Waveforms | IntelliVue MP70 | Waves: AUX
4561 | Intermittent Mandatory Ventilation Status | - |  | Double | VitalSigns | IntelliVue MP70 | VENT: IMV
4562 | Left Cardiac Work | - |  | Double | VitalSigns | IntelliVue MP70 | LCW
4564 | Right Cardiac Work | - |  | Double | VitalSigns | IntelliVue MP70 | RCW
4565 | Serum Creatinine | - |  | Double | VitalSigns | IntelliVue MP70 | SCrea
4566 | Urine Plasma Osmolarity Ratio | - |  | Double | VitalSigns | IntelliVue MP70 | U/POsm
4567 | Urine Sodium Potassium Ratio | - |  | Double | VitalSigns | IntelliVue MP70 | UrNa/K
4568 | Urine Volume | - |  | Double | VitalSigns | IntelliVue MP70 | UrVol
4574 | Left Cardiac Work Index | - |  | Double | VitalSigns | IntelliVue MP70 | LCWI
4576 | Right Cardiac Work Index | - |  | Double | VitalSigns | IntelliVue MP70 | RCWI
4577 | Light Intensity (SvO2 Module) | - |  | Double | VitalSigns | IntelliVue MP70 | LI
4578 | SpO2 SvO2 Difference | - |  | Double | VitalSigns | IntelliVue MP70 | Sp-vO2
4581 | Normalized Post Replacement Volume | - |  | Double | VitalSigns | IntelliVue MP70 | VdpoNh
4582 | Normalized Effluent Volume | - | This variable is associated with ChannelID identifier. | Double | VitalSigns | IntelliVue MP70 | Veff/EffDn
4584 | Normalized Total Replacement Volume (PBP) | - | This variable is associated with ChannelID identifier. | Double | VitalSigns | IntelliVue MP70 | VrePN/VrePNh
4585 | Normalized Replacement Volume (Pre) | - |  | Double | VitalSigns | IntelliVue MP70 | VdprNh
4586 | Normalized Ultrafiltration Volume | - | This variable is associated with ChannelID identifier. | Double | VitalSigns | IntelliVue MP70 | VufrN/VufrNh
4587 | Patient Fluid Balance Limit Setting | - |  | Double | Settings | IntelliVue MP70 | sEPFL
4588 | Total Replacement Volume (PBP) | - | This variable is associated with ChannelID identifier. | Double | VitalSigns | IntelliVue MP70 | VreP/VreP/h
4589 | Cumulated Liquid Volume (Pre) | - | This variable is associated with ChannelID identifier. | Double | VitalSigns | IntelliVue MP70 | Vdpr/Vdpr/h
4593 | Filtration Fraction Setting | - |  | Double | Settings | IntelliVue MP70 | sFufrP
4616 | Diff.-Temperature between Tblood and Tskin | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
4617 | Diff.-Temperature between Tblood and Trect | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
4618 | Diff.-Temperature between Tblood and Tnaso | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
4619 | Diff.-Temperature between Tblood and Tesop | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
4620 | Diff.-Temperature between Tblood and Tcore | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
4621 | Diff.-Temperature between Tblood and Tart | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
4622 | Diff.-Temperature between Tblood and Tven | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
4623 | Diff.-Temperature between Tskin and Trect | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
4624 | Diff.-Temperature between Tskin and Tnaso | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
4625 | Diff.-Temperature between Tskin and Tesop | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
4626 | Diff.-Temperature between Tskin and Tcore | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
4627 | Diff.-Temperature between Tskin and Tart | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
4628 | Diff.-Temperature between Tskin and Tven | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
4629 | Diff.-Temperature between Trect and Tnaso | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
4630 | Diff.-Temperature between Trect and Tesop | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
4631 | Diff.-Temperature between Trect and Tcore | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
4632 | Diff.-Temperature between Trect and Tart | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
4633 | Diff.-Temperature between Trect and Tven | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
4634 | Diff.-Temperature between Tnaso and Tesop | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
4635 | Diff.-Temperature between Tnaso and Tcore | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
4636 | Diff.-Temperature between Tnaso and Tart | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
4637 | Diff.-Temperature between Tnaso and Tven | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
4638 | Diff.-Temperature between Tesop and Tcore | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
4639 | Diff.-Temperature between Tesop and Tart | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
4640 | Diff.-Temperature between Tesop and Tven | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
4641 | Diff.-Temperature between Tcore and Tart | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
4642 | Diff.-Temperature between Tcore and Tven | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
4643 | Diff.-Temperature between Tart and Tven | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
4644 | Newborn Body Length | - |  | Double | VitalSigns | IntelliVue MP70 | Length
4646 | Newborn Gestational Age | - |  | Double | PatientInfo | IntelliVue MP70 | G.Age
4647 | Calories | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
4648 | Input/Output Totals | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
4649 | Newborn Apgar Scores | - |  | Double | PatientInfo | IntelliVue MP70 | Available only in revF or earlier
4650 | BSA Formula Boyd | - |  | Double | PatientInfo | IntelliVue MP70 | BSA(B)
4651 | BSA Formula Dubois | - |  | Double | PatientInfo | IntelliVue MP70 | BSA(D)
4653 | Pupil Reaction (Left Eye) | - | Light reaction of left eye's pupil. | Double | VitalSigns | IntelliVue MP70 | PRL
4654 | Pupil Reaction (Right Eye) | - | Light reaction of right eye's pupil. | Double | VitalSigns | IntelliVue MP70 | PRR
4661 | Pleth Gain | - |  | Double | DeviceInfo | IntelliVue MP70 | PlGain
4662 | Airway Temperature | - |  | Double | VitalSigns | IntelliVue MP70 | Tairwy
4663 | Ventilation Perfusion Index | - |  | Double | VitalSigns | IntelliVue MP70 | VENT: VQI
4664 | O2 Extraction Index | - |  | Double | VitalSigns | IntelliVue MP70 | O2EI
4665 | Duration Above Base Pressure | - |  | Double | VitalSigns | IntelliVue MP70 | VENT: DPosP
4666 | Inspiration Rise Time | - |  | Double | VitalSigns | IntelliVue MP70 | VENT: RiseTi
4667 | Continuous Positive Airway Pressure (CPAP) | - |  | Double | VitalSigns | IntelliVue MP70 | CPAP
4668 | HFV Amplitude | - |  | Double | VitalSigns | IntelliVue MP70 | VENT: HFVAmp
4669 | Urine Urea | - |  | Double | VitalSigns | IntelliVue MP70 | UrUrea
4670 | Urine pH Value | - |  | Double | VitalSigns | IntelliVue MP70 | UrpH
4671 | Total of CO2 (tCO2) | - | Result of Blood gas Analysis. | Double | VitalSigns | IntelliVue MP70 | tCO2
4672 | Total Bilirubin | - |  | Double | VitalSigns | IntelliVue MP70 | tBili
4673 | Serum Glucose | - |  | Double | VitalSigns | IntelliVue MP70 | SerGlu
4674 | Urine Glucose | - |  | Double | VitalSigns | IntelliVue MP70 | UrGlu
4675 | Direct Bilirubin | - |  | Double | VitalSigns | IntelliVue MP70 | dBili
4676 | Serum Calcium | - |  | Double | VitalSigns | IntelliVue MP70 | SerCa
4677 | Total Serum Calcium | - |  | Double | VitalSigns | IntelliVue MP70 | tSerCa
4678 | Serum Magnesium | - |  | Double | VitalSigns | IntelliVue MP70 | SerMg
4679 | Serum Phosphat | - |  | Double | VitalSigns | IntelliVue MP70 | SerPho
4680 | Serum Kalium (Potassium) | - |  | Double | VitalSigns | IntelliVue MP70 | SerK
4681 | Serum Clorid | - |  | Double | VitalSigns | IntelliVue MP70 | SerCl
4682 | Serum Albumine | - |  | Double | VitalSigns | IntelliVue MP70 | SerAlb
4683 | Urine Clorid | - |  | Double | VitalSigns | IntelliVue MP70 | UrCl
4684 | Serum Globulin | - |  | Double | VitalSigns | IntelliVue MP70 | SerGlo
4685 | Serum Protein | - |  | Double | VitalSigns | IntelliVue MP70 | SerPro
4686 | Serum Urea | - |  | Double | VitalSigns | IntelliVue MP70 | SrUrea
4687 | Partial Thromboplastin Time | - |  | Double | VitalSigns | IntelliVue MP70 | PTT
4688 | Prothrombin Time | - |  | Double | VitalSigns | IntelliVue MP70 | PT
4689 | Thrombin Time | - |  | Double | VitalSigns | IntelliVue MP70 | TT
4690 | Thromboplastine Time (Quick Time) | - |  | Double | VitalSigns | IntelliVue MP70 | Quick
4691 | Alkalische Phosphatase | - |  | Double | VitalSigns | IntelliVue MP70 | AP
4692 | Alpha Amylase | - |  | Double | VitalSigns | IntelliVue MP70 | alphaA
4693 | Cholesterinesterase | - |  | Double | VitalSigns | IntelliVue MP70 | CHE
4694 | Creatinin Kinase | - |  | Double | VitalSigns | IntelliVue MP70 | SerCK
4695 | Creatine Kinase of Type Muscle-Brain | - |  | Double | VitalSigns | IntelliVue MP70 | CK-MB
4696 | Creatine Kinase of Type Muscle | - |  | Double | VitalSigns | IntelliVue MP70 | CK-MM
4697 | Gamma Glutamyl Transpeptidase | - |  | Double | VitalSigns | IntelliVue MP70 | GGT
4698 | Glutamic Oxaloacetic Transaminase | - |  | Double | VitalSigns | IntelliVue MP70 | GOT
4699 | Glutamic Pyruvic Transaminase | - |  | Double | VitalSigns | IntelliVue MP70 | GPT
4700 | Cholesterin | - |  | Double | VitalSigns | IntelliVue MP70 | Chol
4701 | Triglyzeride | - |  | Double | VitalSigns | IntelliVue MP70 | TGL
4702 | Urine Protein | - | (Total) Protein in Urine | Double | VitalSigns | IntelliVue MP70 | UrPro
4703 | Urine Calcium | - |  | Double | VitalSigns | IntelliVue MP70 | UrCa
4704 | Carboxy Hemoglobin | - |  | Double | VitalSigns | IntelliVue MP70 | CO-Hb
4705 | Total Protein | - |  | Double | VitalSigns | IntelliVue MP70 | tPro
4706 | Lactate Dehydrogenase | - |  | Double | VitalSigns | IntelliVue MP70 | LDH
4707 | Aspartin Aminotransferase | - |  | Double | VitalSigns | IntelliVue MP70 | AST
4708 | Alveolarproteinose (Alkaline Phosphatase) | - |  | Double | VitalSigns | IntelliVue MP70 | ALP
4709 | Reticulocyte Count | - |  | Double | VitalSigns | IntelliVue MP70 | RC
4710 | Coagulation Time | - |  | Double | VitalSigns | IntelliVue MP70 | CT
4711 | Erithrocyte Sedimentation Rate | - |  | Double | VitalSigns | IntelliVue MP70 | ESR
4712 | Packed Cell Volume | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
4713 | Kaolin Cephalin Time | - |  | Double | VitalSigns | IntelliVue MP70 | KCT
4715 | Minimum Pressure During Breathing Cycle | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
4716 | Infusion Pump Accumulated Volume | - |  | Double | OtherMeasuredData | IntelliVue MP70 | AccVol
4717 | Transthoracic Impedance | - |  | Double | VitalSigns | Common | T.I.
4718 | Inspired Expired N2O Difference | - |  | Double | VitalSigns | IntelliVue MP70 | VENT: i-eN2O
4719 | Inspired Expired Halothane Difference | - |  | Double | VitalSigns | IntelliVue MP70 | VENT: i-eHAL
4720 | Inspired Expired Enfluran Difference | - |  | Double | VitalSigns | IntelliVue MP70 | VENT: i-eENF
4721 | Inspired Expired Isofluran Difference | - |  | Double | VitalSigns | IntelliVue MP70 | VENT: i-eISO
4722 | Inspired Expired Sevofluran Difference | - |  | Double | VitalSigns | IntelliVue MP70 | VENT: i-eSEV
4723 | Inspired Expired Desfluran Difference | - |  | Double | VitalSigns | IntelliVue MP70 | VENT: i-eDES
4724 | Inspired Expired Agent Difference | - |  | Double | VitalSigns | IntelliVue MP70 | VENT: i-eAGT
4725 | Inspired Expired O2 Difference | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
4726 | Patient Circuit O2 | - | O2 measured in the Patient Circuit | Double | VitalSigns | IntelliVue MP70 | VENT: cktO2
4727 | Barometric Pressure in Sample Chamber | - | This variable is obsolete that may be removed in a future release. It is duplicated with variable 4082. | Double | OtherMeasuredData | IntelliVue MP70 | Obsolete Variables
4728 | Minute Volume | - |  | Double | VitalSigns | IntelliVue MP70 | MINVOL
4729 | Mean Airway Pressure | - |  | Double | VitalSigns | IntelliVue MP70 | MnAwP
4730 | Airway Respiration Rate | - |  | Double | VitalSigns | Common | awRR
4731 | Inspiratory Airway Pressure Waveform | - |  | Double | Waveforms | IntelliVue MP70 | Waves: AWPin
4732 | Inspiratory Flow Waveform | - |  | Double | Waveforms | IntelliVue MP70 | Waves: AWFin
4733 | HFV Inspired Mandatory Minute Volume | - |  | Double | VitalSigns | IntelliVue MP70 | HFMVin
4734 | HFV Gas Transport Coefficient Value | - |  | Double | VitalSigns | IntelliVue MP70 | DCO2
4735 | HFV Inspired Tidal Volume | - |  | Double | VitalSigns | IntelliVue MP70 | VENT: HFTVin
4736 | HFV Tidal Volume | - |  | Double | VitalSigns | IntelliVue MP70 | VENT: HFVTV
4738 | Generic Heart Rate | - | Heart Rate received from an external device | Double | VitalSigns | IntelliVue MP70 | extHR
4739 | PERF from the 2nd SpO2/PLETH module | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
4741 | Percent Alpha (Left) | - |  | Double | VitalSigns | IntelliVue MP70 | EEG: LT %AL
4742 | Percent Beta (Left) | - |  | Double | VitalSigns | IntelliVue MP70 | EEG: LT %BE
4745 | Absolute Alpha (Left) | - |  | Double | VitalSigns | IntelliVue MP70 | EEG: LT AL
4746 | Absolute Beta (Left) | - |  | Double | VitalSigns | IntelliVue MP70 | EEG: LT BE
4747 | Absolute Delta (Left) | - |  | Double | VitalSigns | IntelliVue MP70 | EEG: LT DL
4748 | Absolute Theta (Left) | - |  | Double | VitalSigns | IntelliVue MP70 | EEG: LT TH
4749 | Mean Dominant Frequency (Left) | - |  | Double | VitalSigns | IntelliVue MP70 | EEG: LT MDF
4750 | Median Power Frequency (Left) | - |  | Double | VitalSigns | IntelliVue MP70 | EEG: LT MPF
4751 | Peak Power Frequency (Left) | - |  | Double | VitalSigns | IntelliVue MP70 | EEG: LT PPF
4752 | Spectral Edge Frequency (Left) | - |  | Double | VitalSigns | IntelliVue MP70 | EEG: LT SEF
4753 | Total Power (Left) | - |  | Double | VitalSigns | IntelliVue MP70 | EEG: LT TP
4754 | Scale of the Left Channel EEG Wave | - |  | Double | VitalSigns | IntelliVue MP70 | EEG: LSCALE
4755 | Percent Alpha (Right) | - |  | Double | VitalSigns | IntelliVue MP70 | EEG: RT %AL
4756 | Percent Beta (Right) | - |  | Double | VitalSigns | IntelliVue MP70 | EEG: RT %BE
4759 | Absolute Alpha (Right) | - |  | Double | VitalSigns | IntelliVue MP70 | EEG: RT AL
4760 | Absolute Beta (Right) | - |  | Double | VitalSigns | IntelliVue MP70 | EEG: RT BE
4761 | Absolute Delta (Right) | - |  | Double | VitalSigns | IntelliVue MP70 | EEG: RT DL
4762 | Absolute Theta (Right) | - |  | Double | VitalSigns | IntelliVue MP70 | EEG: RT TH
4763 | Mean Dominant Frequency (Right) | - |  | Double | VitalSigns | IntelliVue MP70 | EEG: RT MDF
4764 | Median Power Frequency (Right) | - |  | Double | VitalSigns | IntelliVue MP70 | EEG: RT MPF
4765 | Peak Power Frequency (Right) | - |  | Double | VitalSigns | IntelliVue MP70 | EEG: RT PPF
4766 | Spectral Edge Frequency (Right) | - |  | Double | VitalSigns | IntelliVue MP70 | EEG: RT SEF
4767 | Total Power (Right) | - |  | Double | VitalSigns | IntelliVue MP70 | EEG: RT TP
4768 | Scale of the Right Channel EEG Wave | - |  | Double | VitalSigns | IntelliVue MP70 | EEG: RSCALE
4769 | Duration of Positive Pressure | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
4772 | Fresh Gas N2O Concentration | - |  | Double | VitalSigns | IntelliVue MP70 | fgN2O
4774 | Liquid Level (Anesthetic Agent Bottle) | - |  | Double | OtherMeasuredData | IntelliVue MP70 | AGTLev
4775 | Liquid Level (ISOflurane Bottle) | - |  | Double | OtherMeasuredData | IntelliVue MP70 | ISOLev
4776 | Liquid Level (ENFlurane Bottle) | - |  | Double | OtherMeasuredData | IntelliVue MP70 | ENFLev
4777 | Liquid Level (HALothane Bottle) | - |  | Double | OtherMeasuredData | IntelliVue MP70 | HALLev
4778 | Liquid Level (DESflurane Bottle) | - |  | Double | OtherMeasuredData | IntelliVue MP70 | DESLev
4779 | Liquid Level (SEVoflurane Bottle) | - |  | Double | OtherMeasuredData | IntelliVue MP70 | SEVLev
4780 | Invasive Blood Pressure Wave | - |  | Double | Waveforms | IntelliVue MP70 | Waves: P
4781 | Urine Shift Volume | - |  | Double | VitalSigns | IntelliVue MP70 | UrVSht
4782 | Urine Flow Rate | - |  | Double | VitalSigns | IntelliVue MP70 | UrFl
4786 | Mean Dominant Frequency (CH1) | - |  | Double | VitalSigns | IntelliVue MP70 | EEG: Frequ1
4787 | Mean Dominant Frequency (CH2) | - |  | Double | VitalSigns | IntelliVue MP70 | EEG: Frequ2
4788 | Peak Power Frequency (CH1) | - |  | Double | VitalSigns | IntelliVue MP70 | EEG: Frequ1
4789 | Peak Power Frequency (CH2) | - |  | Double | VitalSigns | IntelliVue MP70 | EEG: Frequ2
4794 | % of Power in the Alpha Band (CH1) | - |  | Double | VitalSigns | IntelliVue MP70 | EEG: Prcnt1
4795 | % of Power in the Alpha Band (CH2) | - |  | Double | VitalSigns | IntelliVue MP70 | EEG: Prcnt2
4796 | % of Power in the Beta Band (1st EEG) | - |  | Double | VitalSigns | IntelliVue MP70 | EEG: Prcnt1
4797 | % of Power in the Beta Band (2nd EEG) | - |  | Double | VitalSigns | IntelliVue MP70 | EEG: Prcnt2
4800 | Hemoglobin | - |  | Double | VitalSigns | IntelliVue MP70 | Hb
4801 | Patient Temperature | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
4802 | Capillaries pH Value | - |  | Double | VitalSigns | IntelliVue MP70 | pHc
4803 | Adjusted pH (at Patient Temperature) | - | Adjusted pH at Patient Temperature | Double | VitalSigns | IntelliVue MP70 | &pH
4804 | Adjusted Capillaries pH Value | - |  | Double | VitalSigns | IntelliVue MP70 | &pHc
4805 | Partial O2 in the Capillaries | - |  | Double | VitalSigns | IntelliVue MP70 | PcO2
4806 | Adjusted PO2 (at Patient Temperature) | - | Adjusted PO2 at Patient Temperature | Double | VitalSigns | IntelliVue MP70 | &PO2
4807 | Adjusted PcO2 (at Patient Temperature) | - |  | Double | VitalSigns | IntelliVue MP70 | &PcO2
4808 | Venous CO2 Partial Pressure (PvCO2) | - |  | Double | VitalSigns | IntelliVue MP70 | PvCO2
4809 | Partial CO2 in the Capillaries | - |  | Double | VitalSigns | IntelliVue MP70 | PcCO2
4811 | Adjusted PcCO2 (at Patient Temperature) | - |  | Double | VitalSigns | IntelliVue MP70 | &PcCO2
4812 | Calculated Total CO2 (tCO2) | - |  | Double | VitalSigns | IntelliVue MP70 | CALC: 'tCO2
4813 | Calculated SO2 | - | Calculated SaO2 | Double | VitalSigns | IntelliVue MP70 | CALC: 'SaO2
4815 | Calculated SvO2 | - |  | Double | VitalSigns | IntelliVue MP70 | CALC: 'SvO2
4816 | Calculated ScO2 | - |  | Double | VitalSigns | IntelliVue MP70 | CALC: 'ScO2
4817 | Calculated HCO3 | - |  | Double | VitalSigns | IntelliVue MP70 | CALC: HCO3
4818 | Base Excess of Extra Cellular Fluid | - |  | Double | VitalSigns | IntelliVue MP70 | CALC: Beecf
4819 | Calculated Anion Gap | - |  | Double | VitalSigns | IntelliVue MP70 | CALC: AnGap
4821 | Urea | - | Measured by the i-Stat module. | Double | VitalSigns | IntelliVue MP70 | Urea
4822 | Manually Entered FiO2 | - | In the Blood Review window. | Double | Settings | IntelliVue MP70 | Available only in revF or earlier
4823 | Calculated Arterial Blood Base Excess | - |  | Double | VitalSigns | IntelliVue MP70 | BE,B
4824 | Ionized Magnesium | - |  | Double | VitalSigns | IntelliVue MP70 | iMg
4826 | Desired Module Temperature | - | It's the selected/adjusted Temp in i-Stat Window. | Double | Settings | IntelliVue MP70 | Available only in revF or earlier
4827 | Calculated BUN/Creatinine Ratio | - | Calculated by the i-Stat module. | Double | VitalSigns | IntelliVue MP70 | B/Cre
4828 | Calculated Urea/Creatinine Ratio | - | Calculated by the i-Stat module. | Double | VitalSigns | IntelliVue MP70 | U/Cre
4829 | Lactate | - | Measured value by the i-Stat module. | Double | VitalSigns | IntelliVue MP70 | Lact
4830 | Device Simulator Status | - | Device data with the values: Passed and Failed. | Double | DeviceInfo | IntelliVue MP70 | Available only in revF or earlier
4835 | Pressure of Sensor 1 | - | Pressure value of sensor 1. Device specific data. | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
4836 | Pressure of Sensor 2 | - | Pressure value of sensor 2. Device specific data. | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
4837 | Module Temperature | - |  | Double | OtherMeasuredData | IntelliVue MP70 | Available only in revF or earlier
4843 | Air Temperature | - |  | Double | OtherMeasuredData | IntelliVue MP70 | Available only in revF or earlier
4844 | Micro Environment Air Temperature Setting | - |  | Double | Settings | IntelliVue MP70 | sTemp
4847 | Weight of the Urine Disposable Bag | - |  | Double | VitalSigns | IntelliVue MP70 | BagWgt
4848 | Total Urine Volume of the Current Measurement Period | - |  | Double | VitalSigns | IntelliVue MP70 | tUrVol
4849 | Urine Fluid Density | - |  | Double | VitalSigns | IntelliVue MP70 | UrDens
4850 | Urine Fluid Temperature | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
4854 | Ventilation Activated Status | - | Parameter to Informs whether the Patient is Ventilated | Double | VitalSigns | IntelliVue MP70 | VENT: PtVent
4856 | Simplified Acuity Physiologic Score | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
4857 | Value Normal | - | Missing parameters in SAPS parameter set which are considered in normal range. | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
4858 | Aspartate Aminotransferase | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
4859 | Spontaneous Respiration Rate | - |  | Double | VitalSigns | IntelliVue MP70 | SpRR
4864 | Current Fluid in Urine Bag | - |  | Double | VitalSigns | IntelliVue MP70 | BagVol
4865 | Cardiac Function Index (Thermodilution) | - |  | Double | VitalSigns | IntelliVue MP70 | iCFI
4866 | Intrathoracic Blood Volume (Thermodilution) | - |  | Double | VitalSigns | IntelliVue MP70 | iITVB
4867 | Intrathoracic Blood Volume Index (Thermodilution) | - |  | Double | VitalSigns | IntelliVue MP70 | iITBVI
4868 | Extravascular Lung Water (Thermodilution) | - |  | Double | VitalSigns | IntelliVue MP70 | iEVLW
4869 | Extravascular Lung Water Index (Thermodilution) | - |  | Double | VitalSigns | IntelliVue MP70 | iEVLWI
4873 | NMT TOF Count | - |  | Double | VitalSigns | IntelliVue MP70 | NMT: TOFcnt
4874 | NMT TOF Ratio | - |  | Double | VitalSigns | IntelliVue MP70 | NMT: TOFrat
4875 | Twitch Height of the 1Hz/0.1Hz Stimulation Response | - |  | Double | VitalSigns | IntelliVue MP70 | NMT: Twitch
4876 | Post Tetatic Count | - |  | Double | VitalSigns | IntelliVue MP70 | NMT: PTC
4877 | Remaining Time to Next Stimulation | - |  | Double | VitalSigns | IntelliVue MP70 | NMT: RemTi
4878 | TOF Response TOF1 | - |  | Double | VitalSigns | IntelliVue MP70 | NMT: TOF1
4879 | TOF Response TOF2 | - |  | Double | VitalSigns | IntelliVue MP70 | NMT: TOF2
4880 | TOF Response TOF3 | - |  | Double | VitalSigns | IntelliVue MP70 | NMT: TOF3
4881 | TOF Response TOF4 | - |  | Double | VitalSigns | IntelliVue MP70 | NMT: TOF4
4883 | Activated Partial Thromboplastin Time (aPTT Whole Blood Time) | - |  | Double | VitalSigns | IntelliVue MP70 | aPTTWB
4884 | Activated Partial Thromboplastin Time (aPTT Plasma Equivalent Time) | - |  | Double | VitalSigns | IntelliVue MP70 | aPTTPE
4886 | Prothrombin Time (Whole Blood Time) | - |  | Double | VitalSigns | IntelliVue MP70 | PT WB
4887 | Prothrombin Time (Plasma Equivalent Time) | - |  | Double | VitalSigns | IntelliVue MP70 | PT PE
4888 | Prothrombin Time Ratio | - |  | Double | VitalSigns | IntelliVue MP70 | PTrat
4889 | Spontaneous Inspiration Time | - |  | Double | VitalSigns | IntelliVue MP70 | InsTi
4890 | Overdistension Index | - |  | Double | VitalSigns | IntelliVue MP70 | C20/C
4891 | Lung Time Constant | - |  | Double | VitalSigns | IntelliVue MP70 | TC
4892 | Airway Correlation Coefficient | - |  | Double | VitalSigns | IntelliVue MP70 | r
4893 | Rate Volume Ratio | - |  | Double | VitalSigns | IntelliVue MP70 | RVrat
4894 | Normalized Ionized Calcium (Ca) | - |  | Double | VitalSigns | IntelliVue MP70 | iCa(N)
4956 | Peak Inspiratory Flow Setting | - |  | Double | Settings | IntelliVue MP70 | s%PIF
4957 | Trigger Sensitivity Setting | - |  | Double | Settings | IntelliVue MP70 | Available only in revF or earlier
4958 | Plateau Time Setting | - |  | Double | Settings | IntelliVue MP70 | sPltTi
4959 | Multiple Sigh Number Setting | - |  | Double | Settings | IntelliVue MP70 | VENT: sSghNr
4960 | Apnea Tidal Volume Setting | - |  | Double | Settings | IntelliVue MP70 | sATV
4961 | Apnea Respiration Rate Setting | - |  | Double | Settings | IntelliVue MP70 | sARR
4962 | Apnea Peak Flow Setting | - |  | Double | Settings | IntelliVue MP70 | sAPkFl
4963 | Apnea Inspired O2 Setting | - |  | Double | Settings | IntelliVue MP70 | VENT: sAFIO2
4964 | PSV Pressure Setting | - | Pressure Support Ventilation Setting | Double | Settings | IntelliVue MP70 | VENT: sPSV
4966 | Sigh Activation Status | - |  | Double | DeviceInfo | IntelliVue MP70 | VENT: sEnSgh
4967 | Suction O2 Concentration Setting | - |  | Double | Settings | IntelliVue MP70 | VENT: sO2Suc
4968 | PCV Control Pressure Setting | - |  | Double | Settings | IntelliVue MP70 | VENT: sPin
4969 | PCV Inspiratory Time Setting | - |  | Double | Settings | IntelliVue MP70 | VENT: sPVinT
4970 | PCV Apnea Respiration Rate Setting | - |  | Double | Settings | IntelliVue MP70 | VENT: sAPVRR
4971 | PCV Apnea Inspiratory Time Setting | - |  | Double | Settings | IntelliVue MP70 | VENT: sAPVTi
4972 | PCV Apnea O2 Setting | - |  | Double | Settings | IntelliVue MP70 | VENT: sAPVO2
4973 | PCV Apnea High Airway Pressure Setting | - |  | Double | Settings | IntelliVue MP70 | VENT: sAPVhP
4975 | Inspiratory Component of I:E Ratio Setting | - | Pressure Ventilation I component of I:E Ratio setting | Double | Settings | IntelliVue MP70 | sPVI
4976 | Expiratory Component of I:E Ratio Setting | - | Pressure Ventilation E component of I:E Ratio setting | Double | Settings | IntelliVue MP70 | sPVE
4977 | Inspiratory Component of Apnea Ventilation I:E Ratio Setting | - |  | Double | Settings | IntelliVue MP70 | sAPVI
4978 | Expiratory Component of Apnea Ventilation I:E Ratio Setting | - |  | Double | Settings | IntelliVue MP70 | sAPVE
4979 | Setting: Pressure Ventilation Control | - |  | Double | Settings | IntelliVue MP70 | Available only in revF or earlier
4981 | Cycle Time Setting | - |  | Double | Settings | IntelliVue MP70 | sCycTi
4982 | IPPV Ventilation Respiration Rate Setting | - |  | Double | Settings | IntelliVue MP70 | VENT: sIPPV
4983 | IMV Ventilation Respiration Rate Setting | - |  | Double | Settings | IntelliVue MP70 | VENT: sIMV
4984 | Pressure Support PEEP Setting | - |  | Double | Settings | IntelliVue MP70 | VENT: sSPEEP
4987 | Apnea Ventilation Delay | - |  | Double | Settings | IntelliVue MP70 | Available only in revF or earlier
4989 | O2 Monitoring Mode | - |  | Double | DeviceInfo | IntelliVue MP70 | VENT: sO2Mon
4990 | O2 Calibration Mode | - |  | Double | DeviceInfo | IntelliVue MP70 | VENT: sO2Cal
4993 | Inspiratory Time Setting | - |  | Double | Settings | IntelliVue MP70 | VENT: sInsTi
4994 | Expiratory Time Setting | - |  | Double | Settings | IntelliVue MP70 | VENT: sExpTi
4996 | Alarm Percentage on Rise Time Setting | - |  | Double | Settings | IntelliVue MP70 | VENT: sALMRT
4997 | Continuous Positive Airway Pressure (CPAP) Setting | - | This is the device's measurement with label="NLS_NOM_SETT_CONC_AWAY_O2" and physio_id="NOM_CONC_AWAY_O2". | Double | Settings | IntelliVue MP70 | sCPAP
4998 | Ventilator Flow Setting | - |  | Double | Settings | IntelliVue MP70 | VENT: sFlow
5000 | Low Inspiratory Pressure Setting | - |  | Double | Settings | IntelliVue MP70 | sPmin
5001 | HFV Flow Setting | - |  | Double | Settings | IntelliVue MP70 | sHFVFl
5002 | HFV Respiration Rate Setting | - |  | Double | Settings | IntelliVue MP70 | sHFVRR
5005 | Mechanical Ventilation Activated Status | - | Controlled Mechanical Ventilation Setting | Double | DeviceInfo | IntelliVue MP70 | VENT: sCMV
5006 | SIMV Mode | - | Synchronized Intermittent Mandatory Ventilation Setting | Double | DeviceInfo | IntelliVue MP70 | VENT: sSIMV
5007 | Mandatory Inspired Minute Volume Setting | - |  | Double | Settings | IntelliVue MP70 | VENT: sMMV
5009 | Infusion Pump Delivery Rate Setting | - |  | Double | Settings | IntelliVue MP70 | sDRate
5011 | Airway Respiration Rate Setting | - |  | Double | Settings | IntelliVue MP70 | sAWRR
5012 | Trigger Flow/Volume Setting | - | This is the device's measurement with label="NLS_NOM_SETT_VENT_VOL_LUNG_TRAPD" and physio_id="NOM_SETT_VENT_VOL_LUNG_TRAPD". | Double | Settings | IntelliVue MP70 | VENT: sTrVol
5013 | Apnea Alarm Delay | - |  | Double | DeviceInfo | IntelliVue MP70 | sAADel
5014 | HFV Amplitude Setting | - |  | Double | Settings | IntelliVue MP70 | sHFVAm
5015 | Volume Alarm Delay | - |  | Double | Settings | IntelliVue MP70 | sMVDel
5016 | Pressure Increase Rate Setting | - |  | Double | Settings | IntelliVue MP70 | VENT: sPincR
5017 | Volume Warning Setting | - | Delivered by the Evita 2 VueLink Driver. | Double | Settings | IntelliVue MP70 | VENT: sVmax
5026 | Airway Inspiratory Pressure Lower Alarm Limit | - |  | Double | Settings | IntelliVue MP70 | AlarmLimit: loPmax
5028 | Apnea Alarm Delay Setting | - |  | Double | Settings | IntelliVue MP70 | Available only in revF or earlier
5029 | Pressure Ramp Setting (ASB Ramp) | - |  | Double | Settings | IntelliVue MP70 | Available only in revF or earlier
5030 | Applied Tidal Volume Setting | - |  | Double | Settings | IntelliVue MP70 | sTVap
5031 | Assist Pressure Sensitivity Setting | - | Assist Sensitivity. Used by the Bear 1000 ventilator. | Double | Settings | IntelliVue MP70 | sSens
5032 | High Inspiratory Pressure Setting | - |  | Double | Settings | IntelliVue MP70 | Available only in revF or earlier
5033 | Background Flow Setting | - |  | Double | Settings | IntelliVue MP70 | VENT: sBkgFl
5035 | Vaporizer Agent Concentration Setting | - |  | Double | Settings | IntelliVue MP70 | sAGT
5036 | Vaporizer Isoflurane Concentration Setting | - |  | Double | Settings | IntelliVue MP70 | sISO
5037 | Vaporizer Enflurane Concentration Setting | - |  | Double | Settings | IntelliVue MP70 | sENF
5038 | Vaporizer Halothane Concentration Setting | - |  | Double | Settings | IntelliVue MP70 | sHAL
5039 | Vaporizer Desflurane Concentration Setting | - |  | Double | Settings | IntelliVue MP70 | sDES
5040 | Vaporizer Sevoflurane Concentration Setting | - |  | Double | Settings | IntelliVue MP70 | sSEV
5041 | Gas Measurement Sensor Position | - |  | Double | DeviceInfo | IntelliVue MP70 | VENT: sGasPr
5042 | O2 Measurement Sensor Position | - |  | Double | DeviceInfo | IntelliVue MP70 | VENT: sO2Pr
5044 | Inspired Tidal Volume Setting | - |  | Double | Settings | IntelliVue MP70 | sTVin
5045 | Incubator/Warmer Temperature Setting | - |  | Double | Settings | IntelliVue MP70
5046 | Preset Period of UrVol Setting | - |  | Double | Settings | IntelliVue MP70 | sUrTi
5050 | Low Time Setting | - | Setting variable for Evita 4 Airway Pressure Release Ventilation Mode. | Double | Settings | IntelliVue MP70 | VENT: sTlow
5051 | High Time Setting | - | Setting variable for Evita 4 Airway Pressure Release Ventilation Mode. | Double | Settings | IntelliVue MP70 | VENT: sThigh
5052 | Low Pressure Setting | - | Setting variable for Evita 4 Airway Pressure Release Ventilation Mode. | Double | Settings | IntelliVue MP70 | VENT: sPlow
5053 | High Pressure Setting | - | Setting variable for Evita 4 Airway Pressure Release Ventilation Mode. | Double | Settings | IntelliVue MP70 | VENT: sPhigh
5054 | Volume Assist Level Setting | - |  | Double | Settings | IntelliVue MP70 | VENT: sVolas
5055 | CPAP Flow Assist Level Setting | - |  | Double | Settings | IntelliVue MP70 | VENT: sFlas
5056 | Preset Stimulation Current Setting | - |  | Double | Settings | IntelliVue MP70 | sCurnt
5057 | Preset Stimulation Charge Setting | - |  | Double | Settings | IntelliVue MP70 | sChrge
5058 | Preset Stimulation Impulse Duration Setting | - |  | Double | Settings | IntelliVue MP70 | sPulsD
5059 | NMT TOF Repetition Time | - | Preset Train of Four (Slow TOF) Repetition Time Setting | Double | DeviceInfo | IntelliVue MP70 | NMT: sRepTi
5060 | Panting Limit Setting | - |  | Double | Settings | IntelliVue MP70 | VENT: sfmax
5065 | Airway Inspiratory Pressure Upper Alarm Limit | - | Alarm Limit: High Pressure | Double | Settings | IntelliVue MP70 | AlarmLimit: highP
5066 | Sigh Volume | - |  | Double | VitalSigns | IntelliVue MP70 | Vsigh
5069 | Sustained Pressure Alarm Limit | - |  | Double | Settings | IntelliVue MP70 | AlarmLimit: sustP
5072 | Intrathoracic Blood Volume (ITBV) | - |  | Double | VitalSigns | IntelliVue MP70 | ITBV
5073 | Intrathoracic Blood Volume Index (ITBI) | - |  | Double | VitalSigns | IntelliVue MP70 | ITBVI
5074 | Extravascular Lung Water | - |  | Double | VitalSigns | IntelliVue MP70 | EVLW
5075 | Extravascular Lung Water Index | - |  | Double | VitalSigns | IntelliVue MP70 | EVLWI
5076 | Global End Diastolic Volume | - |  | Double | VitalSigns | IntelliVue MP70 | GEDV
5077 | Global End Diastolic Volume Index | - |  | Double | VitalSigns | IntelliVue MP70 | GEDVI
5078 | Cardiac Function Index | - |  | Double | VitalSigns | IntelliVue MP70 | CFI
5079 | Mixed Venous O2 Saturation | - |  | Double | VitalSigns | IntelliVue MP70 | SvO2
5080 | Generic Term for the Transcutaneous Gases | - |  | Double | VitalSigns | IntelliVue MP70 | TCUT: tcGas
5083 | Intrinsic PEEP | - |  | Double | VitalSigns | IntelliVue MP70 | iPEEP
5086 | Positive Peak Inspiratory Pressure | - |  | Double | VitalSigns | IntelliVue MP70 | PIP
5087 | Dead Space to Tidal Volume Ratio | - |  | Double | VitalSigns | IntelliVue MP70 | Vd/Vt
5088 | Lung Volume Trapped | - |  | Double | VitalSigns | IntelliVue MP70 | VENT: TrpVol
5089 | Lung Vital Capacity | - |  | Double | VitalSigns | IntelliVue MP70 | VC
5090 | Generic Lung Compliance | - |  | Double | VitalSigns | IntelliVue MP70 | COMP
5092 | EEG Electromyography | - |  | Double | VitalSigns | IntelliVue MP70 | EEG: EMG
5093 | Suppression Ratio | - |  | Double | VitalSigns | IntelliVue MP70 | EEG: SR
5095 | EEG Peak Power Frequency | - |  | Double | VitalSigns | IntelliVue MP70 | EEG: PPF
5101 | End Tidal N2 | - |  | Double | VitalSigns | IntelliVue MP70 | N2
5102 | Inspired N2 | - |  | Double | VitalSigns | IntelliVue MP70 | N2
5109 | Generic Agent Wave | - |  | Double | Waveforms | IntelliVue MP70 | Waves: AGT
5110 | End Tidal MAC Concentration | - |  | Double | VitalSigns | IntelliVue MP70 | MAC
5111 | Left Ventricular Stroke Volume Index | - |  | Double | VitalSigns | IntelliVue MP70 | LVSWI
5112 | Left Ventricular Stroke Volume | - |  | Double | VitalSigns | IntelliVue MP70 | LVSW
5113 | Right Ventricular Stroke Volume | - |  | Double | VitalSigns | IntelliVue MP70 | RVSW
5114 | Glasgow Coma Score | - |  | Double | VitalSigns | IntelliVue MP70 | GCS
5115 | SubScore of the GCS - Eye Response | - |  | Double | VitalSigns | IntelliVue MP70 | EyeRsp
5116 | SubScore of the GCS - Motoric Response | - |  | Double | VitalSigns | IntelliVue MP70 | MotRsp
5117 | SubScore of the GCS - Verbal Response | - |  | Double | VitalSigns | IntelliVue MP70 | VblRsp
5118 | Head Circumference | - |  | Double | VitalSigns | IntelliVue MP70 | HC
5119 | pH in Arterial Blood | - |  | Double | VitalSigns | IntelliVue MP70 | Ph
5120 | Arterial Blood Hemoglobin (HB) | - |  | Double | VitalSigns | IntelliVue MP70 | Hb
5121 | pH in Venous Blood | - |  | Double | VitalSigns | IntelliVue MP70 | pHv
5122 | Urine Natrium | - |  | Double | VitalSigns | IntelliVue MP70 | UrNa
5123 | Natrium in Serum | - |  | Double | VitalSigns | IntelliVue MP70 | SerNa
5124 | pH in the Blood Plasma | - |  | Double | VitalSigns | IntelliVue MP70 | pH
5125 | Blood Natrium (Sodium) | - |  | Double | VitalSigns | IntelliVue MP70 | Na
5126 | Kalium (Potassium) | - |  | Double | VitalSigns | IntelliVue MP70 | K
5127 | Generic Glucose Parameter | - |  | Double | VitalSigns | IntelliVue MP70 | Glu
5128 | Partial CO2 | - |  | Double | VitalSigns | IntelliVue MP70 | PCO2
5129 | Partial O2 | - |  | Double | VitalSigns | IntelliVue MP70 | PO2
5130 | Base Excess of Blood | - |  | Double | VitalSigns | IntelliVue MP70 | BE
5131 | Percent Alveolarvenous Shunt Qs/Qt | - |  | Double | VitalSigns | IntelliVue MP70 | Qs/Qt
5132 | Extravascular Thermal Volume Index | - |  | Double | VitalSigns | IntelliVue MP70 | ETVI
5133 | VPB Rate | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
5134 | ECG Lead V1 Wave | - |  | Double | Waveforms | IntelliVue MP70 | Waves: V1
5135 | ECG Lead V2 Wave | - |  | Double | Waveforms | IntelliVue MP70 | Waves: V2
5136 | ECG Lead V3 Wave | - |  | Double | Waveforms | IntelliVue MP70 | Waves: V3
5137 | ECG Lead V4 Wave | - |  | Double | Waveforms | IntelliVue MP70 | Waves: V4
5138 | ECG Lead V5 Wave | - |  | Double | Waveforms | IntelliVue MP70 | Waves: V5
5139 | ECG Lead V6 Wave | - |  | Double | Waveforms | IntelliVue MP70 | Waves: V6
5140 | ECG Lead I Wave | - |  | Double | Waveforms | IntelliVue MP70 | Waves: I
5141 | ECG Lead II Wave | - |  | Double | Waveforms | IntelliVue MP70 | Waves: II
5142 | ECG Lead III Wave | - |  | Double | Waveforms | IntelliVue MP70 | Waves: III
5143 | ECG Lead aVR Wave | - |  | Double | Waveforms | IntelliVue MP70 | Waves: aVR
5144 | ECG Lead aVL Wave | - |  | Double | Waveforms | IntelliVue MP70 | Waves: aVL
5145 | ECG Lead aVF Wave | - |  | Double | Waveforms | IntelliVue MP70 | Waves: aVF
5185 | Mean Dominant Frequency (CH3) | - |  | Double | VitalSigns | IntelliVue MP70 | MDF3
5186 | Mean Dominant Frequency (CH4) | - |  | Double | VitalSigns | IntelliVue MP70 | MDF4
5189 | NMT DBS Count | - |  | Double | VitalSigns | IntelliVue MP70 | DBScnt
5191 | Peak Power Frequency (CH3) | - |  | Double | VitalSigns | IntelliVue MP70 | PPF3
5192 | Peak Power Frequency (CH4) | - |  | Double | VitalSigns | IntelliVue MP70 | PPF4
5497 | Pump Infusion Mode | - |  | Double | DeviceInfo | IntelliVue MP70 | sPuMod
5507 | Remaining Time Counter | - | Time to Elapse Counter used by the i-Stat module. | Double | OtherMeasuredData | IntelliVue MP70 | Elapse
5598 | APRV Low Pressure Synchronisation Time Setting | - |  | Double | Settings | IntelliVue MP70 | sTloS
5603 | APRV High Pressure Synchronisation Time Setting | - |  | Double | Settings | IntelliVue MP70 | sThiS
5614 | Acoustic Respiration Rate | - |  | Double | VitalSigns | IntelliVue MP70 | RRac
5616 | Regional O2 Saturation (rSO2 2) | - |  | Double | VitalSigns | Common
5617 | Regional O2 Saturation (rSO2 3) | - |  | Double | VitalSigns | Common
5620 | Regional O2 Saturation (rSO2 4) | - |  | Double | VitalSigns | Common
5632 | Regional O2 Saturation (rSO2 5) | - |  | Double | VitalSigns | IntelliVue MP70 | SO2 5
5633 | Regional O2 Saturation (rSO2 6) | - |  | Double | VitalSigns | IntelliVue MP70 | SO2 6
5637 | O2 Delivery Index (Thermodilution) | - |  | Double | VitalSigns | IntelliVue MP70 | iDO2I
5659 | Noninvasive Venous Hemoglobin (SpHbv) | - |  | Double | VitalSigns | IntelliVue MP70 | SpHbv
5660 | Noninvasive Arterial Hemoglobin (SpHba) | - |  | Double | VitalSigns | IntelliVue MP70 | SpHba
5661 | Generic Pulse Rate | - |  | Double | VitalSigns | IntelliVue MP70 | iAvgPR
5680 | Corrected Total Hemoglobin | - |  | Double | VitalSigns | IntelliVue MP70 | cSpHb
5681 | Corrected Arterial Total Hemoglobin | - |  | Double | VitalSigns | IntelliVue MP70 | cSpHb
5682 | Corrected Venous Total Hemoglobin | - |  | Double | VitalSigns | IntelliVue MP70 | cSpHbv
5694 | Systemic Vascular Resistance (Thermodilution) | - |  | Double | VitalSigns | IntelliVue MP70 | iSVR
5695 | Systemic Vascular Resistance Index (Thermodilution) | - |  | Double | VitalSigns | IntelliVue MP70 | iSVRI
5707 | Pulse Rate (Abdominal) | - |  | Double | VitalSigns | IntelliVue MP70 | Pulse
5719 | ECG Lead V3R | - |  | Double | VitalSigns | IntelliVue MP70 | ECG: ST
5720 | ECG Lead V4R | - |  | Double | VitalSigns | IntelliVue MP70 | ECG: ST
5721 | ECG Lead V5R | - |  | Double | VitalSigns | IntelliVue MP70 | ECG: ST
5763 | Inspired NO | - |  | Double | VitalSigns | IntelliVue MP70 | NO
5767 | ECG Lead V9 | - |  | Double | VitalSigns | IntelliVue MP70 | ECG: ST
5774 | ECG Lead V7 | - |  | Double | VitalSigns | IntelliVue MP70 | ECG: ST
5775 | ECG Lead V8 | - |  | Double | VitalSigns | IntelliVue MP70 | ECG: ST
5807 | Apnea Ventilation I:E Ratio Setting | unitless | This variable is associated with ChannelID identifier. This variable is calculated from variable 4977 and 4978, which is sent as a string in the format "1:xx.xx". | String | Settings | IntelliVue MP70 | Calculated Variables
5817 | Assisted Respiration Rate | - |  | Double | VitalSigns | IntelliVue MP70 | Rrasst
5818 | Correlation Coefficient Squared | - |  | Double | VitalSigns | IntelliVue MP70 | r2
5837 | ECG Lead V6R | - |  | Double | VitalSigns | IntelliVue MP70 | ECG: ST
5873 | Citrate Solution Concentration | - |  | Double | Settings | IntelliVue MP70 | sCitCo
5874 | Calcium Solution Concentration | - |  | Double | Settings | IntelliVue MP70 | sCConc
5886 | Citrate Acid Concentration | - |  | Double | Settings | IntelliVue MP70 | sCSCAC
5887 | Citrate Load Setting | - |  | Double | Settings | IntelliVue MP70 | sCitLo
5949 | Backup Respiration Rate Setting | - |  | Double | Settings | IntelliVue MP70 | sRRbak
5958 | Inspired Expired Pressure Difference | - |  | Double | VitalSigns | IntelliVue MP70 | sPi-e
5959 | Backup Inspired Expired Pressure Difference Setting | - |  | Double | Settings | IntelliVue MP70 | sPi-eB
6124 | Femoral Arterial Pressure Mean | - |  | Double | VitalSigns | IntelliVue MP70 | IBP: FAP
6125 | Femoral Arterial Pressure Systolic | - |  | Double | VitalSigns | IntelliVue MP70 | IBP: FAP
6126 | Femoral Arterial Pressure Diastolic | - |  | Double | VitalSigns | IntelliVue MP70 | IBP: FAP
6281 | Tidal Volume Lower Alarm Limit | - |  | Double | Settings | IntelliVue MP70 | AlarmLimit: lowTV
6283 | Minute Volume Lower Alarm Limit | - |  | Double | Settings | IntelliVue MP70 | AlarmLimit: lowMV
6322 | ECG Lead V7 Wave | - |  | Double | Waveforms | IntelliVue MP70 | Waves: V7
6323 | ECG Lead V8 Wave | - |  | Double | Waveforms | IntelliVue MP70 | Waves: V8
6324 | ECG Lead V9 Wave | - |  | Double | Waveforms | IntelliVue MP70 | Waves: V9
6326 | Airway Aux Wave | - |  | Double | Waveforms | IntelliVue MP70 | Waves: AWPx
6331 | Fetal DECG Wave | - |  | Double | Waveforms | IntelliVue MP70 | Waves: DECG
6339 | Patient BSA Formula | unitless | Consult detailed help for more information on this variable | Double | PatientInfo | Common | Patient Demographics
6343 | Platelets (Thrombocyte Count) | - |  | Double | VitalSigns | IntelliVue MP70 | Plts
6388 | Endotracheal Tube Type | - |  | Double | DeviceInfo | IntelliVue MP70 | sTube
6390 | Minute Volume Upper Alarm Limit | - |  | Double | Settings | IntelliVue MP70 | AlarmLimit: highMV
6394 | Pulse Rate from Plethysmogram | - |  | Double | VitalSigns | Common | Pulse
6396 | ECG Ectopic Status | - | Consult detailed help for more information on this variable | Double | VitalSigns | Common | Ectopic Status (on IntelliVue)
6399 | ST Lead ES | - |  | Double | VitalSigns | IntelliVue MP70 | ECG: ST
6400 | ST Lead AS | - |  | Double | VitalSigns | IntelliVue MP70 | ECG: ST
6401 | ST Lead AI | - |  | Double | VitalSigns | IntelliVue MP70 | ECG: ST
6402 | ST Index | - |  | Double | VitalSigns | IntelliVue MP70 | ECG: Stindx
6409 | Pulse Pressure | - |  | Double | VitalSigns | IntelliVue MP70 | PP
6410 | Mean Transit Time (MTt) | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
6450 | Alveolar Dead Space | - |  | Double | VitalSigns | IntelliVue MP70 | Vdalv
6451 | Tympanic Temperature | - |  | Double | VitalSigns | IntelliVue MP70 | Ttymp
6452 | Injectate Volume (CO) | - |  | Double | VitalSigns | IntelliVue MP70 | InjVol
6453 | SpO2 (Telemetry) | - |  | Double | VitalSigns | IntelliVue MP70 | %SpO2T
6454 | Pulse Rate (Telemetry) | - |  | Double | VitalSigns | IntelliVue MP70 | PulseT
6455 | Intracranial Pressure ICP1 (Systolic) | - |  | Double | VitalSigns | IntelliVue MP70 | IBP: IC2
6456 | Intracranial Pressure ICP1 (Diastolic) | - |  | Double | VitalSigns | IntelliVue MP70 | IBP: IC1
6457 | Intracranial Pressure ICP1 (Mean) | - |  | Double | VitalSigns | IntelliVue MP70 | IBP: IC3
6458 | Intracranial Pressure ICP2 Systolic | - |  | Double | VitalSigns | IntelliVue MP70 | IBP: IC2
6459 | Intracranial Pressure ICP2 Diastolic | - |  | Double | VitalSigns | IntelliVue MP70 | IBP: IC2
6460 | Intracranial Pressure ICP2 Mean | - |  | Double | VitalSigns | IntelliVue MP70 | IBP: IC2
6461 | Brachial Arterial Pressure (BA) Systolic | - |  | Double | VitalSigns | IntelliVue MP70 | IBP: BAP
6462 | Brachial Arterial Pressure (BA) Diastolic | - |  | Double | VitalSigns | IntelliVue MP70 | IBP: BAP
6463 | Brachial Arterial Pressure (BA) Mean | - |  | Double | VitalSigns | IntelliVue MP70 | IBP: BAP
6464 | Vesical Temperature (Bladder) | - |  | Double | VitalSigns | IntelliVue MP70 | Tvesic
6465 | Cerebral Temperature | - |  | Double | VitalSigns | IntelliVue MP70 | Tcereb
6466 | ST Level for Lead I | - |  | Double | VitalSigns | IntelliVue MP70 | ECG: Rf-I
6467 | ST Level for Lead II | - |  | Double | VitalSigns | IntelliVue MP70 | ECG: Rf-II
6468 | ST Level for Lead III | - |  | Double | VitalSigns | IntelliVue MP70 | ECG: Rf-III
6469 | ST Level for Lead aVR | - |  | Double | VitalSigns | IntelliVue MP70 | ECG: Rf-aVR
6470 | ST Level for Lead aVL | - |  | Double | VitalSigns | IntelliVue MP70 | ECG: Rf-aVL
6471 | ST Level for Lead aVF | - |  | Double | VitalSigns | IntelliVue MP70 | ECG: Rf-aVF
6472 | ST Level for Lead V1 | - |  | Double | VitalSigns | IntelliVue MP70 | ECG: Rf-V1
6473 | ST Level for Lead V2 | - |  | Double | VitalSigns | IntelliVue MP70 | ECG: Rf-V2
6474 | ST Level for Lead V3 | - |  | Double | VitalSigns | IntelliVue MP70 | ECG: Rf-V3
6475 | ST Level for Lead V4 | - |  | Double | VitalSigns | IntelliVue MP70 | ECG: Rf-V4
6476 | ST Level for Lead V5 | - |  | Double | VitalSigns | IntelliVue MP70 | ECG: Rf-V5
6477 | ST Level for Lead V6 | - |  | Double | VitalSigns | IntelliVue MP70 | ECG: Rf-V6
6478 | Prothrombin Time (International Normalized Ratio) | - |  | Double | VitalSigns | IntelliVue MP70 | PT INR
6479 | PSV Tidal Volume | - | Tidal Volume in Pressure Support Ventilation Mode | Double | VitalSigns | IntelliVue MP70 | TVPSV
6480 | Negative Inspiratory Force Index | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
6481 | A-Line ARX Index | - |  | Double | VitalSigns | IntelliVue MP70 | EEG: AAI
6484 | Impedance Cardiography (ICG) Wave | - |  | Double | Waveforms | IntelliVue MP70 | Waves: ICG
6485 | Thoracic Fluid Content | - |  | Double | VitalSigns | IntelliVue MP70 | TFC
6486 | Thoracic Fluid Content Index | - |  | Double | VitalSigns | IntelliVue MP70 | TFI
6487 | Accelerated Cardiac Index | - |  | Double | VitalSigns | IntelliVue MP70 | ACI
6488 | Heart Contractility Index | - |  | Double | VitalSigns | IntelliVue MP70 | HI
6489 | O2 Saturation (Left) | - |  | Double | VitalSigns | IntelliVue MP70 | SO2 l
6490 | O2 Saturation (Right) | - |  | Double | VitalSigns | IntelliVue MP70 | SO2 r
6491 | Pulse Pressure Variation | - |  | Double | VitalSigns | IntelliVue MP70 | PPV
6492 | Downslope Time | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
6493 | Pulmonary Vascular Permeability Index | - | This is the device's measurement with label="NLS_NOM_PERM_VASC_PULM_INDEX" and physio_id="NOM_PERM_VASC_PULM_INDEX". See also variable 8791. | Double | VitalSigns | IntelliVue MP70 | PVPI
6502 | ART Pulse Rate | - | ART pulse rate | Double | VitalSigns | IntelliVue MP70 | Pulse
6543 | Umbilical Arterial Pressure Pulse Rate | - |  | Double | VitalSigns | IntelliVue MP70 | Pulse
6544 | Patient Visit Number | unitless | Patient encounter ID contains the ID of the current visit of the patient | String | PatientInfo | IntelliVue MP70 | Patient Demographics
6555 | Patient Height at Entered Time | - |  | Double | PatientInfo | Common | Patient Demographics
6557 | Patient Weight at Entered Time | - |  | Double | PatientInfo | Common | Patient Demographics
6567 | Patient Body Surface Area (BSA) | - |  | Double | PatientInfo | Common | Patient Demographics
6685 | Inspired O2 Setting (FiO2) | - |  | Double | Settings | IntelliVue MP70 | VENT: sFIO2
6702 | Inspiratory Component of I:E Ratio Setting | - |  | Double | Settings | IntelliVue MP70 | sI
6703 | Expiratory Component of I:E Ratio Setting | - |  | Double | Settings | IntelliVue MP70 | sE
6735 | Lung Elastance | - |  | Double | VitalSigns | IntelliVue MP70 | E
6739 | Backup Inspiratory Time Setting | - |  | Double | Settings | IntelliVue MP70 | sInTiB
6744 | Pressure Support Level Above Phigh | - |  | Double | Settings | IntelliVue MP70 | sdPin
6746 | Total Respiration Duration Setting | - |  | Double | Settings | IntelliVue MP70 | sBreaT
6772 | Heater Power | - |  | Double | DeviceInfo | IntelliVue MP70 | sPower
6861 | Capillary Refill Time | - |  | Double | VitalSigns | IntelliVue MP70 | CRT
6935 | Estimated Patient Circuit Leakage | - |  | Double | OtherMeasuredData | IntelliVue MP70 | LkPat
6936 | Estimated Ventilation Total Leakage | - |  | Double | OtherMeasuredData | IntelliVue MP70 | LkTot
6967 | Ventilation Mode Additional Information | - |  | Double | DeviceInfo | IntelliVue MP70 | sVSupp
6995 | Backup Plateau Time Setting | - |  | Double | Settings | IntelliVue MP70 | sPltTA
7000 | ECG Lead V3R Wave | - |  | Double | Waveforms | IntelliVue MP70 | Waves: V3R
7001 | ECG Lead V4R Wave | - |  | Double | Waveforms | IntelliVue MP70 | Waves: V4R
7002 | ECG Lead V5R Wave | - |  | Double | Waveforms | IntelliVue MP70 | Waves: V5R
7003 | ECG Lead V6R Wave | - |  | Double | Waveforms | IntelliVue MP70 | Waves: V6R
7019 | Signal Quality Index (BIS) | - | This variable is associated with ChannelID identifier. | Double | OtherMeasuredData | IntelliVue MP70 | SQI L/SQI R
7106 | Continuous NIBP Systolic | - |  | Double | VitalSigns | IntelliVue MP70 | cNBP SYS
7107 | Continuous NIBP Diastolic | - |  | Double | VitalSigns | IntelliVue MP70 | cNBP DIA
7108 | Continuous NIBP Mean | - |  | Double | VitalSigns | IntelliVue MP70 | cNBP MEAN
7118 | Next Intervention Time | - |  | Double | OtherMeasuredData | IntelliVue MP70 | InTtvTi
7315 | Total Power (CH3) | - |  | Double | VitalSigns | IntelliVue MP70 | TP3
7316 | Total Power (CH4) | - |  | Double | VitalSigns | IntelliVue MP70 | TP4
7335 | Total Power (Left) | - |  | Double | VitalSigns | IntelliVue MP70 | TP L
7336 | Total Power (Right) | - |  | Double | VitalSigns | IntelliVue MP70 | TP R
7414 | PCV I:E Ratio Setting | unitless | This variable is associated with ChannelID identifier. This variable is calculated from variable 4975 and 4976, which is sent as a string in the format "1:xx.xx". | String | Settings | IntelliVue MP70 | Calculated Variables
7434 | Inspiratory CO2 (FiCO2) | - |  | Double | VitalSigns | IntelliVue MP70 | VENT: FICO2
7532 | End Tidal NO2 | - |  | Double | VitalSigns | IntelliVue MP70 | NO2
7535 | Fresh Gas Total Flow | - |  | Double | OtherMeasuredData | IntelliVue MP70 | fgFlow
7545 | Bolus Volume Setting | - |  | Double | Settings | IntelliVue MP70 | sVBol
7557 | Generic NIBP Pressure | - | Available on IntelliVue devices with firmware revision G or later. | Double | VitalSigns | Common | NBP
7587 | Endotracheal Tube Diameter | - |  | Double | DeviceInfo | IntelliVue MP70 | sTubeD
7640 | Mandatory Respiration Rate | - |  | Double | VitalSigns | IntelliVue MP70 | VENT: MRR
7641 | Mandatory Minute Volume | - |  | Double | VitalSigns | IntelliVue MP70 | VENT: MMV
7642 | Mandatory Tidal Volume | - |  | Double | VitalSigns | IntelliVue MP70 | VENT: MTV
7644 | N2O Consumption | - |  | Double | OtherMeasuredData | IntelliVue MP70 | AccN2O
7765 | Device Power Up Time | s | Number of seconds since the Monitor device's last power-on. | Double | DeviceInfo | Common | System Information
7768 | O2 Consumption | - |  | Double | VitalSigns | IntelliVue MP70 | AccO2
7770 | Air Consumption | - |  | Double | VitalSigns | IntelliVue MP70 | AccAir
7773 | Arterial Blood Pressure (ABP) | - | Available on IntelliVue devices with firmware revision G or later. Arterial Blood Pressure (NOM_PRESS_BLD_ART_ABP). | Double | VitalSigns | Common | IBP: ABP
7777 | ABP Pulse Rate | - | ABP pulse rate | Double | VitalSigns | IntelliVue MP70 | Pulse
7797 | Signal Noise Ratio (CCO) | - |  | Double | VitalSigns | IntelliVue MP70 | SNR
7803 | Spontaneous Airway Respiration Rate | - |  | Double | VitalSigns | IntelliVue MP70 | SpAWRR
7804 | Synchronized Respiration Rate | - |  | Double | VitalSigns | IntelliVue MP70 | RRsync
7805 | Mechanical Respiration Rate | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
7806 | Daily Urine Output | - |  | Double | VitalSigns | IntelliVue MP70 | U/O
7809 | Activated Partial Thromboplastin Time Ratio | - |  | Double | VitalSigns | IntelliVue MP70 | PTTrat
7810 | Global Ejection Fraction | - | This is the device's measurement with label="NLS_NOM_FRACT_EJECT" and physio_id="NOM_FRACT_EJECT". See also 8790. | Double | VitalSigns | IntelliVue MP70 | GEF
7811 | Generic CO2 Parameter | - | This variable contains "PECO2, Partial expired CO2", which is measured from label="NLS_NOM_VENT_AWAY_CO2_EXP", physio_id="NOM_VENT_AWAY_CO2_EXP". | Double | VitalSigns | IntelliVue MP70 | VENT: PECO2
7823 | Central Venous O2 Saturation (ScvO2) | - |  | Double | VitalSigns | IntelliVue MP70 | ScvO2
7858 | Suppression Ratio (BIS) | - | This variable is associated with ChannelID identifier. | Double | VitalSigns | IntelliVue MP70 | SR L/SR R
7865 | Right Ventricular Stroke Work Index | - |  | Double | VitalSigns | IntelliVue MP70 | RVSWI
7870 | Base Excess of Extracellular Fluid | - |  | Double | VitalSigns | IntelliVue MP70 | BEecf
7878 | Rise Time Setting | - |  | Double | Settings | IntelliVue MP70 | sRmpTi
7879 | Time Duration at Low Pressure (Tlow) | - |  | Double | VitalSigns | IntelliVue MP70 | Th/TL
7880 | Time Duration at High Pressure (Thigh) | - |  | Double | VitalSigns | IntelliVue MP70 | Th/TL
7904 | Spontaneous Tidal Volume | - |  | Double | VitalSigns | IntelliVue MP70 | VENT: SpTV
7905 | Cardiac Troponin l | - |  | Double | VitalSigns | IntelliVue MP70 | cTnI
7906 | Cardiopulmonary Bypass Status | - |  | Double | VitalSigns | IntelliVue MP70 | CPB
7907 | Cardiac Brain Natriuretic Peptide | - |  | Double | VitalSigns | IntelliVue MP70 | BNP
7908 | Mean Dominant Frequency | - |  | Double | VitalSigns | IntelliVue MP70 | EEG: MDF
7983 | Filtration Current Volume | - |  | Double | VitalSigns | IntelliVue MP70 | Vufr/h
7986 | Calcium Flow Rate Setting | - |  | Double | Settings | IntelliVue MP70 | sCSR
7987 | Calcium Compensation Setting | - |  | Double | Settings | IntelliVue MP70 | sCComp
8189 | Cerebral State Index (CSI) | - |  | Double | VitalSigns | IntelliVue MP70 | CSI
8201 | EEG Amplitude | - | This variable is associated with ChannelID identifier. | Double | VitalSigns | IntelliVue MP70 | aEEG1/aEEG2/aEEG3/aEEG4
8339 | Patient Middle Names | unitless | The Middle Name attribute contains the middle name of the patient. | String | PatientInfo | IntelliVue MP70 | Patient Demographics
8350 | Ao Pulse Rate | - |  | Double | VitalSigns | IntelliVue MP70 | Pulse
8351 | PAP Pulse Rate | - |  | Double | VitalSigns | IntelliVue MP70 | Pulse
8352 | CVP Pulse Rate | - |  | Double | VitalSigns | IntelliVue MP70 | Pulse
8353 | RAP Pulse Rate | - |  | Double | VitalSigns | IntelliVue MP70 | Pulse
8354 | LAP Pulse Rate | - |  | Double | VitalSigns | IntelliVue MP70 | Pulse
8355 | ICP Pulse Rate | - |  | Double | VitalSigns | IntelliVue MP70 | Pulse
8356 | Umbilical Venous Pressure Pulse Rate | - |  | Double | VitalSigns | IntelliVue MP70 | Pulse
8357 | FAP Pulse Rate | - |  | Double | VitalSigns | IntelliVue MP70 | Pulse
8358 | BAP Pulse Rate | - |  | Double | VitalSigns | IntelliVue MP70 | Pulse
8359 | IC1 Pulse Rate | - |  | Double | VitalSigns | IntelliVue MP70 | Pulse
8360 | Intracranial Pressure ICP2 Pulse Rate | - |  | Double | VitalSigns | IntelliVue MP70 | Pulse
8361 | Pulse Rate (P1) | - |  | Double | VitalSigns | IntelliVue MP70 | Pulse
8362 | Pulse Rate (P2) | - |  | Double | VitalSigns | IntelliVue MP70 | Pulse
8363 | Pulse Rate (P3) | - |  | Double | VitalSigns | IntelliVue MP70 | Pulse
8364 | Pulse Rate (P4) | - |  | Double | VitalSigns | IntelliVue MP70 | Pulse
8365 | Pulse Rate from Plethysmogram (Left) | - |  | Double | VitalSigns | IntelliVue MP70 | Pulse
8367 | Pleth Wave (Right) | - |  | Double | Waveforms | IntelliVue MP70 | Waves: PLETHr
8368 | Pleth Wave (Left) | - |  | Double | Waveforms | IntelliVue MP70 | Waves: PLETHl
8369 | Calculated SO2 | - |  | Double | VitalSigns | IntelliVue MP70 | CALC: 'SO2
8370 | Airway Respiration Rate (Ohmeda) | - |  | Double | VitalSigns | IntelliVue MP70 | VENT: RRaw
8371 | Airway Respiration Rate Setting (Ohmeda) | - |  | Double | Settings | IntelliVue MP70 | VENT: sRRaw
8372 | Fresh Gas O2 Concentration | - |  | Double | OtherMeasuredData | IntelliVue MP70 | fgO2
8373 | Manually entered FIO2 or Flow | - | This variable contains measurement from label "NLS_NOM_CONC_AWAY_O2" and observed value "NOM_CONC_AWAY_O2". | Double | VitalSigns | IntelliVue MP70 | O2
8374 | Generic O2 Parameter | - | This is the device's measurement with label="NLS_NOM_SETT_CONC_AWAY_O2" and physio_id="NOM_CONC_AWAY_O2". | Double |  | IntelliVue MP70 | sO2
8375 | Fresh Gas O2 Flow Setting | - | Mixer parameter. | Double | Settings | IntelliVue MP70 | sfgO2
8376 | Airway Pressure Numeric Value (Ohmeda) | - |  | Double | VitalSigns | IntelliVue MP70 | AWP
8377 | Inspiratory Peak Flow | - |  | Double | VitalSigns | IntelliVue MP70 | PIF
8378 | Maximum Pressure During Breathing Cycle | - |  | Double | VitalSigns | IntelliVue MP70 | VENT: Pmax
8379 | Peak Pressure | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
8380 | Peak Inspiratory Flow Setting | - |  | Double | Settings | IntelliVue MP70 | sPIF
8381 | Airway Maximum Pressure Setting | - |  | Double | Settings | IntelliVue MP70 | VENT: sPmax
8382 | Peak Inspiratory Pressure Setting | - |  | Double | Settings | IntelliVue MP70 | sPIP
8383 | Setting: Inspiratory Pressure (Ohmeda) | - |  | Double | Settings | IntelliVue MP70 | Available only in revF or earlier
8384 | Mean Airway Pressure (Ohmeda) | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
8385 | EEG Wave (Left) | - |  | Double | Waveforms | IntelliVue MP70 | Waves: LT EEG
8386 | EEG Wave (Right) | - |  | Double | Waveforms | IntelliVue MP70 | Waves: RT EEG
8387 | % of Power in the Delta Band - 1st EEG | - |  | Double | VitalSigns | IntelliVue MP70 | EEG: Prcnt1
8388 | % of Power in the Delta Band - 2nd EEG | - |  | Double | VitalSigns | IntelliVue MP70 | EEG: Prcnt2
8389 | Percent Delta - Left Side | - |  | Double | VitalSigns | IntelliVue MP70 | EEG: LT %DL
8390 | Percent Delta - Right Side | - |  | Double | VitalSigns | IntelliVue MP70 | EEG: RT %DL
8391 | % of Power in the Theta Band (1st EEG) | - |  | Double | VitalSigns | IntelliVue MP70 | EEG: Prcnt1
8392 | % of Power in the Theta Band (2nd EEG) | - |  | Double | VitalSigns | IntelliVue MP70 | EEG: Prcnt2
8393 | Percent Theta - Left Side | - |  | Double | VitalSigns | IntelliVue MP70 | EEG: LT %TH
8394 | Percent Theta - Right Side | - |  | Double | VitalSigns | IntelliVue MP70 | EEG: RT %TH
8395 | Generic Calculation Constant | - | This variable contains measurement from Philips label "NLS_NOM_NUM_CALC_CONST" and observed value "NOM_METRIC_NOS". See also variable 8493. | Double | VitalSigns | IntelliVue MP70 | CompCt
8396 | Pressure Ventilation Control Pressure | - |  | Double | VitalSigns | IntelliVue MP70 | VENT: PVcP
8397 | Right Heart Ejection Fraction | - |  | Double | VitalSigns | IntelliVue MP70 | REF
8398 | White Blood Count (Leucocyte Count) | - |  | Double | VitalSigns | IntelliVue MP70 | WBC
8399 | Red Blood Count (Erithrocyte Count) | - |  | Double | VitalSigns | IntelliVue MP70 | RBC
8400 | Mean Corpuscular Volume | - |  | Double | VitalSigns | IntelliVue MP70 | MCV
8401 | Corpuscular Mean Hemoglobin Content | - |  | Double | VitalSigns | IntelliVue MP70 | MCH
8402 | Ferrum | - |  | Double | VitalSigns | IntelliVue MP70 | Fe
8403 | Fetal Hemoglobin | - |  | Double | VitalSigns | IntelliVue MP70 | HbF
8404 | MetHemoglobin | - |  | Double | VitalSigns | IntelliVue MP70 | Met-Hb
8405 | Airway Volume Wave | - |  | Double | Waveforms | IntelliVue MP70 | Waves: AWV
8406 | Expiratory Peak Flow | - |  | Double | VitalSigns | IntelliVue MP70 | PEF
8407 | Expiratory Peak Flow | - |  | Double | VitalSigns | IntelliVue MP70 | eeFlow
8408 | The V stands for Volume | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
8409 | The R stands for Rate | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
8410 | Adjusted PaCO2 (at Patient Temperature) | - |  | Double | VitalSigns | IntelliVue MP70 | &PaCO2
8411 | Incubator Heating Power | - | Power Required to set Air&Pat Temp in the Incubator | Double | OtherMeasuredData | IntelliVue MP70 | INCUBATOR: Power
8412 | Actual Patient Age | - | Actual patient age, provided by a module, measured in years. Please note that this Patient Age will be available only when the corresponding module is connected to the Monitor. Variable 169 contains the patient age as provided by the Monitor itself (calculated from the birth date entered on the Monitor device.) | Double | PatientInfo | IntelliVue MP70 | Age
8413 | The P stands for Pressure | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
8414 | PCV Apnea Control Pressure Setting | - |  | Double | Settings | IntelliVue MP70 | VENT: sAPVcP
8415 | Fresh Gas Air Flow Setting | - | Mixer parameter. | Double | Settings | IntelliVue MP70 | sfgAir
8416 | PEEP/CPAP Lower Alarm Limit | - |  | Double | Settings | IntelliVue MP70 | AlarmLimit: loPEEP
8417 | Tidal Volume Upper Alarm Limit | - |  | Double | Settings | IntelliVue MP70 | AlarmLimit: highTV
8418 | T-Tre Temperature difference between Temp and Trect | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
8419 | T-Tbl Temperature difference between Temp and Tblood | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
8420 | T-Tco Temperature difference between Temp and Tcore | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
8421 | T-Tsk Temperature difference between Temp and Tskin | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
8422 | T-Tes Temperature difference between Temp and Tesoph | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
8423 | T-Tna Temperature difference between Temp and Tnaso | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
8424 | T-Tar Temperature difference between Temp and Tart | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
8425 | T-Tve Temperature difference between Temp and Tven | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
8426 | T-Taw Temperature difference between Temp and Tairwy | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
8427 | Tre-T Temperature difference between Trect and Temp | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
8428 | Tre-bl Temperature difference between Trect and Tblood | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
8429 | Tre-sk Temperature difference between Trect and Tskin | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
8430 | Tre-aw Temperature difference between Trect and Tairwy | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
8431 | Temperature difference between Tblood and Temp | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
8433 | Tbl-aw Temperature difference between Tblood and Tairwy | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
8434 | Tco-T Temperature difference between Tcore and Temp | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
8435 | Tco-re Temperature difference between Tcore and Trect | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
8436 | Tco-bl Temperature difference between Tcore and Tblood | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
8437 | Tco-sk Temperature difference between Tcore and Tskin | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
8438 | Tco-es Temperature difference between Tcore and Tesoph | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
8439 | Tco-na Temperature difference between Tcore and Tnaso | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
8440 | Tco-aw Temperature difference between Tcore and Tairwy | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
8441 | Tsk-T Temperature difference between Tskin and Temp | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
8442 | Tsk-bl Temperature difference between Tskin and Tblood | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
8443 | Tsk-aw Temperature difference between Tskin and Tairwy | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
8444 | Tes-T Temperature difference between Tesoph and Temp | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
8445 | Tes-re Temperature difference between Tesoph and Trect | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
8446 | Tes-bl Temperature difference between Tesoph and Tblood | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
8447 | Tes-sk Temperature difference between Tesoph and Tskin | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
8448 | Tes-na Temperature difference between Tesoph and Tnaso | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
8449 | Tes-aw Temperature difference between Tesoph and Tairwy | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
8450 | Tna-T Temperature difference between Tnaso and Temp | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
8451 | Tna-re Temperature difference between Tnaso and Trect | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
8452 | Tna-bl Temperature difference between Tnaso and Tblood | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
8453 | Tna-sk Temperature difference between Tnaso and Tskin | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
8454 | Tna-aw Temperature difference between Tnaso and Tairwy | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
8455 | Tar-T Temperature difference between Tart and Temp | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
8456 | Tar-re Temperature difference between Tart and Trect | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
8457 | Tar-bl Temperature difference between Tart and Tblood | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
8458 | Tar-co Temperature difference between Tart and Tcore | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
8459 | Tar-sk Temperature difference between Tart and Tskin | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
8460 | Tar-es Temperature difference between Tart and Tesoph | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
8461 | Tar-na Temperature difference between Tart and Tnaso | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
8462 | Tar-aw Temperature difference between Tart and Tairwy | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
8463 | Tve-T Temperature difference between Tven and Temp | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
8464 | Tve-re Temperature difference between Tven and Trect | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
8465 | Tve-bl Temperature difference between Tven and Tblood | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
8466 | Tve-co Temperature difference between Tven and Tcore | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
8467 | Tve-sk Temperature difference between Tven and Tskin | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
8468 | Tve-es Temperature difference between Tven and Tesoph | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
8469 | Tve-na Temperature difference between Tven and Tnaso | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
8470 | MAC Concentration (Awake) | - |  | Double | VitalSigns | IntelliVue MP70 | MAC
8471 | Tve-ar Temperature difference between Tven and Tart | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
8472 | Tve-aw Temperature difference between Tven and Tairwy | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
8473 | Taw-T Temperature difference between Tairwy and Temp | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
8474 | Taw-re Temperature difference between Tairwy and Trect | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
8475 | Taw-bl Temperature difference between Tairwy and Tblood | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
8476 | Taw-co Temperature difference between Tairwy and Tcore | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
8477 | Taw-sk Temperature difference between Tairwy and Tskin | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
8478 | Taw-es Temperature difference between Tairwy and Tesoph | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
8479 | Taw-na Temperature difference between Tairwy and Tnaso | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
8480 | Taw-ar Temperature difference between Tairwy and Tart | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
8481 | Taw-ve Temperature difference between Tairwy and Tven | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
8483 | Generic Agent 1 Wave | - |  | Double | Waveforms | IntelliVue MP70 | Waves: AGT1
8484 | End Tidal Generic Agent 1 | - |  | Double | VitalSigns | IntelliVue MP70 | AGT1
8485 | Inspired Generic Agent 1 | - |  | Double | VitalSigns | IntelliVue MP70 | AGT1
8486 | Generic Agent 2 Wave | - |  | Double | Waveforms | IntelliVue MP70 | Waves: AGT2
8487 | End Tidal Generic Agent 2 | - |  | Double | VitalSigns | IntelliVue MP70 | AGT2
8488 | Inspired Generic Agent 2 | - |  | Double | VitalSigns | IntelliVue MP70 | AGT2
8489 | Adjusted Arterial Blood pH Value | - |  | Double | VitalSigns | IntelliVue MP70 | &pHa
8490 | Adjusted Venous Blood pH Value | - |  | Double | VitalSigns | IntelliVue MP70 | &pHv
8491 | Adjusted PvCO2 (at Patient Temperature) | - |  | Double | VitalSigns | IntelliVue MP70 | &PvCO2
8492 | Adjusted PvO2 (at Patient Temperature) | - |  | Double | VitalSigns | IntelliVue MP70 | &PvO2
8493 | Generic Calculation Constant | - | This variable contains parameter from Philips label "NLS_NOM_NUM_CATHETER_CONST" and observed value "NOM_NUM_CATHETER_CONST". See also variable 8395. | Double | VitalSigns | IntelliVue MP70 | CathCt
8501 | ECG Lead V Wave | - |  | Double | Waveforms | IntelliVue MP70 | Waves: V
8502 | ECG Lead MCL Wave | - |  | Double | Waveforms | IntelliVue MP70 | Waves: MCL
8503 | ECG Lead MCL1 Wave | - |  | Double | Waveforms | IntelliVue MP70 | Waves: MCL1
8510 | MAC Concentration | - |  | Double | VitalSigns | IntelliVue MP70 | MAC
8511 | Pleth Wave (Telemetry) | - |  | Double | Waveforms | IntelliVue MP70 | Waves: PlethT
8513 | NIBP Systolic (Telemetry) | - |  | Double | VitalSigns | IntelliVue MP70 | NBP_T SYS
8514 | NIBP Diastolic (Telemetry) | - |  | Double | VitalSigns | IntelliVue MP70 | NBP_T DIA
8515 | NIBP Mean (Telemetry) | - |  | Double | VitalSigns | IntelliVue MP70 | NBP_T MEAN
8517 | Generic Pressure 1 (P1) Systolic | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
8518 | Generic Pressure 1 (P1) Diastolic | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
8519 | Generic Pressure 1 (P1) Mean | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
8521 | Generic Pressure 2 (P2) Systolic | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
8522 | Generic Pressure 2 (P2) Diastolic | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
8523 | Generic Pressure 2 (P2) Mean | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
8525 | Generic Pressure 3 (P3) Systolic | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
8526 | Generic Pressure 3 (P3) Diastolic | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
8527 | Generic Pressure 3 (P3) Mean | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
8529 | Generic Pressure 4 (P4) Systolic | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
8530 | Generic Pressure 4 (P4) Diastolic | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
8531 | Generic Pressure 4 (P4) Mean | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
8535 | Incubator Humidity | - |  | Double | OtherMeasuredData | IntelliVue MP70 | INCUBATOR: Hum
8536 | Predictive Oral Temperature | - |  | Double | VitalSigns | IntelliVue MP70 | pToral
8537 | Predictive Rectal Temperature | - |  | Double | VitalSigns | IntelliVue MP70 | pTrect
8538 | Predictive Axillary Temperature | - |  | Double | VitalSigns | IntelliVue MP70 | pTaxil
8539 | Perfusion (Telemetry) | - |  | Double | VitalSigns | IntelliVue MP70 | Perf T
8541 | Right to Left Heart Shunt | - |  | Double | VitalSigns | IntelliVue MP70 | RLShnt
8542 | Adjusted PcO2 (at Patient Temperature) | - |  | Double | VitalSigns | IntelliVue MP70 | &PCO2
8543 | Adjusted Arterial PaO2 (at Patient Temperature) | - |  | Double | VitalSigns | IntelliVue MP70 | &PaO2
8544 | Pulse Rate from Plethysmogram (Right) | - |  | Double | VitalSigns | IntelliVue MP70 | Pulse
8549 | Patient O2 Consumption per Body Surface | - |  | Double | VitalSigns | IntelliVue MP70 | VO2m²
8550 | Patient O2 Consumption per Body Weight | - |  | Double | VitalSigns | IntelliVue MP70 | VO2kg
8560 | Ideal Patient Weight | - |  | Double | PatientInfo | IntelliVue MP70 | IBW
8562 | End Flow Setting | - |  | Double | Settings | IntelliVue MP70 | sEndFl
8564 | PSV Backup Delay Setting | - |  | Double | Settings | IntelliVue MP70 | sPSVbd
8640 | Technical Alarm Internal Source ID | unitless | Consult detailed help for more information on this variable | Double (Multiset) | Alarms | Common | Alarms
8641 | Technical Alarm Source Type | unitless |  | Double (Multiset) | Alarms | Common | Alarms
8642 | Device Technical Alarm | unitless | Consult detailed help for more information on this variable | Double (Multiset) | Alarms | Common | Alarms
8643 | Patient Alarm Internal Source ID | unitless | Consult detailed help for more information on this variable | Double (Multiset) | Alarms | Common | Alarms
8644 | Patient Alarm Source Type | unitless |  | Double (Multiset) | Alarms | Common | Alarms
8645 | Patient Alarm | unitless | Consult detailed help for more information on this variable | Double (Multiset) | Alarms | Common | Alarms
8656 | Accoustic Respiration Rate Wave | - |  | Double | Waveforms | IntelliVue MP70 | Waves: acRESP
8703 | Blood Temperature Wave | - |  | Double | Waveforms | IntelliVue MP70 | Waves: Tblood
8714 | Non-specific Pressure 1 Wave | - |  | Double | Waveforms | IntelliVue MP70 | Waves: P1
8715 | Generic Invasive Blood Pressure P1 | - | Available on IntelliVue devices. | Double | VitalSigns | IntelliVue MP70 | IBP: P1
8716 | Generic Invasive Blood Pressure P1 Systolic | - |  | Double | VitalSigns | IntelliVue MP70 | IBP: P1
8717 | Generic Invasive Blood Pressure P1 Diastolic | - |  | Double | VitalSigns | IntelliVue MP70 | IBP: P1
8718 | Generic Invasive Blood Pressure P1 Mean | - |  | Double | VitalSigns | IntelliVue MP70 | IBP: P1
8719 | Non-specific Pressure 2 Wave | - |  | Double | Waveforms | IntelliVue MP70 | Waves: P2
8720 | Generic Invasive Blood Pressure P2 | - | Available on IntelliVue devices. | Double | VitalSigns | IntelliVue MP70 | IBP: P2
8721 | Generic Invasive Blood Pressure P2 Systolic | - |  | Double | VitalSigns | IntelliVue MP70 | IBP: P2
8722 | Generic Invasive Blood Pressure P2 Diastolic | - |  | Double | VitalSigns | IntelliVue MP70 | IBP: P2
8723 | Generic Invasive Blood Pressure P2 Mean | - |  | Double | VitalSigns | IntelliVue MP70 | IBP: P2
8724 | Non-specific Pressure 3 Wave | - |  | Double | Waveforms | IntelliVue MP70 | Waves: P3
8725 | Generic Invasive Blood Pressure P3 | - | Available on IntelliVue devices. | Double | VitalSigns | IntelliVue MP70 | IBP: P3
8726 | Generic Invasive Blood Pressure P3 Systolic | - |  | Double | VitalSigns | IntelliVue MP70 | IBP: P3
8727 | Generic Invasive Blood Pressure P3 Diastolic | - |  | Double | VitalSigns | IntelliVue MP70 | IBP: P3
8728 | Generic Invasive Blood Pressure P3 Mean | - |  | Double | VitalSigns | IntelliVue MP70 | IBP: P3
8729 | Non-specific Pressure 4 Wave | - |  | Double | Waveforms | IntelliVue MP70 | Waves: P4
8730 | Generic Invasive Blood Pressure P4 | - | Available on IntelliVue devices. | Double | VitalSigns | IntelliVue MP70 | IBP: P4
8731 | Generic Invasive Blood Pressure P4 Systolic | - |  | Double | VitalSigns | IntelliVue MP70 | IBP: P4
8732 | Generic Invasive Blood Pressure P4 Diastolic | - |  | Double | VitalSigns | IntelliVue MP70 | IBP: P4
8733 | Generic Invasive Blood Pressure P4 Mean | - |  | Double | VitalSigns | IntelliVue MP70 | IBP: P4
8734 | Non-specific Pressure 5 Wave | - |  | Double | Waveforms | IntelliVue MP70 | Waves: P_5
8735 | Generic Invasive Blood Pressure P5 | - | Available on IntelliVue devices. | Double | VitalSigns | IntelliVue MP70 | IBP: P5
8736 | Generic Invasive Blood Pressure P5 Systolic | - |  | Double | VitalSigns | IntelliVue MP70 | IBP: P5
8737 | Generic Invasive Blood Pressure P5 Diastolic | - |  | Double | VitalSigns | IntelliVue MP70 | IBP: P5
8738 | Generic Invasive Blood Pressure P5 Mean | - |  | Double | VitalSigns | IntelliVue MP70 | IBP: P5
8739 | Non-specific Pressure 6 Wave | - |  | Double | Waveforms | IntelliVue MP70 | Waves: P_6
8740 | Generic Invasive Blood Pressure P6 | - | Available on IntelliVue devices. | Double | VitalSigns | IntelliVue MP70 | IBP: P6
8741 | Generic Invasive Blood Pressure P6 Systolic | - |  | Double | VitalSigns | IntelliVue MP70 | IBP: P6
8742 | Generic Invasive Blood Pressure P6 Diastolic | - |  | Double | VitalSigns | IntelliVue MP70 | IBP: P6
8743 | Generic Invasive Blood Pressure P6 Mean | - |  | Double | VitalSigns | IntelliVue MP70 | IBP: P6
8744 | Non-specific Pressure 7 Wave | - |  | Double | Waveforms | IntelliVue MP70 | Waves: P_7
8745 | Generic Invasive Blood Pressure P7 | - | Available on IntelliVue devices. | Double | VitalSigns | IntelliVue MP70 | IBP: P7
8746 | Generic Invasive Blood Pressure P7 Systolic | - |  | Double | VitalSigns | IntelliVue MP70 | IBP: P7
8747 | Generic Invasive Blood Pressure P7 Diastolic | - |  | Double | VitalSigns | IntelliVue MP70 | IBP: P7
8748 | Generic Invasive Blood Pressure P7 Mean | - |  | Double | VitalSigns | IntelliVue MP70 | IBP: P7
8749 | Non-specific Pressure 8 Wave | - |  | Double | Waveforms | IntelliVue MP70 | Waves: P_8
8750 | Generic Invasive Blood Pressure P8 | - | Available on IntelliVue devices. | Double | VitalSigns | IntelliVue MP70 | IBP: P8
8751 | Generic Invasive Blood Pressure P8 Systolic | - |  | Double | VitalSigns | IntelliVue MP70 | IBP: P8
8752 | Generic Invasive Blood Pressure P8 Diastolic | - |  | Double | VitalSigns | IntelliVue MP70 | IBP: P8
8753 | Generic Invasive Blood Pressure P8 Mean | - |  | Double | VitalSigns | IntelliVue MP70 | IBP: P8
8754 | Intra-Uterine Pressure (IUP) Wave | - |  | Double | Waveforms | IntelliVue MP70 | Waves: IUP
8755 | Intra-Uterine Pressure | - |  | Double | VitalSigns | IntelliVue MP70 | IBP: IUP
8756 | Intra-Uterine Pressure Systolic | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
8757 | Intra-Uterine Pressure Diastolic | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
8758 | Intra-Uterine Pressure Mean | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
8760 | Unspecified Blood Pressure | - |  | Double | VitalSigns | IntelliVue MP70 | IBP: P
8761 | Unspecified Blood Pressure Systolic | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
8762 | Unspecified Blood Pressure Diastolic | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
8763 | Unspecified Blood Pressure Mean | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
8788 | Airway Flow Wave | - |  | Double | Waveforms | IntelliVue MP70 | Waves: AWF
8789 | Auxiliary Wave/Parameter | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
8790 | Global Ejection Fraction (2) | - | This is the device's measurement with label="NLS_NOM_EMFC_GEF" and physio_id="NOM_METRIC_NOS". See also 7810. | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
8791 | Pulmonary Vascular Permeability Index (2) | - | This is the device's measurement with label="NLS_NOM_EMFC_PVPI" and physio_id="NOM_METRIC_NOS". See also variable 6493. | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
8793 | Pressure Drop | - |  | Double | OtherMeasuredData | IntelliVue MP70 | DP
8803 | Replacement Pre-delivered Rate Setting | - |  | Double | Settings | IntelliVue MP70 | sRPrFl
8804 | Replacement Post-delivered Rate Setting | - |  | Double | Settings | IntelliVue MP70 | sRPoFl
8805 | Pre-blood Flow Rate Setting | - |  | Double | Settings | IntelliVue MP70 | sPBPFl
8806 | Patient Weight Balance Setting | - |  | Double | Settings | IntelliVue MP70 | sWTbal
8808 | Anticoagulant Bolus Interval Setting | - |  | Double | Settings | IntelliVue MP70 | sBolIn
8809 | Anticoagulant Infusion Rate Setting | - |  | Double | Settings | IntelliVue MP70 | sHepRt
8810 | Ultrafiltration Ratio Setting | - |  | Double | Settings | IntelliVue MP70 | sUFR
8811 | Therapy Elapsed Time | - |  | Double | OtherMeasuredData | IntelliVue MP70 | RunTi
8812 | Airway Volume | - |  | Double | VitalSigns | IntelliVue MP70 | Available only in revF or earlier
8813 | Processed Cumulated Blood Volume | - |  | Double | VitalSigns | IntelliVue MP70 | BVolPr
8814 | Replacement Pre-delivered Cumulated Volume | - |  | Double | VitalSigns | IntelliVue MP70 | RpreC
8815 | Replacement Post-delivered Cumulated Volume | - | This variable is associated with ChannelID identifier. | Double | VitalSigns | IntelliVue MP70 | Vdpo/h / RpostC
8816 | Pre-blood Cumulated Volume | - | This variable is associated with ChannelID identifier. | Double | VitalSigns | IntelliVue MP70 | VPBP/h / PreBC
8835 | Syringe Speed | - |  | Double | OtherMeasuredData | IntelliVue MP70 | Vsyr/h
8851 | Maximum Volume Gain Setting | - |  | Double | Settings | IntelliVue MP70 | sMaxG
8853 | Patient Fluid Balance | - |  | Double | VitalSigns | IntelliVue MP70 | Wtcum
8857 | Arterial Blood Pressure (ABP) Wave | - |  | Double | Waveforms | IntelliVue MP70 | Waves: ABP
8858 | Arterial Blood Pressure (ART) Wave | - |  | Double | Waveforms | IntelliVue MP70 | Waves: ART
8859 | Pulmonary Arterial Pressure (PAP) Wave | - |  | Double | Waveforms | IntelliVue MP70 | Waves: PAP
8860 | Central Venous Pressure (CVP) Wave | - |  | Double | Waveforms | IntelliVue MP70 | Waves: CVP
8861 | Right Atrial Pressure (RAP) Wave | - |  | Double | Waveforms | IntelliVue MP70 | Waves: RAP
8862 | Left Atrial Pressure (LAP) Wave | - |  | Double | Waveforms | IntelliVue MP70 | Waves: LAP
8863 | Intra-cranial Pressure (ICP) Wave | - |  | Double | Waveforms | IntelliVue MP70 | Waves: ICP
8864 | Umbilical Arterial Pressure (UAP) Wave | - |  | Double | Waveforms | IntelliVue MP70 | Waves: UAP
8865 | Umbilical Venous Pressure (UVP) Wave | - |  | Double | Waveforms | IntelliVue MP70 | Waves: UVP
8866 | Airway CO2 Level Waveform | - |  | Double | Waveforms | IntelliVue MP70 | Waves: CO2
8867 | Airway O2 Wave | - |  | Double | Waveforms | IntelliVue MP70 | Waves: O2
8868 | Airway N2 Wave | - |  | Double | Waveforms | IntelliVue MP70 | Waves: N2
8869 | Generic Agent Wave (2nd) | - |  | Double | Waveforms | IntelliVue MP70 | Waves: AGTs
8870 | Impedance Respiration Wave | - |  | Double | Waveforms | IntelliVue MP70 | Waves: Resp
8871 | Brachial Arterial Pressure (BAP) Wave | - |  | Double | Waveforms | IntelliVue MP70 | Waves: BAP
8872 | Intracranial Pressure 1 (IC1) Wave | - |  | Double | Waveforms | IntelliVue MP70 | Waves: IC1
8873 | Intracranial Pressure 2 (IC2) Wave | - |  | Double | Waveforms | IntelliVue MP70 | Waves: IC2
8874 | Femoral Arterial Pressure (FAP) Wave | - |  | Double | Waveforms | IntelliVue MP70 | Waves: FAP
8875 | Vector ECG Wave | - |  | Double | Waveforms | IntelliVue MP70 | Waves: vECG
8889 | Transcutaneous Heater Power | - |  | Double | OtherMeasuredData | IntelliVue MP70 | HeatPa
8939 | BIS Standard Deviation | - | This variable is associated with ChannelID identifier. | Double | VitalSigns | IntelliVue MP70 | sdBISl/sdBISr
8940 | EMG Standard Deviation (BIS) | - | This variable is associated with ChannelID identifier. | Double | VitalSigns | IntelliVue MP70 | sdEMGl/sdEMGr
8941 | Asymmetric Index | - | This variable is associated with ChannelID identifier. | Double | VitalSigns | IntelliVue MP70 | ASYM/ASYM L/ASYM R
9037 | O2 Consumption Index Estimate | - |  | Double | VitalSigns | IntelliVue MP70 | iVO2Ie
9110 | Aortic Blood Pressure (Ao) | - | Available on IntelliVue devices with firmware revision G or later. Arterial Blood Pressure in the Aorta (NOM_PRESS_BLD_AORT). | Double | VitalSigns | Common | IBP: Ao
9111 | Umbilical Arterial Pressure (UAP) | - | Available on IntelliVue devices with firmware revision G or later. | Double | VitalSigns | Common | IBP: UAP
9112 | Umbilical Venous Pressure (UVP) | - | Available on IntelliVue devices with firmware revision G or later. | Double | VitalSigns | Common | IBP: UVP
9154 | Humidifier Activation Status | - |  | Double | DeviceInfo | IntelliVue MP70 | sHmCtl
9155 | Nebulizer Time Setting | - |  | Double | Settings | IntelliVue MP70 | sNebTi
9165 | Edi Diaphragm Electrical Activity Waveform | - |  | Double | Waveforms | IntelliVue MP70 | Waves: EDI
9166 | Peak Diaphragm Electrical Activity (Edi) | - |  | Double | VitalSigns | IntelliVue MP70 | Edimax
9167 | Minimum Diaphragm Electrical Activity (Edi) | - |  | Double | VitalSigns | IntelliVue MP70 | Edimin
9168 | % Edi Trigger In NAVA | - |  | Double | VitalSigns | IntelliVue MP70 | %EdiTr
9169 | % Edi Cycle Off In NAVA | - |  | Double | VitalSigns | IntelliVue MP70 | %EdiCy
9170 | NAVA Level Setting | - |  | Double | Settings | IntelliVue MP70 | sNAVA
9171 | Edi Trigger Setting | - |  | Double | Settings | IntelliVue MP70 | sEdiTr
9196 | Patient Body Mass Index (BMI) | - |  | Double | PatientInfo | IntelliVue MP70 | BMI
9270 | Maximum Positive Auxiliary Pressure | - |  | Double | VitalSigns | IntelliVue MP70 | Pxpeak
9271 | Mean Positive Auxiliary Pressure | - |  | Double | VitalSigns | IntelliVue MP70 | Pxmean
9272 | Minimum Positive Auxiliary Pressure | - |  | Double | VitalSigns | IntelliVue MP70 | Pxmin
9286 | Patient CO2 Production Indexed By Body Surface | - |  | Double | VitalSigns | IntelliVue MP70 | VCO2m²
9287 | Patient CO2 Production Indexed By Body Weight | - |  | Double | VitalSigns | IntelliVue MP70 | VCO2kg
9288 | Tidal Volume Indexed By Body Weight | - |  | Double | VitalSigns | IntelliVue MP70 | TV/kg
9314 | Functional Residual Capacity | - |  | Double | VitalSigns | IntelliVue MP70 | FRC
9315 | Corrected QT Interval (QTc) | - |  | Double | VitalSigns | IntelliVue MP70 | ECG: QTc
9317 | Airway N2 Concentration | - | Measurement from VueLink | Double | VitalSigns | IntelliVue MP70 | N2
9318 | Airway Agent Concentration | - | Measurement from VueLink. This variable is associated with a SourceID identifier. | Double | VitalSigns | IntelliVue MP70 | AGT
9321 | Respiration Rate (Spirometry) | - |  | Double | VitalSigns | IntelliVue MP70 | RRspir
9323 | QT Heart Rate (QT-HR) | - | QT heart rate | Double | VitalSigns | IntelliVue MP70 | ECG: QT-HR
9324 | Delta QTc | - |  | Double | VitalSigns | IntelliVue MP70 | ECG: DeltaQTc
9329 | Humidification Type Setting | - |  | Double | Settings | IntelliVue MP70 | sHumid
9336 | Normalized Rapid Shallow Breathing Index | - |  | Double | VitalSigns | IntelliVue MP70 | RSBInm
9337 | Percent of Spontaneous Inspiratory Time | - |  | Double | VitalSigns | IntelliVue MP70 | Tin/Tt
9408 | Energy Balance Calculated by Metabolics | - |  | Double | VitalSigns | IntelliVue MP70 | dE
9437 | Alveolar Volume | - |  | Double | VitalSigns | IntelliVue MP70 | Valv
9438 | Apparatus Dead Space | - |  | Double | VitalSigns | IntelliVue MP70 | appVd
9439 | Pleth (Pre-Ductal) | - | Pulse Rate from Plethysmogram (pre ductal) | Double | VitalSigns | IntelliVue MP70 | Pulse
9440 | Pleth (Post-Ductal) | - | Pulse Rate from Plethysmogram (post ductal) | Double | VitalSigns | IntelliVue MP70 | Pulse
9441 | Relative Perfusion (Pre-Ductal) | - |  | Double | VitalSigns | IntelliVue MP70 | PerfPr
9442 | Relative Perfusion (Post-Ductal) | - |  | Double | VitalSigns | IntelliVue MP70 | PerfPo
9444 | Cardioplegia Main Pump Time of Current Plegia Session | - |  | Double | VitalSigns | IntelliVue MP70 | CpOnTi
9445 | Cardioplegia Main Pump Time Since Stop of Plegia Session | - |  | Double | VitalSigns | IntelliVue MP70 | CpOffT
9446 | Cardioplegia Main Pump Total Plegia Time | - |  | Double | VitalSigns | IntelliVue MP70 | CpPlTi
9447 | Cardioplegia Slave Pump Time Since Start of Current Plegia Session | - |  | Double | VitalSigns | IntelliVue MP70 | CsOnTi
9448 | Cardioplegia Slave Pump Time Since Stop of Plegia Session | - |  | Double | VitalSigns | IntelliVue MP70 | CsOffT
9449 | Cardioplegia Slave Pump Total Plegia Time | - |  | Double | VitalSigns | IntelliVue MP70 | CsPlTi
9453 | PAV-based Lung Compliance | - |  | Double | VitalSigns | IntelliVue MP70 | Cpav
9454 | PAV-based Patient Resistance | - |  | Double | VitalSigns | IntelliVue MP70 | Rpav
9455 | PAV-based Total Resistance Estimate | - |  | Double | VitalSigns | IntelliVue MP70 | Rtot
9456 | PAV-based Lung Elastance | - |  | Double | VitalSigns | IntelliVue MP70 | Epav
9467 | Backup Inspiratory Pressure Setting | - |  | Double | Settings | IntelliVue MP70 | sPIbak
9476 | Percent of Leakage | - |  | Double | OtherMeasuredData | IntelliVue MP70 | %Leak
9485 | Maximum Inspiratory Time Setting | - |  | Double | Settings | IntelliVue MP70 | sTimax
9585 | Plateau Time | - |  | Double | VitalSigns | IntelliVue MP70 | PlatTi
9586 | QT Baseline | - |  | Double | VitalSigns | IntelliVue MP70 | ECG: QT Bsl
9587 | QT Heart Rate Baseline | - |  | Double | VitalSigns | IntelliVue MP70 | ECG: QTHRBl
9588 | High Density Lipoprotein | - |  | Double | VitalSigns | IntelliVue MP70 | HDL
9589 | Low Density Lipoprotein | - |  | Double | VitalSigns | IntelliVue MP70 | LDL
9590 | C-reactive Protein | - |  | Double | VitalSigns | IntelliVue MP70 | CRP
9591 | Urine Hemoglobin | - |  | Double | VitalSigns | IntelliVue MP70 | UrHb
9595 | Noninvasive Total Hemoglobin (SpHb) | - |  | Double | VitalSigns | IntelliVue MP70 | SpHb
9596 | Noninvasive Oxygen Content (SpOC) | - |  | Double | VitalSigns | IntelliVue MP70 | SpOC
9597 | Noninvasive Carboxyhemoglobin (SpCO) | - |  | Double | VitalSigns | IntelliVue MP70 | SpCO
9598 | Noninvasive Methemoglobin (SpMet) | - |  | Double | VitalSigns | IntelliVue MP70 | SpMet
9600 | Pleth Variability Index (PVI) | - |  | Double | VitalSigns | IntelliVue MP70 | PVI
9603 | Regional O2 Saturation (rSO2) | - |  | Double | VitalSigns | Common
9605 | rSO2 Area Under Curve (AUC) | - | This variable is associated with ChannelID identifier. | Double | VitalSigns | IntelliVue MP70 | AUC_1/AUC_2/AUC_3/AUC_4/AUC_5/AUC_6
9610 | Generic O2 Saturation (3) | - |  | Double | VitalSigns | IntelliVue MP70 | SO2 3
9611 | Generic O2 Saturation (4) | - |  | Double | VitalSigns | IntelliVue MP70 | SO2 4
9612 | Tissue O2 Saturation (StO2) | - |  | Double | VitalSigns | IntelliVue MP70 | StO2
9613 | Nitrogen Balance (NsLoss) | - |  | Double | VitalSigns | IntelliVue MP70 | NsLoss
52523 | Inspired O2 Concentration | % | See variable 52527 for partial pressure | Double | VitalSigns | IntelliVue MP70 | O2
52527 | Inspired O2 Partial Pressure | mmHg | This variable is calculated from variable 52523 and will output inspired O2 partial pressure. | Double | VitalSigns

## MX800
ID | Label | Unit | Description | Type | Group | Device | Source
2 | NIBP Systolic | - |  | Double | VitalSigns | Common | NBP
3 | NIBP Diastolic | - |  | Double | VitalSigns | Common | NBP
4 | NIBP Mean | - |  | Double | VitalSigns | Common | NBP
7 | Patient Delta Temperature | - |  | Double | VitalSigns | Common | DeltaTemp
14 | SpO2 | - |  | Double | VitalSigns | Common | SpO2
16 | NIBP Pulse Rate | - |  | Double | VitalSigns | IntelliVue MX800 | Pulse
19 | ECG Heart Rate | - |  | Double | VitalSigns | Common | ECG: HR
22 | Respiration Rate | - |  | Double | VitalSigns | Common | RR
26 | Device Inactivation Status | unitless | Possible values: "enabled" and "audio-paused". See "Multiset Alarms" on help file. | String | DeviceInfo | Common | Alarms
28 | Active Alarms | unitless | See "Multiset Alarms" on help file. | Double (Multiset) | Alarms | Common | Alarms
Consult detailed help for more information on this variable
29 | Latched Alarms | unitless | See "Multiset Alarms" on help file. | Double (Multiset) | Alarms | Common | Alarms
Consult detailed help for more information on this variable
33 | Inactive Alarms | unitless | See "Multiset Alarms" on help file. | Double (Multiset) | Alarms | Common | Alarms
Consult detailed help for more information on this variable
34 | General ST Level 1 | - | Physiological identifier: NOM_ECG_AMPL_ST_I | Double | VitalSigns | Common | ECG: ST
35 | General ST Level 2 | - | Physiological identifier: NOM_ECG_AMPL_ST_II | Double | VitalSigns | Common | ECG: ST
36 | General ST Level 3 | - | Physiological identifier: NOM_ECG_AMPL_ST_III | Double | VitalSigns | Common | ECG: ST
45 | Temperature 3 | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
46 | Temperature 4 | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
48 | O2 Saturation (SO2) | - |  | Double | VitalSigns | IntelliVue MX800 | SO2
113 | Pleth Waveform | - |  | Double | Waveforms | IntelliVue MX800 | Waves: Pleth
127 | ECG Measurement Status | - | ECG Rhythm Status | Double | OtherMeasuredData | Common | ECG Status (on IntelliVue)
Consult detailed help for more information on this variable
169 | Patient Age | - | Patient age. The Monitor device calculates the age of the patient from the birth data manually entered on the Monitor. | Double | PatientInfo | Common | Patient Demographics
170 | Patient Sex | unitless | Possible values: "unknown", "male", "female", and "unspecified". | String | PatientInfo | Common | Patient Demographics
405 | Generic O2 Saturation (1) | - |  | Double | VitalSigns | IntelliVue MX800 | SO2 1
406 | Generic O2 Saturation (2) | - |  | Double | VitalSigns | IntelliVue MX800 | SO2 2
414 | Cardiac Output | - |  | Double | VitalSigns | IntelliVue MX800 | C.O.
465 | Fetal Heart Rate (FHR) | - | This variable is associated with ChannelID identifier. | Double | VitalSigns | IntelliVue MX800 | DFHR/DFHR1/DFHR2/DFHR3
467 | Fetal Beat-to-Beat Rate (FHR) | - | This variable is associated with ChannelID identifier. | Double | VitalSigns | IntelliVue MX800 | bDFHR/bDFHR1/bDFHR2/bDFHR3
469 | TOCO | - |  | Double | VitalSigns | IntelliVue MX800 | Toco
497 | ECG Lead Name | - | Physiological identifier: NOM_ECG_AMPL_ST | Double | OtherMeasuredData | Common | Available only in revF or earlier
501 | Arterial Pressure (ART) Systolic | - |  | Double | VitalSigns | Common | IBP: ART
502 | Arterial Pressure (ART) Diastolic | - |  | Double | VitalSigns | Common | IBP: ART
503 | Arterial Pressure (ART) Mean | - |  | Double | VitalSigns | Common | IBP: ART
512 | Left Atrial Mean Pressure (LAP) | - |  | Double | VitalSigns | Common | IBP: LAP
519 | Central Venous Pressure (CVP) | - | Available on IntelliVue devices with firmware revision G or later. | Double | VitalSigns | Common | IBP: CVP
538 | Device Language | unitless | Consult detailed help for more information on this variable | String | DeviceInfo | Common | System Information
591 | ECG Lead II | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
592 | Airway N2O Concentration | - | Measurement from VueLink | Double | VitalSigns | IntelliVue MX800 | N2O
593 | Inspired Desflurane | - |  | Double | VitalSigns | IntelliVue MX800 | DES
595 | Inspired Sevoflurane | - |  | Double | VitalSigns | IntelliVue MX800 | SEV
597 | Fresh Gas Agent Value | - |  | Double | VitalSigns | IntelliVue MX800 | fgAGT
598 | Fresh Gas Halothane | - |  | Double | VitalSigns | IntelliVue MX800 | fgHAL
599 | Fresh Gas Enflurane | - |  | Double | VitalSigns | IntelliVue MX800 | fgENF
600 | Fresh Gas Isoflurane | - |  | Double | VitalSigns | IntelliVue MX800 | fgISO
601 | Inspired Halothane | - |  | Double | VitalSigns | IntelliVue MX800 | HAL
603 | Inspired Enflurane | - |  | Double | VitalSigns | IntelliVue MX800 | ENF
605 | Inspired Isoflurane | - |  | Double | VitalSigns | IntelliVue MX800 | ISO
608 | End Tidal CO2 | - |  | Double | VitalSigns | Common | CO2
616 | Airway Pressure | - | Measured in the inspiratory path. | Double | VitalSigns | IntelliVue MX800 | AWPin
618 | Airway N2O Wave | - |  | Double | Waveforms | IntelliVue MX800 | Waves: N2O
620 | Halothane Concentration Wave | - |  | Double | Waveforms | IntelliVue MX800 | Waves: HAL
621 | Enflurane Concentration Wave | - |  | Double | Waveforms | IntelliVue MX800 | Waves: ENF
622 | Isoflurane Concentration Wave | - |  | Double | Waveforms | IntelliVue MX800 | Waves: ISO
623 | Desflurane Concentration Wave | - |  | Double | Waveforms | IntelliVue MX800 | Waves: DES
624 | Sevoflurane Concentration Wave | - |  | Double | Waveforms | IntelliVue MX800 | Waves: SEV
625 | Apnea Duration | - | This is the device's measurement with label="NLS_NOM_TIME_PD_APNEA" and physio_id="NOM_TIME_PD_APNEA". | Double | VitalSigns | IntelliVue MX800 | ApneaD
626 | End Tidal CO2 | - |  | Double | VitalSigns | IntelliVue MX800 | CRT
630 | Cardiac Beat-to-Beat Rate | - | Cardiac Beat-to-Beat Rate (NOM_ECG_CARD_BEAT_RATE_BTB). | Double | VitalSigns | Common | ECG: btbHR
633 | Airway Delta O2 Concentration | - |  | Double | VitalSigns | IntelliVue MX800 | VENT: DeltaO2
649 | Cardiac Output Index | - |  | Double | VitalSigns | IntelliVue MX800 | C.I.
650 | Stroke Volume | - |  | Double | VitalSigns | IntelliVue MX800 | SV
651 | Stroke Volume Index | - |  | Double | VitalSigns | IntelliVue MX800 | SI
654 | Systemic Vascular Resistance | - |  | Double | VitalSigns | IntelliVue MX800 | SVR
669 | Normalized Ionized Calcium Concentration (Ca) | - |  | Double | VitalSigns | IntelliVue MX800 | iCa-N
678 | Potassium Ion Concentration (K+) | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
696 | QT Interval (QT) | - |  | Double | VitalSigns | IntelliVue MX800 | ECG: QT
764 | Device Specification ID | unitless | Variables 764, 765 and 4201 must be interpreted conjointly. | String (Multiset) | DeviceInfo | Common | System Information
Consult detailed help for more information on this variable
765 | Device Specification Type | unitless | Variables 764, 765 and 4201 must be interpreted conjointly. | String (Multiset) | DeviceInfo | Common | System Information
Consult detailed help for more information on this variable
794 | Pleth Wave (2) | - |  | Double | Waveforms | IntelliVue MX800 | Waves: PLETH2
854 | Total Number of Used Disposable | - |  | Double | DeviceInfo | IntelliVue MX800 | nFilt
946 | Cumulated Syringe Infused Volume | - |  | Double | OtherMeasuredData | IntelliVue MX800 | SyrC
966 | Current Therapy Time | - |  | Double | OtherMeasuredData | IntelliVue MX800 | FiltTi
1043 | Manually Entered Hematocrit | - |  | Double | Settings | IntelliVue MX800 | scHCTi
1055 | Infusion Drug Name | - |  | Double | PatientInfo | IntelliVue MX800 | sDrug
1072 | Infusion Time Remaining | - |  | Double | VitalSigns | IntelliVue MX800 | InfTim
1095 | Infusion Dose Rate Setting | - |  | Double | Settings | IntelliVue MX800 | sDosRt
1096 | Volume To Be Infused Original Setting | - |  | Double | Settings | IntelliVue MX800 | sVtot
1110 | KVO Rate Setting | - |  | Double | Settings | IntelliVue MX800 | sKVO
1186 | Inspired Minute Volume | - |  | Double | VitalSigns | IntelliVue MX800 | MINVOL
1190 | I:E Ratio | - | This variable contains expiratory component of I:E Ratio assuming that the inspiratory component is "1". | Double | VitalSigns | IntelliVue MX800 | I:E 1:
1208 | SpO2 | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
1220 | Patient Birth Date | unitless | Format: 14 ascii digits string of "YYYYMMDDhhmmss" | String | PatientInfo | Common | Patient Demographics
1234 | SpO2 (Pre-Ductal) | - |  | Double | VitalSigns | IntelliVue MX800 | SpO2pr
1236 | SpO2 (Post-Ductal) | - |  | Double | VitalSigns | IntelliVue MX800 | SpO2po
1244 | ECG Lead III | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
1246 | Hematocrit | - |  | Double | VitalSigns | IntelliVue MX800 | Hct
1250 | Pulmonary Vascular Resistance | - |  | Double | VitalSigns | IntelliVue MX800 | PVR
1251 | Body Surface Area | - |  | Double | PatientInfo | IntelliVue MX800 | BSA
1255 | ECG Lead MCL1 | - | Physiological identifier: NOM_ECG_AMPL_ST_MCL | Double | VitalSigns | Common | ECG: ST
1257 | ECG Lead V1 | - | Physiological identifier: NOM_ECG_AMPL_ST_V1 | Double | VitalSigns | Common | ECG: ST
1258 | ECG Lead V6 | - | Physiological identifier: NOM_ECG_AMPL_ST_V6 | Double | VitalSigns | Common | ECG: ST
1259 | ECG Lead aVL | - | Physiological identifier: NOM_ECG_AMPL_ST_AVL | Double | VitalSigns | Common | ECG: ST
1260 | ECG Lead aVF | - | Physiological identifier: NOM_ECG_AMPL_ST_AVF | Double | VitalSigns | Common | ECG: ST
1261 | ECG Lead aVR | - | Physiological identifier: NOM_ECG_AMPL_ST_AVR | Double | VitalSigns | Common | ECG: ST
1262 | ECG Lead V2 | - | Physiological identifier: NOM_ECG_AMPL_ST_V2 | Double | VitalSigns | Common | ECG: ST
1263 | ECG Lead V3 | - | Physiological identifier: NOM_ECG_AMPL_ST_V3 | Double | VitalSigns | Common | ECG: ST
1264 | ECG Lead V4 | - | Physiological identifier: NOM_ECG_AMPL_ST_V4 | Double | VitalSigns | Common | ECG: ST
1265 | ECG Lead V5 | - | Physiological identifier: NOM_ECG_AMPL_ST_V5 | Double | VitalSigns | Common | ECG: ST
1283 | Patient Temperature | - | This variable is associated with a SourceID Identifier. | Double | VitalSigns | IntelliVue MX800 | Infrared Temp
1286 | Micro Environment Air Temperature | - |  | Double | OtherMeasuredData | IntelliVue MX800 | INCUBATOR: Air T
1306 | Tidal Volume | - |  | Double | VitalSigns | IntelliVue MX800 | TV
1319 | Static Resistance | - |  | Double | VitalSigns | IntelliVue MX800 | Rstat
1321 | Tidal Volume Setting | - |  | Double | Settings | IntelliVue MX800 | sTV
1323 | Pressure Sensitivity Setting | - |  | Double | Settings | IntelliVue MX800 | sSenPr
1324 | Sigh Interval Setting | - |  | Double | Settings | IntelliVue MX800 | sTsigh
1325 | Sigh Volume Setting | - |  | Double | Settings | IntelliVue MX800 | sVsigh
1332 | Pressure Support Level Above PEEP Setting | - |  | Double | Settings | IntelliVue MX800 | sdPEEP
1333 | Spontaneous Minute Volume | - |  | Double | VitalSigns | IntelliVue MX800 | VENT: SpMV
1340 | Negative Inspiratory Pressure | - |  | Double | VitalSigns | IntelliVue MX800 | NgInsP
1342 | Peak Spontaneous Flow | - |  | Double | VitalSigns | IntelliVue MX800 | SpPkFl
1343 | Flow-by Base Flow Setting | - |  | Double | Settings | IntelliVue MX800 | VENT: sBasFl
1344 | Flow Sensitivity Setting | - |  | Double | Settings | IntelliVue MX800 | VENT: sSenFl
1348 | Expiratory Time | - |  | Double | VitalSigns | IntelliVue MX800 | ExpTi
1381 | Inspiratory Flow Setting | - |  | Double | Settings | IntelliVue MX800 | sInsFl
1382 | Expiratory Flow Setting | - |  | Double | Settings | IntelliVue MX800 | sExpFl
1383 | Inspiratory Flow Trigger Setting | - | Delivered by the Evita 2 VueLink Driver. | Double | Settings | IntelliVue MX800 | VENT: sTrgFl
1387 | APRV Low Pressure Setting | - |  | Double | Settings | IntelliVue MX800 | sPSVl
1389 | APRV High Pressure Setting | - |  | Double | Settings | IntelliVue MX800 | sPSVh
1391 | Apnea Time Setting | - |  | Double | Settings | IntelliVue MX800 | sAlnTi
1398 | Minute Volume Setting | - |  | Double | Settings | IntelliVue MX800 | sMV
1418 | Positive End Expiratory Pressure (PEEP) | - |  | Double | VitalSigns | IntelliVue MX800 | VENT: PEEP
1419 | Ventilator Work of Breathing | - |  | Double | VitalSigns | IntelliVue MX800 | WOBv
1422 | Rapid Shallow Breathing Index | - |  | Double | VitalSigns | IntelliVue MX800 | RSBI
1424 | Patient Work of Breathing | - |  | Double | VitalSigns | IntelliVue MX800 | WOBp
1435 | Systemic Vascular Resistance Index | - |  | Double | VitalSigns | IntelliVue MX800 | SVRI
1436 | Pulmonary Vascular Resistance Index | - |  | Double | VitalSigns | IntelliVue MX800 | PVRI
1442 | Continuous Cardiac Output | - |  | Double | VitalSigns | IntelliVue MX800 | CCO
1443 | Continuous Cardiac Output Index | - |  | Double | VitalSigns | IntelliVue MX800 | CCI
1451 | Pump Infusion Status | - |  | Double | DeviceInfo | IntelliVue MX800 | PuStat
1463 | Hemoglobin (Hb) | - | This variable is associated with SourceID identifier. | Double | VitalSigns | IntelliVue MX800 | sHb
1464 | Pleth Wave (Pre-Ductal) | - |  | Double | Waveforms | IntelliVue MX800 | Waves: PLTHpr
1465 | Pleth Wave (Post-Ductal) | - |  | Double | Waveforms | IntelliVue MX800 | Waves: PLTHpo
1466 | Arterial Venous O2 Content Difference | - |  | Double | VitalSigns | IntelliVue MX800 | Ca-vO2
1467 | O2 Delivery (DO2) | - |  | Double | VitalSigns | IntelliVue MX800 | DO2
1469 | O2 Extraction Ratio (O2ER) | - |  | Double | VitalSigns | IntelliVue MX800 | O2ER
1471 | O2 Delivery Index (DO2I) | - |  | Double | VitalSigns | IntelliVue MX800 | DO2I
1472 | O2 Consumption Index (VO2I) | - |  | Double | VitalSigns | IntelliVue MX800 | VO2I
1473 | Oxygenation Ratio (PaO2/FiO2) | - | PaO2 to FIO2 ratio expressed in mmHg to % ratio. | Double | VitalSigns | IntelliVue MX800 | PaFIO2
1475 | Physiologic Dead Space | - |  | Double | VitalSigns | IntelliVue MX800 | VENT: Vd
1477 | Alveolar Ventilation Volume (VA) | - |  | Double | VitalSigns | IntelliVue MX800 | VENT: ALVENT
1478 | Urea Nitrogen Concentration (BUN) | - |  | Double | VitalSigns | IntelliVue MX800 | BUN
1479 | Creatinine Concentration (Crea) | - | Measured by the i-Stat module. | Double | VitalSigns | IntelliVue MX800 | Crea
1480 | Plasma Osmolarity (Osm) | - |  | Double | VitalSigns | IntelliVue MX800 | PlOsm
1482 | Urine Creatinine | - |  | Double | VitalSigns | IntelliVue MX800 | UrCrea
1484 | Urine Potassium (K+) | - |  | Double | VitalSigns | IntelliVue MX800 | UrK
1486 | Urine Osmolarity | - |  | Double | VitalSigns | IntelliVue MX800 | UrOsm
1487 | Blood Urea Nitrogen Creatinine Ratio (BUN/cr) | - |  | Double | VitalSigns | IntelliVue MX800 | BUN/cr
1488 | Osmolar Clearance (COSM) | - |  | Double | VitalSigns | IntelliVue MX800 | COsm
1489 | Creatinine Clearance (CreaCl) | - |  | Double | VitalSigns | IntelliVue MX800 | CreaCl
1490 | Fractional Excretion of Sodium (FeNa) | - |  | Double | VitalSigns | IntelliVue MX800 | FeNa
1493 | Urine Serum Creatinine Ratio | - |  | Double | VitalSigns | IntelliVue MX800 | U/SCr
1495 | Urine Potassium Excretion | - |  | Double | VitalSigns | IntelliVue MX800 | UrKEx
1496 | Urine Sodium Excretion | - |  | Double | VitalSigns | IntelliVue MX800 | UrNaEx
1497 | Free Water Clearance (CH2O) | - |  | Double | VitalSigns | IntelliVue MX800 | CH2O
1520 | Inspired N2O | - |  | Double | VitalSigns | IntelliVue MX800 | N2O
1523 | Premature Ventricular Contractions | - |  | Double | VitalSigns | Common | ECG: PVC
1533 | LAP Systolic | - |  | Double | VitalSigns | Common | IBP: LAP
1534 | LAP Diastolic | - |  | Double | VitalSigns | Common | IBP: LAP
1540 | TcGas Heating Power | - |  | Double | DeviceInfo | IntelliVue MX800 | TCUT: HeatPw
1556 | Abdominal Fetal Heart Rate (FHR) | - |  | Double | VitalSigns | IntelliVue MX800 | aFHR
1557 | Abdominal Fetal Beat-to-Beat Rate (FHR) | - |  | Double | VitalSigns | IntelliVue MX800 | baFHR
1558 | Patient Activity Level (Movement) | - |  | Double | VitalSigns | IntelliVue MX800 | mAct
1559 | Patient Activity Level (Wearable) | - |  | Double | VitalSigns | IntelliVue MX800 | wAct
1569 | Minimum Airway Pressure | - |  | Double | VitalSigns | IntelliVue MX800 | AWPmin
1570 | P0.1 Occlusion Pressure | - |  | Double | VitalSigns | IntelliVue MX800 | VENT: Poccl
1573 | Patient CO2 Production | - | Volume of CO2 produced by respiration. | Double | VitalSigns | IntelliVue MX800 | VCO2
1574 | Relative Dead Space | - | Relative Dead Space Volume. | Double | VitalSigns | IntelliVue MX800 | VENT: relVd
1590 | Procalcitonin (PCT) | - |  | Double | VitalSigns | IntelliVue MX800 | PCT
1607 | Fetal Heart Rate (USound) | - | This variable is associated with ChannelID identifier. | Double | VitalSigns | IntelliVue MX800 | FHR/FHR1/FHR2/FHR3
1609 | Fetal Beat-to-Beat Rate (USound) | - | This variable is associated with ChannelID identifier. | Double | VitalSigns | IntelliVue MX800 | btbFHR/bFHR1/bFHR2/bFHR3
1630 | HFV Mean Pressure Setting | - |  | Double | Settings | IntelliVue MX800 | sHFMAP
1631 | HFV Tidal Volume Setting | - |  | Double | Settings | IntelliVue MX800 | sHFVTV
1645 | PAV-based Resistance Setting | - |  | Double | Settings | IntelliVue MX800 | sRppv
1646 | PAV-based Elastance Setting | - |  | Double | Settings | IntelliVue MX800 | sEppv
1703 | Abdominal TOCO | - |  | Double | VitalSigns | IntelliVue MX800 | aToco
1704 | Chest Movement Respiration Rate | - |  | Double | VitalSigns | IntelliVue MX800 | cmRR
1705 | Patient Movement Respiration Rate | - |  | Double | VitalSigns | IntelliVue MX800 | mRR
1706 | Patient Movement Pulse Rate | - |  | Double | VitalSigns | IntelliVue MX800 | mPulse
1707 | Respiration Signal Quality Index (Movement) | - |  | Double | OtherMeasuredData | IntelliVue MX800 | mSQI
1710 | Cardiac Power Index (CCO) | - |  | Double | VitalSigns | IntelliVue MX800 | CPI
1711 | Cardiac Power Output (CCO) | - |  | Double | VitalSigns | IntelliVue MX800 | CPO
1715 | Central Venous Pressure (Thermodilution) | - |  | Double | VitalSigns | IntelliVue MX800 | iCVP
1716 | Mean Arterial Pressure (Thermodilution) | - |  | Double | VitalSigns | IntelliVue MX800 | iARTm
1718 | Average Blood Temperature (Thermodilution) | - |  | Double | VitalSigns | IntelliVue MX800 | iAvgBT
1723 | O2 Delivery DO2 (Thermodilution) | - |  | Double | VitalSigns | IntelliVue MX800 | iDO2
1724 | O2 Consumption (Thermodilution) | - |  | Double | VitalSigns | IntelliVue MX800 | iVO2
1725 | O2 Consumption Index (Thermodilution) | - |  | Double | VitalSigns | IntelliVue MX800 | iVO2I
1733 | Global End-Diastolic Volume (Thermodilution) | - |  | Double | VitalSigns | IntelliVue MX800 | iGEDV
1734 | Global End-Diastolic Volum Index (Thermodilution) | - |  | Double | VitalSigns | IntelliVue MX800 | iGEDVI
1735 | Pulmonary Vascular Permeability Index (Thermodilution) | - |  | Double | VitalSigns | IntelliVue MX800 | iPVPI
1737 | Stroke Volume (Thermodilution) | - |  | Double | VitalSigns | IntelliVue MX800 | iSV
1738 | Stroke Volume Index (Thermodilution) | - |  | Double | VitalSigns | IntelliVue MX800 | iSI
1739 | O2 Consumption Estimate | - |  | Double | VitalSigns | IntelliVue MX800 | iVO2e
1745 | EEG1 Wave | - |  | Double | Waveforms | IntelliVue MX800 | Waves: EEG1
1746 | EEG2 Wave | - |  | Double | Waveforms | IntelliVue MX800 | Waves: EEG2
1747 | EEG3 Wave | - |  | Double | Waveforms | IntelliVue MX800 | Waves: EEG3
1748 | EEG4 Wave | - |  | Double | Waveforms | IntelliVue MX800 | Waves: EEG4
1751 | Spectral Edge Frequency | - |  | Double | VitalSigns | IntelliVue MX800 | EEG: SEF
1757 | Burst Suppression Ratio | - | This variable is associated with ChannelID identifier. | Double | VitalSigns | IntelliVue MX800 | BSR1/BSR2/BSR3/BSR4
1759 | Patient Energy Expenditure | - |  | Double | VitalSigns | IntelliVue MX800 | EE
1760 | Patient Respiratory Quotient | - |  | Double | VitalSigns | IntelliVue MX800 | RQ
1782 | Spectral Edge Frequency (CH3) | - |  | Double | VitalSigns | IntelliVue MX800 | EEG: Frequ3
1784 | Relative Power In Delta Band (CH3) | - |  | Double | VitalSigns | IntelliVue MX800 | Delta3
1785 | Relative Power In Theta Band (CH3) | - |  | Double | VitalSigns | IntelliVue MX800 | Theta3
1786 | Relative Power In Alpha Band (CH3) | - |  | Double | VitalSigns | IntelliVue MX800 | Alpha3
1787 | Relative Power In Beta Band (CH3) | - |  | Double | VitalSigns | IntelliVue MX800 | Beta3
1790 | Spectral Edge Frequency (CH4) | - |  | Double | VitalSigns | IntelliVue MX800 | SEF4
1792 | Relative Power In Delta Band (CH4) | - |  | Double | VitalSigns | IntelliVue MX800 | Delta4
1793 | Relative Power In Theta Band (CH4) | - |  | Double | VitalSigns | IntelliVue MX800 | Theta4
1794 | Relative Power In Alpha (CH4) | - |  | Double | VitalSigns | IntelliVue MX800 | Alpha4
1795 | Relative Power In Beta Band (CH4) | - |  | Double | VitalSigns | IntelliVue MX800 | Beta4
1911 | Device Unique ID | unitless |  | String | DeviceInfo | Common | System Information
1914 | Patient Bed ID | unitless | This variable contains a printable string identifying the system location. | String | PatientInfo | Common | System Information
1929 | Patient ID | unitless |  | String | PatientInfo | Common | Patient Demographics
1930 | Patient Name | unitless |  | String | PatientInfo | Common | Patient Demographics
1942 | Patient Comment Text | unitless |  | String (Multiset) | PatientInfo | Common | Patient Demographics
1955 | Patient Stress Index | - |  | Double | PatientInfo | IntelliVue MX800 | StrssI
1982 | Generic Patient Temperature | - |  | Double | VitalSigns | IntelliVue MX800 | Tbody
2049 | Breath Cycle Type | - |  | Double | DeviceInfo | IntelliVue MX800 | sCycTy
2077 | Airway Gas Flow Rate | - | Measured in the inspiratory path. | Double | VitalSigns | IntelliVue MX800 | VENT: AWFin
2092 | Inspired Minute Volume | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
2104 | Positive End Expiratory Pressure (PEEP) Setting | - |  | Double | Settings | IntelliVue MX800 | sEPAP
2163 | Spectral Edge Frequency - 1st EEG | - |  | Double | VitalSigns | IntelliVue MX800 | EEG: Frequ1
2167 | Bispectral Index | - | This variable is associated with ChannelID identifier. | Double | VitalSigns | IntelliVue MX800 | BIS L/BIS R
2168 | Total Power (Ch1) | - |  | Double | VitalSigns | IntelliVue MX800 | EEG: TP1
2169 | Spike Rate | - | This variable is associated with ChannelID identifier. | Double | VitalSigns | IntelliVue MX800 | BurstL/BurstR
2173 | Spectral Edge Frequency - 2nd EEG | - |  | Double | VitalSigns | IntelliVue MX800 | EEG: Frequ2
2178 | Total Power (Ch2) | - |  | Double | VitalSigns | IntelliVue MX800 | EEG: TP2
2182 | Narcotrend Index | - |  | Double | VitalSigns | IntelliVue MX800 | NI
2183 | Sharp Transient Intensity | - |  | Double | VitalSigns | IntelliVue MX800 | STI
2188 | Total Power Relative | - |  | Double | VitalSigns | IntelliVue MX800 | TPrel
2200 | Device Hardware Revision | unitless | This variable is obsolete because it may contains incorrect information. Please use variables 764, 765 and 4201 must be used instead. | String | DeviceInfo | Common | Obsolete Variables
2201 | Serial Number | unitless | This variable is obsolete because it may contains incorrect information. Please use variables 764, 765 and 4201 must be used instead. | String | DeviceInfo | Common | Obsolete Variables
2246 | O2 Flush Percentage Setting | - |  | Double | Settings | IntelliVue MX800 | sO2max
2250 | CO2 Tidal Production | - |  | Double | VitalSigns | IntelliVue MX800 | VCO2ti
2254 | Inspiratory Resistance | - |  | Double | VitalSigns | IntelliVue MX800 | Rinsp
2255 | Device Date/Time | unitless | Format: 16 ascii digits string of "YYYYMMDDhhmmssff", where the "ff" is second_fractions that is not used by IntelliVue monitors. | String | DeviceInfo | Common | System Information
2259 | Arterial BP Systolic | - |  | Double | VitalSigns | Common | IBP: ABP
2260 | Arterial BP Diastolic | - |  | Double | VitalSigns | Common | IBP: ABP
2261 | Arterial BP Mean | - |  | Double | VitalSigns | Common | IBP: ABP
2262 | Pulmonary Artery BP - Systolic | - |  | Double | VitalSigns | Common | IBP: PAP
2263 | Pulmonary Artery BP - Diastolic | - |  | Double | VitalSigns | Common | IBP: PAP
2264 | Pulmonary Artery BP Mean | - |  | Double | VitalSigns | Common | IBP: PAP
2271 | Right Atrial Pressure (RAP) | - | Available on IntelliVue devices with firmware revision G or later. | Double | VitalSigns | Common | IBP: RAP
2272 | Left Atrial Pressure (LAP) | - | Available on IntelliVue devices with firmware revision G or later. | Double | VitalSigns | Common | IBP: LAP
2273 | Intra Cranial Pressure (ICP) | - | Available on IntelliVue devices with firmware revision G or later. | Double | VitalSigns | Common | IBP: ICP
2317 | Pulmonary Artery Wedge Pressure | - |  | Double | VitalSigns | IntelliVue MX800 | IBP: PAWP
2321 | Generic NIBP Systolic Pressure | - | This variable is same as variable 2, except that the early estimated values are marked as NonSignificant in this variable. | Double | VitalSigns | Common | NBP
2322 | Generic NIBP Diastolic Pressure | - | This variable is same as variable 3, except that the early estimated values are marked as NonSignificant in this variable. | Double | VitalSigns | Common | NBP
2323 | Generic NIBP Mean Pressure | - | This variable is same as variable 4, except that the early estimated values are marked as NonSignificant in this variable. | Double | VitalSigns | Common | NBP
2324 | Inspired Tidal Volume | - |  | Double | VitalSigns | IntelliVue MX800 | TVin
2325 | Expired Tidal Volume | - |  | Double | VitalSigns | IntelliVue MX800 | TVexp
2326 | Expired Minute Volume | - |  | Double | VitalSigns | IntelliVue MX800 | MINVOL
2341 | Patient Blood Temperature | - |  | Double | VitalSigns | IntelliVue MX800 | Tblood
2342 | Injectate Temperature | - |  | Double | OtherMeasuredData | IntelliVue MX800 | Tinj
2343 | Inspired O2 | - |  | Double | VitalSigns | IntelliVue MX800 | VENT: FIO2
2350 | Inspired O2 Pressure | - | This variable is marked as obsolete because the label is incorrect. Please see variable 52523 for concentration and 52527 for partial pressure. | Double | VitalSigns | IntelliVue MX800 | O2
2351 | End Tidal O2 | - |  | Double | VitalSigns | IntelliVue MX800 | O2
2352 | CO2 (Gas Analyzer) | - |  | Double | VitalSigns | IntelliVue MX800 | CO2
2357 | End Tidal N2O | - |  | Double | VitalSigns | IntelliVue MX800 | N2O
2360 | End Tidal Agent Concentration | - |  | Double | VitalSigns | IntelliVue MX800 | AGT
2362 | End Tidal Desflurane | - |  | Double | VitalSigns | IntelliVue MX800 | DES
2364 | End Tidal Sevoflurane | - |  | Double | VitalSigns | IntelliVue MX800 | SEV
2366 | End Tidal Halothane | - |  | Double | VitalSigns | IntelliVue MX800 | HAL
2368 | End Tidal Enflurane | - |  | Double | VitalSigns | IntelliVue MX800 | ENF
2370 | End Tidal Isoflurane | - |  | Double | VitalSigns | IntelliVue MX800 | ISO
2378 | Fresh Gas N2O Flow | - | Mixer parameter. Setting of fresh gas N2O flow. | Double | DeviceInfo | IntelliVue MX800 | sfgN2O
2379 | Fresh Gas Air Flow | - |  | Double | DeviceInfo | IntelliVue MX800 | fgAir
2391 | Transcutaneous O2 Pressure | - |  | Double | VitalSigns | IntelliVue MX800 | TCUT: tcpO2
2392 | Transcutaneous CO2 Pressure | - |  | Double | VitalSigns | IntelliVue MX800 | TCUT: tcpCO2
2399 | Arterial O2 Saturation (SaO2) | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
2408 | EEG Total Power | - |  | Double | VitalSigns | IntelliVue MX800 | EEG: TP
2409 | Bispectral Index (BIS) | - |  | Double | VitalSigns | IntelliVue MX800 | EEG: BIS
2411 | EEG Burst Suppression Ratio | - |  | Double | VitalSigns | IntelliVue MX800 | EEG: BSI
2420 | Spontaneous Alveolar Minute Volume | - |  | Double | VitalSigns | IntelliVue MX800 | SpMVal
2421 | Airway Dead Space | - |  | Double | VitalSigns | IntelliVue MX800 | awVd
2432 | Expired Spontaneous Tidal Volume | - |  | Double | VitalSigns | IntelliVue MX800 | SpTVex
2434 | Spontaneous Expired Minute Volume | - |  | Double | VitalSigns | IntelliVue MX800 | %SpMV
2440 | Total Alveolar Tidal Volume | - |  | Double | VitalSigns | IntelliVue MX800 | TValv
2447 | Respiration Waveform | - |  | Double | Waveforms | IntelliVue MX800 | Waves: cmResp
2456 | EEG Signal Quality Index | - |  | Double | VitalSigns | IntelliVue MX800 | EEG: SQI
2517 | Wearable Respiration Rate | - |  | Double | VitalSigns | IntelliVue MX800 | wRR
2518 | Wearable Pulse Rate | - |  | Double | VitalSigns | IntelliVue MX800 | wPulse
2519 | Wearable Patient Temperature | - |  | Double | VitalSigns | IntelliVue MX800 | wTemp
2520 | Respiration Signal Quality Index (Wearable) | - |  | Double | OtherMeasuredData | IntelliVue MX800 | wSQI
2521 | Pulse Rate Signal Quality Index (Wearable) | - |  | Double | OtherMeasuredData | IntelliVue MX800 | wSQI
2542 | Airway Pressure Wave | - |  | Double | Waveforms | IntelliVue MX800 | Waves: AWP
2543 | Citrate Blood Concentration Setting | - |  | Double | Settings | IntelliVue MX800 | sCitDo
2547 | Neutrophil Count | - |  | Double | VitalSigns | IntelliVue MX800 | Bands
2580 | Fresh Gas Total Flow Setting | - | Mixer parameter. | Double | Settings | IntelliVue MX800 | sfgFl
2583 | I:E Ratio Setting | unitless | This variable is associated with ChannelID identifier. This variable is calculated from variable 2773, which is sent as a string in the format "1:xx.xx". | String | Settings | IntelliVue MX800 | Calculated Variables
2598 | Patient Admission Status | unitless | Possible values: "empty", "pre-admitted", "admitted", and "discharged". | String | PatientInfo | Common | Patient Demographics
2602 | ECG Lead V | - | Physiological identifier: NOM_ECG_AMPL_ST_V | Double | VitalSigns | Common | ECG: ST
2738 | Plateau Pressure | - | This variable contains parameter from Philips label "NLS_NOM_VENT_PRESS_RESP_PLAT" and observed value "NLS_NOM_VENT_PRESS_RESP_PLAT". | Double | VitalSigns | IntelliVue MX800 | VENT: PEinsp
2773 | I:E Ratio Setting | - |  | Double | Settings | IntelliVue MX800 | sIE 1:
2776 | Total PEEP | - |  | Double | VitalSigns | IntelliVue MX800 | VENT: tPEEP
2787 | APRV Low Time (Tlow) | - |  | Double | VitalSigns | IntelliVue MX800 | Tlow
2821 | System General Status | unitless | This variable identifies the current operating mode of the device. Possible values: "not specified", "monitoring", "demo", and "service". Moreover, "- standby" or "- config" can be appended | String | DeviceInfo | Common | System Information
2845 | Frequency | Hz | This variable describes the frequency of the main power supply. Possible values: "unspecified", "50", and "60". | String | DeviceInfo | Common | System Information
2855 | Device Status | unitless | Possible values: "disconnected", "unassociated", "operating". | String | DeviceInfo | Common | System Information
2879 | Device Software Revision | unitless | This variable is obsolete because it may contains incorrect information. Please use variables 764, 765 and 4201 must be used instead. | String | DeviceInfo | Common | Obsolete Variables
2880 | Maximum Apnea Time Setting | - |  | Double | DeviceInfo | IntelliVue MX800 | sTAmax
2889 | Inspired Agent Concentration | - |  | Double | VitalSigns | IntelliVue MX800 | inAGT
2898 | Patient Weight | - |  | Double | PatientInfo | IntelliVue MX800 | Weight
2901 | Patient First Name | unitless |  | String | PatientInfo | Common | Patient Demographics
2907 | Patient Temperature | - |  | Double | VitalSigns | Common | Temp
2915 | Ejection Fraction | - |  | Double | VitalSigns | IntelliVue MX800 | REF
2916 | End Diastolic Volume | - |  | Double | VitalSigns | IntelliVue MX800 | EDV
2917 | End Diastolic Volume Index | - |  | Double | VitalSigns | IntelliVue MX800 | EDVI
2918 | End Systolic Volume | - |  | Double | VitalSigns | IntelliVue MX800 | ESV
2919 | End Systolic Volume Index | - |  | Double | VitalSigns | IntelliVue MX800 | ESVI
2922 | EEG Electromyography (EMG) | - | This variable is associated with ChannelID identifier. | Double | VitalSigns | IntelliVue MX800 | EMG L/EMG R
2923 | Generic EEG Wave | - |  | Double | Waveforms | IntelliVue MX800 | Waves: EEG
2924 | ECG Pad Wave | - |  | Double | Waveforms | IntelliVue MX800 | Waves: Pads
2925 | ECG Paddle Wave | - |  | Double | Waveforms | IntelliVue MX800 | Waves: Paddle
2930 | Unspecific Pressure Pulse Rate | - |  | Double | VitalSigns | IntelliVue MX800 | Pulse
2932 | Patient Temperature 1 | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
2933 | Patient Temperature 2 | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
2954 | APRV High Pressure (Phigh) | - |  | Double | VitalSigns | IntelliVue MX800 | Phigh
2959 | APRV Low Pressure (Plow) | - |  | Double | VitalSigns | IntelliVue MX800 | Plow
2964 | I:E Ratio | unitless | This variable is associated with ChannelID identifier This variable is calculated from variable 1190 and sent as a string in the format "1:xx.xx". | String | VitalSigns | IntelliVue MX800 | Calculated Variables
2968 | Plateau Airway Pressure | - |  | Double | VitalSigns | IntelliVue MX800 | Pplat
2984 | Plateau Pressure Setting | - |  | Double | Settings | IntelliVue MX800 | sPplat
2991 | FiO2 Upper Alarm Limit | - |  | Double | Settings | IntelliVue MX800 | AlarmLimit: highO2
2992 | FiO2 Lower Alarm Limit | - |  | Double | Settings | IntelliVue MX800 | AlarmLimit: lowO2
2995 | Device Altitude | m |  | Double | DeviceInfo | Common | System Information
3003 | Blood Flow Rate Setting | - |  | Double | Settings | IntelliVue MX800 | sBFl
3005 | Dialysate Flow Rate Setting | - |  | Double | Settings | IntelliVue MX800 | sDialF
3007 | Effluent Flow Rate Setting | - |  | Double | Settings | IntelliVue MX800 | sEfflF
3024 | Access Pod Pressure | - |  | Double | OtherMeasuredData | IntelliVue MX800 | AccP
3025 | Return Pod Pressure | - |  | Double | OtherMeasuredData | IntelliVue MX800 | RetP
3026 | Filter Pod Pressure | - |  | Double | OtherMeasuredData | IntelliVue MX800 | FiltP
3027 | Effluent Pod Pressure | - |  | Double | OtherMeasuredData | IntelliVue MX800 | EfflP
3029 | Transmembrane Pressure | - |  | Double | VitalSigns | IntelliVue MX800 | TMP
3062 | Hb Calculation Offset | - |  | Double | VitalSigns | IntelliVue MX800 | OffsHb
3091 | Patient Time Since Last Ward Admission | - |  | Double | PatientInfo | IntelliVue MX800 | TWrdAd
3092 | Patient Total Admission Time | - |  | Double | PatientInfo | IntelliVue MX800 | WardAd
3101 | Patient Drainage Volume | - |  | Double | VitalSigns | IntelliVue MX800 | DrnVol
3168 | Fluid Removal Actual Rate | - |  | Double | VitalSigns | IntelliVue MX800 | Vrem/h
3204 | Sum of Minimum Alveolar Concentration (MAC) | - |  | Double | VitalSigns | IntelliVue MX800 | MAC
3266 | HFV Frequency Setting | - |  | Double | Settings | IntelliVue MX800 | sHFFrq
3268 | Ventilator Base Flow Setting | - |  | Double | Settings | IntelliVue MX800 | sfgFmn
3271 | Blood Pressure Difference | - |  | Double | VitalSigns | IntelliVue MX800 | IBP: DeltaP
3272 | Generic Blood Pressure Difference 1 | - |  | Double | VitalSigns | IntelliVue MX800 | IBP: DeltaP1
3273 | Generic Blood Pressure Difference 2 | - |  | Double | VitalSigns | IntelliVue MX800 | IBP: DeltaP2
3278 | Patient Height | - |  | Double | PatientInfo | IntelliVue MX800 | Height
3291 | Patient Temperature Setting | - |  | Double | Settings | IntelliVue MX800 | sTskin
3385 | Arterial O2 Content (CaO2) | - |  | Double | VitalSigns | IntelliVue MX800 | CaO2
3391 | Alveolar Arterial O2 Difference (AaDO2) | - |  | Double | VitalSigns | IntelliVue MX800 | AaDO2
3400 | Venous O2 Content (CvO2) | - |  | Double | VitalSigns | IntelliVue MX800 | CvO2
3415 | Ionized Calcium Concentration (Ca) | - |  | Double | VitalSigns | IntelliVue MX800 | iCa
3421 | Chloride Concentration (Cl-) | - |  | Double | VitalSigns | IntelliVue MX800 | Cl
3435 | Activated Clotting Time | - | Measured by the i-Stat module. | Double | VitalSigns | IntelliVue MX800 | ACT
3446 | Generic Temperature | - |  | Double | VitalSigns | IntelliVue MX800 | T1
3447 | Generic Temperature 2 | - |  | Double | VitalSigns | IntelliVue MX800 | T2
3448 | Generic Temperature 3 | - |  | Double | VitalSigns | IntelliVue MX800 | T3
3449 | Generic Temperature 4 | - |  | Double | VitalSigns | IntelliVue MX800 | T4
3474 | Expiratory Resistance | - |  | Double | VitalSigns | IntelliVue MX800 | Rexp
3490 | TcGas Site Timer | - |  | Double | OtherMeasuredData | IntelliVue MX800 | TCUT: SitTim
3491 | O2 Saturation (SO2) | - |  | Double | VitalSigns | IntelliVue MX800 | SaO2
3496 | TcGas Transducer Temperature | - |  | Double | DeviceInfo | IntelliVue MX800 | TCUT: SensrT
3502 | Peak Expiratory Flow Setting | - |  | Double | Settings | IntelliVue MX800 | s%PEF
3503 | Auto Flow Active Status | - |  | Double | DeviceInfo | IntelliVue MX800 | sFlCtr
3506 | Inspiratory Minimum CO2 | - |  | Double | VitalSigns | Common | CO2
3507 | Expired Tidal Volume Indexed by Body Weight | - |  | Double | VitalSigns | IntelliVue MX800 | TVe/kg
3555 | ATC Activation Status | - |  | Double | DeviceInfo | IntelliVue MX800 | sTubeC
3570 | Skin Temperature | - |  | Double | VitalSigns | Common | Tskin
3571 | Rectal Temperature | - |  | Double | VitalSigns | Common | Trect
3572 | Nasopharyngeal Temperature | - |  | Double | VitalSigns | Common | Tnaso
3573 | Esophageal Temperature | - |  | Double | VitalSigns | Common | Tesoph
3574 | Core Temperature | - |  | Double | VitalSigns | Common | Tcore
3575 | Arterial Temperature | - |  | Double | VitalSigns | Common | Tart
3576 | Venous Temperature | - |  | Double | VitalSigns | Common | Tven
3583 | Generic Core Temperature 1 | - |  | Double | VitalSigns | IntelliVue MX800 | T1Core
3584 | Generic Core Temperature 2 | - |  | Double | VitalSigns | IntelliVue MX800 | T2Core
3591 | Device Number | unitless | Device Model Number (example: "M8000") | String | DeviceInfo | Common | System Information
3635 | Drug Delivery Rate | - |  | Double | DeviceInfo | IntelliVue MX800 | sDlvRt
3637 | Drug Concentration | - |  | Double | PatientInfo | IntelliVue MX800 | sDConc
3639 | Manufacturers Name | unitless | "Philips" | String | DeviceInfo | Common | System Information
3640 | Pressure Alarm Delay | - |  | Double | DeviceInfo | IntelliVue MX800 | sPADel
3646 | Arterial Blood Pressure (ART) | - | Available on IntelliVue devices with firmware revision G or later. Arterial Blood Pressure (NOM_PRESS_BLD_ART). | Double | VitalSigns | Common | IBP: ART
3647 | Pulmonary Artery Blood Pressure (PA) | - | Available on IntelliVue devices with firmware revision G or later. | Double | VitalSigns | Common | IBP: PAP
3704 | Arterial/Main Pump Flow | - |  | Double | VitalSigns | IntelliVue MX800 | HLMfl
3710 | Slave Pump Flow | - |  | Double | VitalSigns | IntelliVue MX800 | SlvPfl
3715 | Suction Pump Flow | - |  | Double | VitalSigns | IntelliVue MX800 | SucPfl
3720 | Auxiliary Pump Flow | - |  | Double | VitalSigns | IntelliVue MX800 | AuxPfl
3723 | Auxiliary Pump Time Since Start of Current Plegia Session | - |  | Double | VitalSigns | IntelliVue MX800 | AxOnTi
3724 | Auxiliary Pump Time Since Stop of Plegia Session | - |  | Double | VitalSigns | IntelliVue MX800 | AxOffT
3725 | Auxiliary Pump Delivered Volume in Current Plegia Session | - |  | Double | VitalSigns | IntelliVue MX800 | AxDVol
3726 | Auxiliary Pump Total Volume | - |  | Double | VitalSigns | IntelliVue MX800 | AxTVol
3727 | Auxiliary Pump Total Plegia Time | - |  | Double | VitalSigns | IntelliVue MX800 | AxPlTi
3730 | Cardioplegia Main Pump Flow | - |  | Double | VitalSigns | IntelliVue MX800 | PlePfl
3733 | Cardioplegia Main Pump Delivered Volume of Current Plegia Session | - |  | Double | VitalSigns | IntelliVue MX800 | CpDVol
3734 | Cardioplegia Main Pump Total Volume | - |  | Double | VitalSigns | IntelliVue MX800 | CpTVol
3737 | Cardioplegia Slave Pump Flow | - |  | Double | VitalSigns | IntelliVue MX800 | SplPfl
3740 | Cardioplegia Slave Pump Delivered Volume in Current Plegia Session | - |  | Double | VitalSigns | IntelliVue MX800 | CsDVol
3741 | Cardioplegia Slave Pump Total Volume | - |  | Double | VitalSigns | IntelliVue MX800 | CsTVol
3742 | Cycle Flow Sensitivity Setting | - |  | Double | Settings | IntelliVue MX800 | sFlSeC
3761 | Hydrocarbon Plasma Concentration (HCO3) | - |  | Double | VitalSigns | IntelliVue MX800 | HCO3
3781 | Corpuscular Mean Hemoglobin Concentration | - |  | Double | VitalSigns | IntelliVue MX800 | MCHC
3785 | Arterial O2 Partial Pressure | - |  | Double | VitalSigns | IntelliVue MX800 | PaO2
3786 | Venous O2 Partial Pressure | - |  | Double | VitalSigns | IntelliVue MX800 | PvO2
3787 | Arterial CO2 Partial Pressure | - |  | Double | VitalSigns | IntelliVue MX800 | PaCO2
3789 | Arterial O2 Saturation (SaO2) | - | Obsolete variable that may be removed in a future release. The label of this variable is not correct. Please use variable 14 instead. | Double | VitalSigns | Common | Obsolete Variables
3793 | Hematocrit (HCT) | - | Post-Dialysis Hematocrit | Double | VitalSigns | IntelliVue MX800 | cHCTo
3804 | Respiratory Rate (Flow) | - |  | Double | VitalSigns | IntelliVue MX800 | RRawf
3805 | Respiratory Rate (CO2) | - |  | Double | VitalSigns | IntelliVue MX800 | RRCO2
3853 | Patient Type | unitless | Possible values: "unspecified", "adult", "pediatric", and "neonatal". | String | PatientInfo | Common | Patient Demographics
3854 | Patient Pace Maker Status | unitless | Possible values: "not paced", "paced gen". | String | PatientInfo | Common | Patient Demographics
3869 | Heart Rate Signal Quality | - |  | Double | OtherMeasuredData | IntelliVue MX800 | mSQI
3917 | Tracheal Pressure | - |  | Double | VitalSigns | IntelliVue MX800 | Ptrach
3921 | Ventilator Leakage | - |  | Double | VitalSigns | IntelliVue MX800 | Lksys
3961 | Plateau Time Setting | - | This variable is obsolete that may be removed in a future release. The label of this variable is incorrect, please use variable 9585 instead. | Double | Settings | IntelliVue MX800 | Obsolete Variables
3963 | Warmer Mattress Temperature | - |  | Double | OtherMeasuredData | IntelliVue MX800 | Tmatt
3968 | Spontaneous Fraction of Minute Volume | - |  | Double | VitalSigns | IntelliVue MX800 | SpMVe
4004 | Breathing System Leakage | - |  | Double | OtherMeasuredData | IntelliVue MX800 | VENT: Leak
4006 | Inspired MAC Concentration | - |  | Double | VitalSigns | IntelliVue MX800 | MAC
4008 | Inspiratory 2nd Agent | - |  | Double | VitalSigns | IntelliVue MX800 | inAGTs
4009 | Expiratory 2nd Agent | - | This is the device's measurement with label="NLS_NOM_EMFC_etAGTs" and physio_id="NOM_CONC_AWAY_AGENT_ET". | Double | VitalSigns | IntelliVue MX800 | etAGTs
4012 | Volume Assist Level Setting | - |  | Double | Settings | IntelliVue MX800 | sWOBva
4014 | Backup Tidal Volume Setting | - |  | Double | Settings | IntelliVue MX800 | sTVbak
4016 | Flow Acceleration Setting | - |  | Double | Settings | IntelliVue MX800 | sFlacc
4017 | Cumulated Halothane Consumption | - |  | Double | OtherMeasuredData | IntelliVue MX800 | AccHAL
4018 | Cumulated Enflurane Consumption | - |  | Double | OtherMeasuredData | IntelliVue MX800 | AccENF
4019 | Cumulated Isoflurane Consumption | - |  | Double | OtherMeasuredData | IntelliVue MX800 | AccISO
4020 | Cumulated Desflurane Consumption | - |  | Double | OtherMeasuredData | IntelliVue MX800 | AccDes
4021 | Cumulated Sevoflurane Consumption | - |  | Double | OtherMeasuredData | IntelliVue MX800 | AccSEV
4026 | Air Humidity Setting | - |  | Double | Settings | IntelliVue MX800 | sHum
4028 | Mattress Temperature Setting | - |  | Double | Settings | IntelliVue MX800 | sTmatt
4029 | Inspired Agent Setting | - |  | Double | Settings | IntelliVue MX800 | sInAGT
4030 | Expired Agent Setting | - | End-Tidal agent concentration | Double | Settings | IntelliVue MX800 | sEtAGT
4037 | Inspired Desflurane Setting | - |  | Double | Settings | IntelliVue MX800 | sInDES
4038 | Expired Desflurane Setting | - | End-Tidal desflurane concentration | Double | Settings | IntelliVue MX800 | sEtDES
4039 | Inspired Sevoflurane Setting | - |  | Double | Settings | IntelliVue MX800 | sInSEV
4040 | Expired Sevoflurane Setting | - | End-Tidal sevoflurane concentration | Double | Settings | IntelliVue MX800 | sEtSEV
4079 | Fresh Gas Sevoflurane | - |  | Double | VitalSigns | IntelliVue MX800 | fgSEV
4080 | Fresh Gas Desflurane | - |  | Double | VitalSigns | IntelliVue MX800 | fgDES
4081 | Ambient Temperature | - |  | Double | OtherMeasuredData | IntelliVue MX800 | Tamb
4082 | Ambient Pressure | - |  | Double | OtherMeasuredData | IntelliVue MX800 | PB
4084 | Pulse Rate ( Unspecified Source ) | - |  | Double | VitalSigns | Common | Pulse
4089 | Central Venous Pressure Diastolic | - |  | Double | VitalSigns | Common | IBP: CVP
4093 | Central Venous Mean Pressure | - |  | Double | VitalSigns | Common | IBP: CVP
4094 | Central Venous Pressure Systolic | - |  | Double | VitalSigns | Common | IBP: CVP
4120 | Expiratory Volume Wave | - |  | Double | Waveforms | IntelliVue MX800 | Waves: AWVex
4143 | Apnea Ventilation I:E Ratio Setting (Number) | - |  | Double | Settings | IntelliVue MX800 | sBIE1:
4145 | Dynamic Plateau Airway Pressure | - |  | Double | VitalSigns | IntelliVue MX800 | PplDyn
4158 | Transcutaneous Relative Heater Power | - |  | Double | OtherMeasuredData | IntelliVue MX800 | HeatPr
4171 | Spontaneous Respiratory Rate Setting | - |  | Double | Settings | IntelliVue MX800 | sRRmax
4201 | Product Description | unitless | Variables 764, 765 and 4201 must be interpreted conjointly. | String (Multiset) | DeviceInfo | Common | System Information
Consult detailed help for more information on this variable
4226 | Replacement Cumulated Volume | - |  | Double | VitalSigns | IntelliVue MX800 | Vrep/h
4227 | Dialysate Cumulated Volume | - | This variable is associated with ChannelID identifier. | Double | VitalSigns | IntelliVue MX800 | DialC/Vdly/h
4228 | Effluent Cumulated Volume | - | This variable is associated with ChannelID identifier. | Double | VitalSigns | IntelliVue MX800 | Veff/h / EfflC
4229 | Fluid Removal Cumulated Volume | - |  | Double | VitalSigns | IntelliVue MX800 | Vufr
4352 | Base Pressure | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
4369 | PEEP/CPAP Setting | - |  | Double | Settings | IntelliVue MX800 | VENT: sPEEP
4370 | Sigh Rate Setting | - |  | Double | Settings | IntelliVue MX800 | VENT: sSghR
4371 | Sigh Tidal Volume Setting | - |  | Double | Settings | IntelliVue MX800 | VENT: sSghTV
4375 | Data Averaging Interval Setting | - |  | Double | Settings | IntelliVue MX800 | Available only in revF or earlier
4376 | Chart Interval Setting | - |  | Double | Settings | IntelliVue MX800 | Available only in revF or earlier
4383 | Static Lung Compliance | - |  | Double | VitalSigns | IntelliVue MX800 | Cstat
4384 | Static Lung Resistance | - |  | Double | VitalSigns | IntelliVue MX800 | Raw
4385 | Dynamic Lung Compliance | - |  | Double | VitalSigns | IntelliVue MX800 | Cdyn
4386 | Dynamic Lung Resistance | - | NLS_NOM_RES_AWAY_DYN | Double | VitalSigns | IntelliVue MX800 | Rdyn
4388 | Apnea Inspiratory Pressure Setting | - |  | Double | Settings | IntelliVue MX800 | sIPApn/sAInP
4393 | Patient Hose Compliance | - |  | Double | OtherMeasuredData | IntelliVue MX800 | Chos
4401 | Oximeter Averaging Interval Setting | - |  | Double | Settings | IntelliVue MX800 | Available only in revF or earlier
4457 | System Alarm Status | unitless | Overall alert state of the system. | Double | Alarms | Common | Alarms
Consult detailed help for more information on this variable
4458 | Patient Alarm Priority | unitless | Overall Patient alarm priority. This is the current highest priority value of currently active Patient Alarms. | Double | Alarms | Common | Alarms
Consult detailed help for more information on this variable
4459 | Technical Alarm Priority | unitless | Overall Technical alarm priority. This is the current highest priority value of currently active Technical Alarms. | Double | Alarms | Common | Alarms
Consult detailed help for more information on this variable
4460 | Audio Alarm Priority | unitless | Overall audible alarm priority. This is the current highest priority value of currently active audible alarms. | Double | Alarms | Common | Alarms
Consult detailed help for more information on this variable
4461 | Technical Alarm Source ID | unitless |  | Double (Multiset) | Alarms | Common | Alarms
4462 | Technical Alarm Global State | unitless | Consult detailed help for more information on this variable | Double (Multiset) | Alarms | Common | Alarms
4463 | Technical Alarm Priority | unitless | Consult detailed help for more information on this variable | Double (Multiset) | Alarms | Common | Alarms
4464 | Technical Alarm Flag | unitless | Consult detailed help for more information on this variable | Double (Multiset) | Alarms | Common | Alarms
4465 | Technical Alarm Priority | unitless |  | Double (Multiset) | Alarms | Common | Alarms
4466 | Technical Alarm Description | unitless |  | String (Multiset) | Alarms | Common | Alarms
4467 | Patient Alarm Source ID | unitless |  | Double (Multiset) | Alarms | Common | Alarms
4468 | Patient Alarm Global State | unitless | Consult detailed help for more information on this variable | Double (Multiset) | Alarms | Common | Alarms
4469 | Patient Alarm Priority | unitless | Consult detailed help for more information on this variable | Double (Multiset) | Alarms | Common | Alarms
4470 | Patient Alarm Flag | unitless | Consult detailed help for more information on this variable | Double (Multiset) | Alarms | Common | Alarms
4471 | Patient Alarm Priority | unitless |  | Double (Multiset) | Alarms | Common | Alarms
4472 | Patient Alarm Description | unitless |  | String (Multiset) | Alarms | Common | Alarms
4473 | Device Nomenclature Revision | unitless | Format: major.minor (such as 1.0) | String | DeviceInfo | Common | System Information
4474 | Device Text Catalog Revision | unitless | Format: "MM.NN mm.nn", where "MM.NN" is text catalog revision and "mm.nn" is language revision. | String | DeviceInfo | Common | System Information
4475 | Device Data Format | unitless | Always 11 (Unicode). | String | DeviceInfo | Common | System Information
4476 | Device Operating Area | unitless | Possible values: "unspecified", "operating room", "intensive care unit", "neonatal intensive care unit", and "cardiology care unit". | String | DeviceInfo | Common | System Information
4496 | Patient O2 Consumption (VO2) | - |  | Double | VitalSigns | IntelliVue MX800 | VO2
4497 | Patient CO2 Production (VCO2) | - |  | Double | VitalSigns | IntelliVue MX800 | bbvCO2
4507 | Arterial O2 Saturation (Left) | - |  | Double | VitalSigns | IntelliVue MX800 | SpO2 l
4508 | Arterial O2 Saturation (Right) | - |  | Double | VitalSigns | IntelliVue MX800 | SpO2 r
4509 | SpO2 Difference | - |  | Double | VitalSigns | IntelliVue MX800 | DeltaSpO2
4510 | Perfusion Indicator | - |  | Double | VitalSigns | IntelliVue MX800 | Perf
4511 | Relative Perfusion (Left) | - |  | Double | VitalSigns | IntelliVue MX800 | Perf l
4512 | Relative Perfusion (Right) | - |  | Double | VitalSigns | IntelliVue MX800 | Perf r
4513 | Ao Systolic | - |  | Double | VitalSigns | Common | IBP: Ao
4514 | Ao Diastolic | - |  | Double | VitalSigns | Common | IBP: Ao
4515 | Ao Mean | - |  | Double | VitalSigns | Common | IBP: Ao
4516 | RAP Systolic | - |  | Double | VitalSigns | Common | IBP: RAP
4517 | RAP Diastolic | - |  | Double | VitalSigns | Common | IBP: RAP
4518 | RAP Mean | - |  | Double | VitalSigns | Common | IBP: RAP
4519 | Intracranial Pressure ICP Systolic | - |  | Double | VitalSigns | Common | IBP: ICP
4520 | Intracranial Pressure ICP Diastolic | - |  | Double | VitalSigns | Common | IBP: ICP
4521 | ICP Mean | - |  | Double | VitalSigns | Common | IBP: ICP
4522 | UAP Systolic | - |  | Double | VitalSigns | Common | IBP: UAP
4523 | UAP Diastolic | - |  | Double | VitalSigns | Common | IBP: UAP
4524 | UAP Mean | - |  | Double | VitalSigns | Common | IBP: UAP
4525 | UVP Systolic | - |  | Double | VitalSigns | Common | IBP: UVP
4526 | UVP Diastolic | - |  | Double | VitalSigns | Common | IBP: UVP
4527 | UVP Mean | - |  | Double | VitalSigns | Common | IBP: UVP
4528 | Unspecific Pressure Systolic | - |  | Double | VitalSigns | Common | IBP: P
4529 | Unspecific Pressure Diastolic | - |  | Double | VitalSigns | Common | IBP: P
4530 | Unspecific Pressure Mean | - |  | Double | VitalSigns | Common | IBP: P
4531 | Cerebral Perfusion Pressure (CPP) | - |  | Double | VitalSigns | IntelliVue MX800 | IBP: CPP
4532 | Stroke Volume Variation | - |  | Double | VitalSigns | IntelliVue MX800 | SVV
4533 | Index of Left Ventricular Contractility | - | This variable contains measurement from label "NLS_NOM_GRAD_PRESS_BLD_AORT_POS_MAX" and observed value "NOM_GRAD_PRESS_BLD_AORT_POS_MAX". | Double | VitalSigns | IntelliVue MX800 | dPmax
4556 | Aortic Blood Pressure (Ao) Wave | - |  | Double | Waveforms | IntelliVue MX800 | Waves: Ao
4560 | Auxiliary Waveform | - |  | Double | Waveforms | IntelliVue MX800 | Waves: AUX
4561 | Intermittent Mandatory Ventilation Status | - |  | Double | VitalSigns | IntelliVue MX800 | VENT: IMV
4562 | Left Cardiac Work | - |  | Double | VitalSigns | IntelliVue MX800 | LCW
4564 | Right Cardiac Work | - |  | Double | VitalSigns | IntelliVue MX800 | RCW
4565 | Serum Creatinine | - |  | Double | VitalSigns | IntelliVue MX800 | SCrea
4566 | Urine Plasma Osmolarity Ratio | - |  | Double | VitalSigns | IntelliVue MX800 | U/POsm
4567 | Urine Sodium Potassium Ratio | - |  | Double | VitalSigns | IntelliVue MX800 | UrNa/K
4568 | Urine Volume | - |  | Double | VitalSigns | IntelliVue MX800 | UrVol
4574 | Left Cardiac Work Index | - |  | Double | VitalSigns | IntelliVue MX800 | LCWI
4576 | Right Cardiac Work Index | - |  | Double | VitalSigns | IntelliVue MX800 | RCWI
4577 | Light Intensity (SvO2 Module) | - |  | Double | VitalSigns | IntelliVue MX800 | LI
4578 | SpO2 SvO2 Difference | - |  | Double | VitalSigns | IntelliVue MX800 | Sp-vO2
4581 | Normalized Post Replacement Volume | - |  | Double | VitalSigns | IntelliVue MX800 | VdpoNh
4582 | Normalized Effluent Volume | - | This variable is associated with ChannelID identifier. | Double | VitalSigns | IntelliVue MX800 | Veff/EffDn
4584 | Normalized Total Replacement Volume (PBP) | - | This variable is associated with ChannelID identifier. | Double | VitalSigns | IntelliVue MX800 | VrePN/VrePNh
4585 | Normalized Replacement Volume (Pre) | - |  | Double | VitalSigns | IntelliVue MX800 | VdprNh
4586 | Normalized Ultrafiltration Volume | - | This variable is associated with ChannelID identifier. | Double | VitalSigns | IntelliVue MX800 | VufrN/VufrNh
4587 | Patient Fluid Balance Limit Setting | - |  | Double | Settings | IntelliVue MX800 | sEPFL
4588 | Total Replacement Volume (PBP) | - | This variable is associated with ChannelID identifier. | Double | VitalSigns | IntelliVue MX800 | VreP/VreP/h
4589 | Cumulated Liquid Volume (Pre) | - | This variable is associated with ChannelID identifier. | Double | VitalSigns | IntelliVue MX800 | Vdpr/Vdpr/h
4593 | Filtration Fraction Setting | - |  | Double | Settings | IntelliVue MX800 | sFufrP
4616 | Diff.-Temperature between Tblood and Tskin | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
4617 | Diff.-Temperature between Tblood and Trect | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
4618 | Diff.-Temperature between Tblood and Tnaso | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
4619 | Diff.-Temperature between Tblood and Tesop | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
4620 | Diff.-Temperature between Tblood and Tcore | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
4621 | Diff.-Temperature between Tblood and Tart | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
4622 | Diff.-Temperature between Tblood and Tven | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
4623 | Diff.-Temperature between Tskin and Trect | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
4624 | Diff.-Temperature between Tskin and Tnaso | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
4625 | Diff.-Temperature between Tskin and Tesop | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
4626 | Diff.-Temperature between Tskin and Tcore | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
4627 | Diff.-Temperature between Tskin and Tart | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
4628 | Diff.-Temperature between Tskin and Tven | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
4629 | Diff.-Temperature between Trect and Tnaso | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
4630 | Diff.-Temperature between Trect and Tesop | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
4631 | Diff.-Temperature between Trect and Tcore | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
4632 | Diff.-Temperature between Trect and Tart | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
4633 | Diff.-Temperature between Trect and Tven | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
4634 | Diff.-Temperature between Tnaso and Tesop | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
4635 | Diff.-Temperature between Tnaso and Tcore | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
4636 | Diff.-Temperature between Tnaso and Tart | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
4637 | Diff.-Temperature between Tnaso and Tven | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
4638 | Diff.-Temperature between Tesop and Tcore | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
4639 | Diff.-Temperature between Tesop and Tart | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
4640 | Diff.-Temperature between Tesop and Tven | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
4641 | Diff.-Temperature between Tcore and Tart | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
4642 | Diff.-Temperature between Tcore and Tven | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
4643 | Diff.-Temperature between Tart and Tven | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
4644 | Newborn Body Length | - |  | Double | VitalSigns | IntelliVue MX800 | Length
4646 | Newborn Gestational Age | - |  | Double | PatientInfo | IntelliVue MX800 | G.Age
4647 | Calories | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
4648 | Input/Output Totals | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
4649 | Newborn Apgar Scores | - |  | Double | PatientInfo | IntelliVue MX800 | Available only in revF or earlier
4650 | BSA Formula Boyd | - |  | Double | PatientInfo | IntelliVue MX800 | BSA(B)
4651 | BSA Formula Dubois | - |  | Double | PatientInfo | IntelliVue MX800 | BSA(D)
4653 | Pupil Reaction (Left Eye) | - | Light reaction of left eye's pupil. | Double | VitalSigns | IntelliVue MX800 | PRL
4654 | Pupil Reaction (Right Eye) | - | Light reaction of right eye's pupil. | Double | VitalSigns | IntelliVue MX800 | PRR
4661 | Pleth Gain | - |  | Double | DeviceInfo | IntelliVue MX800 | PlGain
4662 | Airway Temperature | - |  | Double | VitalSigns | IntelliVue MX800 | Tairwy
4663 | Ventilation Perfusion Index | - |  | Double | VitalSigns | IntelliVue MX800 | VENT: VQI
4664 | O2 Extraction Index | - |  | Double | VitalSigns | IntelliVue MX800 | O2EI
4665 | Duration Above Base Pressure | - |  | Double | VitalSigns | IntelliVue MX800 | VENT: DPosP
4666 | Inspiration Rise Time | - |  | Double | VitalSigns | IntelliVue MX800 | VENT: RiseTi
4667 | Continuous Positive Airway Pressure (CPAP) | - |  | Double | VitalSigns | IntelliVue MX800 | CPAP
4668 | HFV Amplitude | - |  | Double | VitalSigns | IntelliVue MX800 | VENT: HFVAmp
4669 | Urine Urea | - |  | Double | VitalSigns | IntelliVue MX800 | UrUrea
4670 | Urine pH Value | - |  | Double | VitalSigns | IntelliVue MX800 | UrpH
4671 | Total of CO2 (tCO2) | - | Result of Blood gas Analysis. | Double | VitalSigns | IntelliVue MX800 | tCO2
4672 | Total Bilirubin | - |  | Double | VitalSigns | IntelliVue MX800 | tBili
4673 | Serum Glucose | - |  | Double | VitalSigns | IntelliVue MX800 | SerGlu
4674 | Urine Glucose | - |  | Double | VitalSigns | IntelliVue MX800 | UrGlu
4675 | Direct Bilirubin | - |  | Double | VitalSigns | IntelliVue MX800 | dBili
4676 | Serum Calcium | - |  | Double | VitalSigns | IntelliVue MX800 | SerCa
4677 | Total Serum Calcium | - |  | Double | VitalSigns | IntelliVue MX800 | tSerCa
4678 | Serum Magnesium | - |  | Double | VitalSigns | IntelliVue MX800 | SerMg
4679 | Serum Phosphat | - |  | Double | VitalSigns | IntelliVue MX800 | SerPho
4680 | Serum Kalium (Potassium) | - |  | Double | VitalSigns | IntelliVue MX800 | SerK
4681 | Serum Clorid | - |  | Double | VitalSigns | IntelliVue MX800 | SerCl
4682 | Serum Albumine | - |  | Double | VitalSigns | IntelliVue MX800 | SerAlb
4683 | Urine Clorid | - |  | Double | VitalSigns | IntelliVue MX800 | UrCl
4684 | Serum Globulin | - |  | Double | VitalSigns | IntelliVue MX800 | SerGlo
4685 | Serum Protein | - |  | Double | VitalSigns | IntelliVue MX800 | SerPro
4686 | Serum Urea | - |  | Double | VitalSigns | IntelliVue MX800 | SrUrea
4687 | Partial Thromboplastin Time | - |  | Double | VitalSigns | IntelliVue MX800 | PTT
4688 | Prothrombin Time | - |  | Double | VitalSigns | IntelliVue MX800 | PT
4689 | Thrombin Time | - |  | Double | VitalSigns | IntelliVue MX800 | TT
4690 | Thromboplastine Time (Quick Time) | - |  | Double | VitalSigns | IntelliVue MX800 | Quick
4691 | Alkalische Phosphatase | - |  | Double | VitalSigns | IntelliVue MX800 | AP
4692 | Alpha Amylase | - |  | Double | VitalSigns | IntelliVue MX800 | alphaA
4693 | Cholesterinesterase | - |  | Double | VitalSigns | IntelliVue MX800 | CHE
4694 | Creatinin Kinase | - |  | Double | VitalSigns | IntelliVue MX800 | SerCK
4695 | Creatine Kinase of Type Muscle-Brain | - |  | Double | VitalSigns | IntelliVue MX800 | CK-MB
4696 | Creatine Kinase of Type Muscle | - |  | Double | VitalSigns | IntelliVue MX800 | CK-MM
4697 | Gamma Glutamyl Transpeptidase | - |  | Double | VitalSigns | IntelliVue MX800 | GGT
4698 | Glutamic Oxaloacetic Transaminase | - |  | Double | VitalSigns | IntelliVue MX800 | GOT
4699 | Glutamic Pyruvic Transaminase | - |  | Double | VitalSigns | IntelliVue MX800 | GPT
4700 | Cholesterin | - |  | Double | VitalSigns | IntelliVue MX800 | Chol
4701 | Triglyzeride | - |  | Double | VitalSigns | IntelliVue MX800 | TGL
4702 | Urine Protein | - | (Total) Protein in Urine | Double | VitalSigns | IntelliVue MX800 | UrPro
4703 | Urine Calcium | - |  | Double | VitalSigns | IntelliVue MX800 | UrCa
4704 | Carboxy Hemoglobin | - |  | Double | VitalSigns | IntelliVue MX800 | CO-Hb
4705 | Total Protein | - |  | Double | VitalSigns | IntelliVue MX800 | tPro
4706 | Lactate Dehydrogenase | - |  | Double | VitalSigns | IntelliVue MX800 | LDH
4707 | Aspartin Aminotransferase | - |  | Double | VitalSigns | IntelliVue MX800 | AST
4708 | Alveolarproteinose (Alkaline Phosphatase) | - |  | Double | VitalSigns | IntelliVue MX800 | ALP
4709 | Reticulocyte Count | - |  | Double | VitalSigns | IntelliVue MX800 | RC
4710 | Coagulation Time | - |  | Double | VitalSigns | IntelliVue MX800 | CT
4711 | Erithrocyte Sedimentation Rate | - |  | Double | VitalSigns | IntelliVue MX800 | ESR
4712 | Packed Cell Volume | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
4713 | Kaolin Cephalin Time | - |  | Double | VitalSigns | IntelliVue MX800 | KCT
4715 | Minimum Pressure During Breathing Cycle | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
4716 | Infusion Pump Accumulated Volume | - |  | Double | OtherMeasuredData | IntelliVue MX800 | AccVol
4717 | Transthoracic Impedance | - |  | Double | VitalSigns | Common | T.I.
4718 | Inspired Expired N2O Difference | - |  | Double | VitalSigns | IntelliVue MX800 | VENT: i-eN2O
4719 | Inspired Expired Halothane Difference | - |  | Double | VitalSigns | IntelliVue MX800 | VENT: i-eHAL
4720 | Inspired Expired Enfluran Difference | - |  | Double | VitalSigns | IntelliVue MX800 | VENT: i-eENF
4721 | Inspired Expired Isofluran Difference | - |  | Double | VitalSigns | IntelliVue MX800 | VENT: i-eISO
4722 | Inspired Expired Sevofluran Difference | - |  | Double | VitalSigns | IntelliVue MX800 | VENT: i-eSEV
4723 | Inspired Expired Desfluran Difference | - |  | Double | VitalSigns | IntelliVue MX800 | VENT: i-eDES
4724 | Inspired Expired Agent Difference | - |  | Double | VitalSigns | IntelliVue MX800 | VENT: i-eAGT
4725 | Inspired Expired O2 Difference | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
4726 | Patient Circuit O2 | - | O2 measured in the Patient Circuit | Double | VitalSigns | IntelliVue MX800 | VENT: cktO2
4727 | Barometric Pressure in Sample Chamber | - | This variable is obsolete that may be removed in a future release. It is duplicated with variable 4082. | Double | OtherMeasuredData | IntelliVue MX800 | Obsolete Variables
4728 | Minute Volume | - |  | Double | VitalSigns | IntelliVue MX800 | MINVOL
4729 | Mean Airway Pressure | - |  | Double | VitalSigns | IntelliVue MX800 | MnAwP
4730 | Airway Respiration Rate | - |  | Double | VitalSigns | Common | awRR
4731 | Inspiratory Airway Pressure Waveform | - |  | Double | Waveforms | IntelliVue MX800 | Waves: AWPin
4732 | Inspiratory Flow Waveform | - |  | Double | Waveforms | IntelliVue MX800 | Waves: AWFin
4733 | HFV Inspired Mandatory Minute Volume | - |  | Double | VitalSigns | IntelliVue MX800 | HFMVin
4734 | HFV Gas Transport Coefficient Value | - |  | Double | VitalSigns | IntelliVue MX800 | DCO2
4735 | HFV Inspired Tidal Volume | - |  | Double | VitalSigns | IntelliVue MX800 | VENT: HFTVin
4736 | HFV Tidal Volume | - |  | Double | VitalSigns | IntelliVue MX800 | VENT: HFVTV
4738 | Generic Heart Rate | - | Heart Rate received from an external device | Double | VitalSigns | IntelliVue MX800 | extHR
4739 | PERF from the 2nd SpO2/PLETH module | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
4741 | Percent Alpha (Left) | - |  | Double | VitalSigns | IntelliVue MX800 | EEG: LT %AL
4742 | Percent Beta (Left) | - |  | Double | VitalSigns | IntelliVue MX800 | EEG: LT %BE
4745 | Absolute Alpha (Left) | - |  | Double | VitalSigns | IntelliVue MX800 | EEG: LT AL
4746 | Absolute Beta (Left) | - |  | Double | VitalSigns | IntelliVue MX800 | EEG: LT BE
4747 | Absolute Delta (Left) | - |  | Double | VitalSigns | IntelliVue MX800 | EEG: LT DL
4748 | Absolute Theta (Left) | - |  | Double | VitalSigns | IntelliVue MX800 | EEG: LT TH
4749 | Mean Dominant Frequency (Left) | - |  | Double | VitalSigns | IntelliVue MX800 | EEG: LT MDF
4750 | Median Power Frequency (Left) | - |  | Double | VitalSigns | IntelliVue MX800 | EEG: LT MPF
4751 | Peak Power Frequency (Left) | - |  | Double | VitalSigns | IntelliVue MX800 | EEG: LT PPF
4752 | Spectral Edge Frequency (Left) | - |  | Double | VitalSigns | IntelliVue MX800 | EEG: LT SEF
4753 | Total Power (Left) | - |  | Double | VitalSigns | IntelliVue MX800 | EEG: LT TP
4754 | Scale of the Left Channel EEG Wave | - |  | Double | VitalSigns | IntelliVue MX800 | EEG: LSCALE
4755 | Percent Alpha (Right) | - |  | Double | VitalSigns | IntelliVue MX800 | EEG: RT %AL
4756 | Percent Beta (Right) | - |  | Double | VitalSigns | IntelliVue MX800 | EEG: RT %BE
4759 | Absolute Alpha (Right) | - |  | Double | VitalSigns | IntelliVue MX800 | EEG: RT AL
4760 | Absolute Beta (Right) | - |  | Double | VitalSigns | IntelliVue MX800 | EEG: RT BE
4761 | Absolute Delta (Right) | - |  | Double | VitalSigns | IntelliVue MX800 | EEG: RT DL
4762 | Absolute Theta (Right) | - |  | Double | VitalSigns | IntelliVue MX800 | EEG: RT TH
4763 | Mean Dominant Frequency (Right) | - |  | Double | VitalSigns | IntelliVue MX800 | EEG: RT MDF
4764 | Median Power Frequency (Right) | - |  | Double | VitalSigns | IntelliVue MX800 | EEG: RT MPF
4765 | Peak Power Frequency (Right) | - |  | Double | VitalSigns | IntelliVue MX800 | EEG: RT PPF
4766 | Spectral Edge Frequency (Right) | - |  | Double | VitalSigns | IntelliVue MX800 | EEG: RT SEF
4767 | Total Power (Right) | - |  | Double | VitalSigns | IntelliVue MX800 | EEG: RT TP
4768 | Scale of the Right Channel EEG Wave | - |  | Double | VitalSigns | IntelliVue MX800 | EEG: RSCALE
4769 | Duration of Positive Pressure | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
4772 | Fresh Gas N2O Concentration | - |  | Double | VitalSigns | IntelliVue MX800 | fgN2O
4774 | Liquid Level (Anesthetic Agent Bottle) | - |  | Double | OtherMeasuredData | IntelliVue MX800 | AGTLev
4775 | Liquid Level (ISOflurane Bottle) | - |  | Double | OtherMeasuredData | IntelliVue MX800 | ISOLev
4776 | Liquid Level (ENFlurane Bottle) | - |  | Double | OtherMeasuredData | IntelliVue MX800 | ENFLev
4777 | Liquid Level (HALothane Bottle) | - |  | Double | OtherMeasuredData | IntelliVue MX800 | HALLev
4778 | Liquid Level (DESflurane Bottle) | - |  | Double | OtherMeasuredData | IntelliVue MX800 | DESLev
4779 | Liquid Level (SEVoflurane Bottle) | - |  | Double | OtherMeasuredData | IntelliVue MX800 | SEVLev
4780 | Invasive Blood Pressure Wave | - |  | Double | Waveforms | IntelliVue MX800 | Waves: P
4781 | Urine Shift Volume | - |  | Double | VitalSigns | IntelliVue MX800 | UrVSht
4782 | Urine Flow Rate | - |  | Double | VitalSigns | IntelliVue MX800 | UrFl
4786 | Mean Dominant Frequency (CH1) | - |  | Double | VitalSigns | IntelliVue MX800 | EEG: Frequ1
4787 | Mean Dominant Frequency (CH2) | - |  | Double | VitalSigns | IntelliVue MX800 | EEG: Frequ2
4788 | Peak Power Frequency (CH1) | - |  | Double | VitalSigns | IntelliVue MX800 | EEG: Frequ1
4789 | Peak Power Frequency (CH2) | - |  | Double | VitalSigns | IntelliVue MX800 | EEG: Frequ2
4794 | % of Power in the Alpha Band (CH1) | - |  | Double | VitalSigns | IntelliVue MX800 | EEG: Prcnt1
4795 | % of Power in the Alpha Band (CH2) | - |  | Double | VitalSigns | IntelliVue MX800 | EEG: Prcnt2
4796 | % of Power in the Beta Band (1st EEG) | - |  | Double | VitalSigns | IntelliVue MX800 | EEG: Prcnt1
4797 | % of Power in the Beta Band (2nd EEG) | - |  | Double | VitalSigns | IntelliVue MX800 | EEG: Prcnt2
4800 | Hemoglobin | - |  | Double | VitalSigns | IntelliVue MX800 | Hb
4801 | Patient Temperature | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
4802 | Capillaries pH Value | - |  | Double | VitalSigns | IntelliVue MX800 | pHc
4803 | Adjusted pH (at Patient Temperature) | - | Adjusted pH at Patient Temperature | Double | VitalSigns | IntelliVue MX800 | &pH
4804 | Adjusted Capillaries pH Value | - |  | Double | VitalSigns | IntelliVue MX800 | &pHc
4805 | Partial O2 in the Capillaries | - |  | Double | VitalSigns | IntelliVue MX800 | PcO2
4806 | Adjusted PO2 (at Patient Temperature) | - | Adjusted PO2 at Patient Temperature | Double | VitalSigns | IntelliVue MX800 | &PO2
4807 | Adjusted PcO2 (at Patient Temperature) | - |  | Double | VitalSigns | IntelliVue MX800 | &PcO2
4808 | Venous CO2 Partial Pressure (PvCO2) | - |  | Double | VitalSigns | IntelliVue MX800 | PvCO2
4809 | Partial CO2 in the Capillaries | - |  | Double | VitalSigns | IntelliVue MX800 | PcCO2
4811 | Adjusted PcCO2 (at Patient Temperature) | - |  | Double | VitalSigns | IntelliVue MX800 | &PcCO2
4812 | Calculated Total CO2 (tCO2) | - |  | Double | VitalSigns | IntelliVue MX800 | CALC: 'tCO2
4813 | Calculated SO2 | - | Calculated SaO2 | Double | VitalSigns | IntelliVue MX800 | CALC: 'SaO2
4815 | Calculated SvO2 | - |  | Double | VitalSigns | IntelliVue MX800 | CALC: 'SvO2
4816 | Calculated ScO2 | - |  | Double | VitalSigns | IntelliVue MX800 | CALC: 'ScO2
4817 | Calculated HCO3 | - |  | Double | VitalSigns | IntelliVue MX800 | CALC: HCO3
4818 | Base Excess of Extra Cellular Fluid | - |  | Double | VitalSigns | IntelliVue MX800 | CALC: Beecf
4819 | Calculated Anion Gap | - |  | Double | VitalSigns | IntelliVue MX800 | CALC: AnGap
4821 | Urea | - | Measured by the i-Stat module. | Double | VitalSigns | IntelliVue MX800 | Urea
4822 | Manually Entered FiO2 | - | In the Blood Review window. | Double | Settings | IntelliVue MX800 | Available only in revF or earlier
4823 | Calculated Arterial Blood Base Excess | - |  | Double | VitalSigns | IntelliVue MX800 | BE,B
4824 | Ionized Magnesium | - |  | Double | VitalSigns | IntelliVue MX800 | iMg
4826 | Desired Module Temperature | - | It's the selected/adjusted Temp in i-Stat Window. | Double | Settings | IntelliVue MX800 | Available only in revF or earlier
4827 | Calculated BUN/Creatinine Ratio | - | Calculated by the i-Stat module. | Double | VitalSigns | IntelliVue MX800 | B/Cre
4828 | Calculated Urea/Creatinine Ratio | - | Calculated by the i-Stat module. | Double | VitalSigns | IntelliVue MX800 | U/Cre
4829 | Lactate | - | Measured value by the i-Stat module. | Double | VitalSigns | IntelliVue MX800 | Lact
4830 | Device Simulator Status | - | Device data with the values: Passed and Failed. | Double | DeviceInfo | IntelliVue MX800 | Available only in revF or earlier
4835 | Pressure of Sensor 1 | - | Pressure value of sensor 1. Device specific data. | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
4836 | Pressure of Sensor 2 | - | Pressure value of sensor 2. Device specific data. | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
4837 | Module Temperature | - |  | Double | OtherMeasuredData | IntelliVue MX800 | Available only in revF or earlier
4843 | Air Temperature | - |  | Double | OtherMeasuredData | IntelliVue MX800 | Available only in revF or earlier
4844 | Micro Environment Air Temperature Setting | - |  | Double | Settings | IntelliVue MX800 | sTemp
4847 | Weight of the Urine Disposable Bag | - |  | Double | VitalSigns | IntelliVue MX800 | BagWgt
4848 | Total Urine Volume of the Current Measurement Period | - |  | Double | VitalSigns | IntelliVue MX800 | tUrVol
4849 | Urine Fluid Density | - |  | Double | VitalSigns | IntelliVue MX800 | UrDens
4850 | Urine Fluid Temperature | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
4854 | Ventilation Activated Status | - | Parameter to Informs whether the Patient is Ventilated | Double | VitalSigns | IntelliVue MX800 | VENT: PtVent
4856 | Simplified Acuity Physiologic Score | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
4857 | Value Normal | - | Missing parameters in SAPS parameter set which are considered in normal range. | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
4858 | Aspartate Aminotransferase | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
4859 | Spontaneous Respiration Rate | - |  | Double | VitalSigns | IntelliVue MX800 | SpRR
4864 | Current Fluid in Urine Bag | - |  | Double | VitalSigns | IntelliVue MX800 | BagVol
4865 | Cardiac Function Index (Thermodilution) | - |  | Double | VitalSigns | IntelliVue MX800 | iCFI
4866 | Intrathoracic Blood Volume (Thermodilution) | - |  | Double | VitalSigns | IntelliVue MX800 | iITVB
4867 | Intrathoracic Blood Volume Index (Thermodilution) | - |  | Double | VitalSigns | IntelliVue MX800 | iITBVI
4868 | Extravascular Lung Water (Thermodilution) | - |  | Double | VitalSigns | IntelliVue MX800 | iEVLW
4869 | Extravascular Lung Water Index (Thermodilution) | - |  | Double | VitalSigns | IntelliVue MX800 | iEVLWI
4873 | NMT TOF Count | - |  | Double | VitalSigns | IntelliVue MX800 | NMT: TOFcnt
4874 | NMT TOF Ratio | - |  | Double | VitalSigns | IntelliVue MX800 | NMT: TOFrat
4875 | Twitch Height of the 1Hz/0.1Hz Stimulation Response | - |  | Double | VitalSigns | IntelliVue MX800 | NMT: Twitch
4876 | Post Tetatic Count | - |  | Double | VitalSigns | IntelliVue MX800 | NMT: PTC
4877 | Remaining Time to Next Stimulation | - |  | Double | VitalSigns | IntelliVue MX800 | NMT: RemTi
4878 | TOF Response TOF1 | - |  | Double | VitalSigns | IntelliVue MX800 | NMT: TOF1
4879 | TOF Response TOF2 | - |  | Double | VitalSigns | IntelliVue MX800 | NMT: TOF2
4880 | TOF Response TOF3 | - |  | Double | VitalSigns | IntelliVue MX800 | NMT: TOF3
4881 | TOF Response TOF4 | - |  | Double | VitalSigns | IntelliVue MX800 | NMT: TOF4
4883 | Activated Partial Thromboplastin Time (aPTT Whole Blood Time) | - |  | Double | VitalSigns | IntelliVue MX800 | aPTTWB
4884 | Activated Partial Thromboplastin Time (aPTT Plasma Equivalent Time) | - |  | Double | VitalSigns | IntelliVue MX800 | aPTTPE
4886 | Prothrombin Time (Whole Blood Time) | - |  | Double | VitalSigns | IntelliVue MX800 | PT WB
4887 | Prothrombin Time (Plasma Equivalent Time) | - |  | Double | VitalSigns | IntelliVue MX800 | PT PE
4888 | Prothrombin Time Ratio | - |  | Double | VitalSigns | IntelliVue MX800 | PTrat
4889 | Spontaneous Inspiration Time | - |  | Double | VitalSigns | IntelliVue MX800 | InsTi
4890 | Overdistension Index | - |  | Double | VitalSigns | IntelliVue MX800 | C20/C
4891 | Lung Time Constant | - |  | Double | VitalSigns | IntelliVue MX800 | TC
4892 | Airway Correlation Coefficient | - |  | Double | VitalSigns | IntelliVue MX800 | r
4893 | Rate Volume Ratio | - |  | Double | VitalSigns | IntelliVue MX800 | RVrat
4894 | Normalized Ionized Calcium (Ca) | - |  | Double | VitalSigns | IntelliVue MX800 | iCa(N)
4956 | Peak Inspiratory Flow Setting | - |  | Double | Settings | IntelliVue MX800 | s%PIF
4957 | Trigger Sensitivity Setting | - |  | Double | Settings | IntelliVue MX800 | Available only in revF or earlier
4958 | Plateau Time Setting | - |  | Double | Settings | IntelliVue MX800 | sPltTi
4959 | Multiple Sigh Number Setting | - |  | Double | Settings | IntelliVue MX800 | VENT: sSghNr
4960 | Apnea Tidal Volume Setting | - |  | Double | Settings | IntelliVue MX800 | sATV
4961 | Apnea Respiration Rate Setting | - |  | Double | Settings | IntelliVue MX800 | sARR
4962 | Apnea Peak Flow Setting | - |  | Double | Settings | IntelliVue MX800 | sAPkFl
4963 | Apnea Inspired O2 Setting | - |  | Double | Settings | IntelliVue MX800 | VENT: sAFIO2
4964 | PSV Pressure Setting | - | Pressure Support Ventilation Setting | Double | Settings | IntelliVue MX800 | VENT: sPSV
4966 | Sigh Activation Status | - |  | Double | DeviceInfo | IntelliVue MX800 | VENT: sEnSgh
4967 | Suction O2 Concentration Setting | - |  | Double | Settings | IntelliVue MX800 | VENT: sO2Suc
4968 | PCV Control Pressure Setting | - |  | Double | Settings | IntelliVue MX800 | VENT: sPin
4969 | PCV Inspiratory Time Setting | - |  | Double | Settings | IntelliVue MX800 | VENT: sPVinT
4970 | PCV Apnea Respiration Rate Setting | - |  | Double | Settings | IntelliVue MX800 | VENT: sAPVRR
4971 | PCV Apnea Inspiratory Time Setting | - |  | Double | Settings | IntelliVue MX800 | VENT: sAPVTi
4972 | PCV Apnea O2 Setting | - |  | Double | Settings | IntelliVue MX800 | VENT: sAPVO2
4973 | PCV Apnea High Airway Pressure Setting | - |  | Double | Settings | IntelliVue MX800 | VENT: sAPVhP
4975 | Inspiratory Component of I:E Ratio Setting | - | Pressure Ventilation I component of I:E Ratio setting | Double | Settings | IntelliVue MX800 | sPVI
4976 | Expiratory Component of I:E Ratio Setting | - | Pressure Ventilation E component of I:E Ratio setting | Double | Settings | IntelliVue MX800 | sPVE
4977 | Inspiratory Component of Apnea Ventilation I:E Ratio Setting | - |  | Double | Settings | IntelliVue MX800 | sAPVI
4978 | Expiratory Component of Apnea Ventilation I:E Ratio Setting | - |  | Double | Settings | IntelliVue MX800 | sAPVE
4979 | Setting: Pressure Ventilation Control | - |  | Double | Settings | IntelliVue MX800 | Available only in revF or earlier
4981 | Cycle Time Setting | - |  | Double | Settings | IntelliVue MX800 | sCycTi
4982 | IPPV Ventilation Respiration Rate Setting | - |  | Double | Settings | IntelliVue MX800 | VENT: sIPPV
4983 | IMV Ventilation Respiration Rate Setting | - |  | Double | Settings | IntelliVue MX800 | VENT: sIMV
4984 | Pressure Support PEEP Setting | - |  | Double | Settings | IntelliVue MX800 | VENT: sSPEEP
4987 | Apnea Ventilation Delay | - |  | Double | Settings | IntelliVue MX800 | Available only in revF or earlier
4989 | O2 Monitoring Mode | - |  | Double | DeviceInfo | IntelliVue MX800 | VENT: sO2Mon
4990 | O2 Calibration Mode | - |  | Double | DeviceInfo | IntelliVue MX800 | VENT: sO2Cal
4993 | Inspiratory Time Setting | - |  | Double | Settings | IntelliVue MX800 | VENT: sInsTi
4994 | Expiratory Time Setting | - |  | Double | Settings | IntelliVue MX800 | VENT: sExpTi
4996 | Alarm Percentage on Rise Time Setting | - |  | Double | Settings | IntelliVue MX800 | VENT: sALMRT
4997 | Continuous Positive Airway Pressure (CPAP) Setting | - | This is the device's measurement with label="NLS_NOM_SETT_CONC_AWAY_O2" and physio_id="NOM_CONC_AWAY_O2". | Double | Settings | IntelliVue MX800 | sCPAP
4998 | Ventilator Flow Setting | - |  | Double | Settings | IntelliVue MX800 | VENT: sFlow
5000 | Low Inspiratory Pressure Setting | - |  | Double | Settings | IntelliVue MX800 | sPmin
5001 | HFV Flow Setting | - |  | Double | Settings | IntelliVue MX800 | sHFVFl
5002 | HFV Respiration Rate Setting | - |  | Double | Settings | IntelliVue MX800 | sHFVRR
5005 | Mechanical Ventilation Activated Status | - | Controlled Mechanical Ventilation Setting | Double | DeviceInfo | IntelliVue MX800 | VENT: sCMV
5006 | SIMV Mode | - | Synchronized Intermittent Mandatory Ventilation Setting | Double | DeviceInfo | IntelliVue MX800 | VENT: sSIMV
5007 | Mandatory Inspired Minute Volume Setting | - |  | Double | Settings | IntelliVue MX800 | VENT: sMMV
5009 | Infusion Pump Delivery Rate Setting | - |  | Double | Settings | IntelliVue MX800 | sDRate
5011 | Airway Respiration Rate Setting | - |  | Double | Settings | IntelliVue MX800 | sAWRR
5012 | Trigger Flow/Volume Setting | - | This is the device's measurement with label="NLS_NOM_SETT_VENT_VOL_LUNG_TRAPD" and physio_id="NOM_SETT_VENT_VOL_LUNG_TRAPD". | Double | Settings | IntelliVue MX800 | VENT: sTrVol
5013 | Apnea Alarm Delay | - |  | Double | DeviceInfo | IntelliVue MX800 | sAADel
5014 | HFV Amplitude Setting | - |  | Double | Settings | IntelliVue MX800 | sHFVAm
5015 | Volume Alarm Delay | - |  | Double | Settings | IntelliVue MX800 | sMVDel
5016 | Pressure Increase Rate Setting | - |  | Double | Settings | IntelliVue MX800 | VENT: sPincR
5017 | Volume Warning Setting | - | Delivered by the Evita 2 VueLink Driver. | Double | Settings | IntelliVue MX800 | VENT: sVmax
5026 | Airway Inspiratory Pressure Lower Alarm Limit | - |  | Double | Settings | IntelliVue MX800 | AlarmLimit: loPmax
5028 | Apnea Alarm Delay Setting | - |  | Double | Settings | IntelliVue MX800 | Available only in revF or earlier
5029 | Pressure Ramp Setting (ASB Ramp) | - |  | Double | Settings | IntelliVue MX800 | Available only in revF or earlier
5030 | Applied Tidal Volume Setting | - |  | Double | Settings | IntelliVue MX800 | sTVap
5031 | Assist Pressure Sensitivity Setting | - | Assist Sensitivity. Used by the Bear 1000 ventilator. | Double | Settings | IntelliVue MX800 | sSens
5032 | High Inspiratory Pressure Setting | - |  | Double | Settings | IntelliVue MX800 | Available only in revF or earlier
5033 | Background Flow Setting | - |  | Double | Settings | IntelliVue MX800 | VENT: sBkgFl
5035 | Vaporizer Agent Concentration Setting | - |  | Double | Settings | IntelliVue MX800 | sAGT
5036 | Vaporizer Isoflurane Concentration Setting | - |  | Double | Settings | IntelliVue MX800 | sISO
5037 | Vaporizer Enflurane Concentration Setting | - |  | Double | Settings | IntelliVue MX800 | sENF
5038 | Vaporizer Halothane Concentration Setting | - |  | Double | Settings | IntelliVue MX800 | sHAL
5039 | Vaporizer Desflurane Concentration Setting | - |  | Double | Settings | IntelliVue MX800 | sDES
5040 | Vaporizer Sevoflurane Concentration Setting | - |  | Double | Settings | IntelliVue MX800 | sSEV
5041 | Gas Measurement Sensor Position | - |  | Double | DeviceInfo | IntelliVue MX800 | VENT: sGasPr
5042 | O2 Measurement Sensor Position | - |  | Double | DeviceInfo | IntelliVue MX800 | VENT: sO2Pr
5044 | Inspired Tidal Volume Setting | - |  | Double | Settings | IntelliVue MX800 | sTVin
5045 | Incubator/Warmer Temperature Setting | - |  | Double | Settings | IntelliVue MX800
5046 | Preset Period of UrVol Setting | - |  | Double | Settings | IntelliVue MX800 | sUrTi
5050 | Low Time Setting | - | Setting variable for Evita 4 Airway Pressure Release Ventilation Mode. | Double | Settings | IntelliVue MX800 | VENT: sTlow
5051 | High Time Setting | - | Setting variable for Evita 4 Airway Pressure Release Ventilation Mode. | Double | Settings | IntelliVue MX800 | VENT: sThigh
5052 | Low Pressure Setting | - | Setting variable for Evita 4 Airway Pressure Release Ventilation Mode. | Double | Settings | IntelliVue MX800 | VENT: sPlow
5053 | High Pressure Setting | - | Setting variable for Evita 4 Airway Pressure Release Ventilation Mode. | Double | Settings | IntelliVue MX800 | VENT: sPhigh
5054 | Volume Assist Level Setting | - |  | Double | Settings | IntelliVue MX800 | VENT: sVolas
5055 | CPAP Flow Assist Level Setting | - |  | Double | Settings | IntelliVue MX800 | VENT: sFlas
5056 | Preset Stimulation Current Setting | - |  | Double | Settings | IntelliVue MX800 | sCurnt
5057 | Preset Stimulation Charge Setting | - |  | Double | Settings | IntelliVue MX800 | sChrge
5058 | Preset Stimulation Impulse Duration Setting | - |  | Double | Settings | IntelliVue MX800 | sPulsD
5059 | NMT TOF Repetition Time | - | Preset Train of Four (Slow TOF) Repetition Time Setting | Double | DeviceInfo | IntelliVue MX800 | NMT: sRepTi
5060 | Panting Limit Setting | - |  | Double | Settings | IntelliVue MX800 | VENT: sfmax
5065 | Airway Inspiratory Pressure Upper Alarm Limit | - | Alarm Limit: High Pressure | Double | Settings | IntelliVue MX800 | AlarmLimit: highP
5066 | Sigh Volume | - |  | Double | VitalSigns | IntelliVue MX800 | Vsigh
5069 | Sustained Pressure Alarm Limit | - |  | Double | Settings | IntelliVue MX800 | AlarmLimit: sustP
5072 | Intrathoracic Blood Volume (ITBV) | - |  | Double | VitalSigns | IntelliVue MX800 | ITBV
5073 | Intrathoracic Blood Volume Index (ITBI) | - |  | Double | VitalSigns | IntelliVue MX800 | ITBVI
5074 | Extravascular Lung Water | - |  | Double | VitalSigns | IntelliVue MX800 | EVLW
5075 | Extravascular Lung Water Index | - |  | Double | VitalSigns | IntelliVue MX800 | EVLWI
5076 | Global End Diastolic Volume | - |  | Double | VitalSigns | IntelliVue MX800 | GEDV
5077 | Global End Diastolic Volume Index | - |  | Double | VitalSigns | IntelliVue MX800 | GEDVI
5078 | Cardiac Function Index | - |  | Double | VitalSigns | IntelliVue MX800 | CFI
5079 | Mixed Venous O2 Saturation | - |  | Double | VitalSigns | IntelliVue MX800 | SvO2
5080 | Generic Term for the Transcutaneous Gases | - |  | Double | VitalSigns | IntelliVue MX800 | TCUT: tcGas
5083 | Intrinsic PEEP | - |  | Double | VitalSigns | IntelliVue MX800 | iPEEP
5086 | Positive Peak Inspiratory Pressure | - |  | Double | VitalSigns | IntelliVue MX800 | PIP
5087 | Dead Space to Tidal Volume Ratio | - |  | Double | VitalSigns | IntelliVue MX800 | Vd/Vt
5088 | Lung Volume Trapped | - |  | Double | VitalSigns | IntelliVue MX800 | VENT: TrpVol
5089 | Lung Vital Capacity | - |  | Double | VitalSigns | IntelliVue MX800 | VC
5090 | Generic Lung Compliance | - |  | Double | VitalSigns | IntelliVue MX800 | COMP
5092 | EEG Electromyography | - |  | Double | VitalSigns | IntelliVue MX800 | EEG: EMG
5093 | Suppression Ratio | - |  | Double | VitalSigns | IntelliVue MX800 | EEG: SR
5095 | EEG Peak Power Frequency | - |  | Double | VitalSigns | IntelliVue MX800 | EEG: PPF
5101 | End Tidal N2 | - |  | Double | VitalSigns | IntelliVue MX800 | N2
5102 | Inspired N2 | - |  | Double | VitalSigns | IntelliVue MX800 | N2
5109 | Generic Agent Wave | - |  | Double | Waveforms | IntelliVue MX800 | Waves: AGT
5110 | End Tidal MAC Concentration | - |  | Double | VitalSigns | IntelliVue MX800 | MAC
5111 | Left Ventricular Stroke Volume Index | - |  | Double | VitalSigns | IntelliVue MX800 | LVSWI
5112 | Left Ventricular Stroke Volume | - |  | Double | VitalSigns | IntelliVue MX800 | LVSW
5113 | Right Ventricular Stroke Volume | - |  | Double | VitalSigns | IntelliVue MX800 | RVSW
5114 | Glasgow Coma Score | - |  | Double | VitalSigns | IntelliVue MX800 | GCS
5115 | SubScore of the GCS - Eye Response | - |  | Double | VitalSigns | IntelliVue MX800 | EyeRsp
5116 | SubScore of the GCS - Motoric Response | - |  | Double | VitalSigns | IntelliVue MX800 | MotRsp
5117 | SubScore of the GCS - Verbal Response | - |  | Double | VitalSigns | IntelliVue MX800 | VblRsp
5118 | Head Circumference | - |  | Double | VitalSigns | IntelliVue MX800 | HC
5119 | pH in Arterial Blood | - |  | Double | VitalSigns | IntelliVue MX800 | Ph
5120 | Arterial Blood Hemoglobin (HB) | - |  | Double | VitalSigns | IntelliVue MX800 | Hb
5121 | pH in Venous Blood | - |  | Double | VitalSigns | IntelliVue MX800 | pHv
5122 | Urine Natrium | - |  | Double | VitalSigns | IntelliVue MX800 | UrNa
5123 | Natrium in Serum | - |  | Double | VitalSigns | IntelliVue MX800 | SerNa
5124 | pH in the Blood Plasma | - |  | Double | VitalSigns | IntelliVue MX800 | pH
5125 | Blood Natrium (Sodium) | - |  | Double | VitalSigns | IntelliVue MX800 | Na
5126 | Kalium (Potassium) | - |  | Double | VitalSigns | IntelliVue MX800 | K
5127 | Generic Glucose Parameter | - |  | Double | VitalSigns | IntelliVue MX800 | Glu
5128 | Partial CO2 | - |  | Double | VitalSigns | IntelliVue MX800 | PCO2
5129 | Partial O2 | - |  | Double | VitalSigns | IntelliVue MX800 | PO2
5130 | Base Excess of Blood | - |  | Double | VitalSigns | IntelliVue MX800 | BE
5131 | Percent Alveolarvenous Shunt Qs/Qt | - |  | Double | VitalSigns | IntelliVue MX800 | Qs/Qt
5132 | Extravascular Thermal Volume Index | - |  | Double | VitalSigns | IntelliVue MX800 | ETVI
5133 | VPB Rate | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
5134 | ECG Lead V1 Wave | - |  | Double | Waveforms | IntelliVue MX800 | Waves: V1
5135 | ECG Lead V2 Wave | - |  | Double | Waveforms | IntelliVue MX800 | Waves: V2
5136 | ECG Lead V3 Wave | - |  | Double | Waveforms | IntelliVue MX800 | Waves: V3
5137 | ECG Lead V4 Wave | - |  | Double | Waveforms | IntelliVue MX800 | Waves: V4
5138 | ECG Lead V5 Wave | - |  | Double | Waveforms | IntelliVue MX800 | Waves: V5
5139 | ECG Lead V6 Wave | - |  | Double | Waveforms | IntelliVue MX800 | Waves: V6
5140 | ECG Lead I Wave | - |  | Double | Waveforms | IntelliVue MX800 | Waves: I
5141 | ECG Lead II Wave | - |  | Double | Waveforms | IntelliVue MX800 | Waves: II
5142 | ECG Lead III Wave | - |  | Double | Waveforms | IntelliVue MX800 | Waves: III
5143 | ECG Lead aVR Wave | - |  | Double | Waveforms | IntelliVue MX800 | Waves: aVR
5144 | ECG Lead aVL Wave | - |  | Double | Waveforms | IntelliVue MX800 | Waves: aVL
5145 | ECG Lead aVF Wave | - |  | Double | Waveforms | IntelliVue MX800 | Waves: aVF
5185 | Mean Dominant Frequency (CH3) | - |  | Double | VitalSigns | IntelliVue MX800 | MDF3
5186 | Mean Dominant Frequency (CH4) | - |  | Double | VitalSigns | IntelliVue MX800 | MDF4
5189 | NMT DBS Count | - |  | Double | VitalSigns | IntelliVue MX800 | DBScnt
5191 | Peak Power Frequency (CH3) | - |  | Double | VitalSigns | IntelliVue MX800 | PPF3
5192 | Peak Power Frequency (CH4) | - |  | Double | VitalSigns | IntelliVue MX800 | PPF4
5497 | Pump Infusion Mode | - |  | Double | DeviceInfo | IntelliVue MX800 | sPuMod
5507 | Remaining Time Counter | - | Time to Elapse Counter used by the i-Stat module. | Double | OtherMeasuredData | IntelliVue MX800 | Elapse
5598 | APRV Low Pressure Synchronisation Time Setting | - |  | Double | Settings | IntelliVue MX800 | sTloS
5603 | APRV High Pressure Synchronisation Time Setting | - |  | Double | Settings | IntelliVue MX800 | sThiS
5614 | Acoustic Respiration Rate | - |  | Double | VitalSigns | IntelliVue MX800 | RRac
5616 | Regional O2 Saturation (rSO2 2) | - |  | Double | VitalSigns | Common
5617 | Regional O2 Saturation (rSO2 3) | - |  | Double | VitalSigns | Common
5620 | Regional O2 Saturation (rSO2 4) | - |  | Double | VitalSigns | Common
5632 | Regional O2 Saturation (rSO2 5) | - |  | Double | VitalSigns | IntelliVue MX800 | SO2 5
5633 | Regional O2 Saturation (rSO2 6) | - |  | Double | VitalSigns | IntelliVue MX800 | SO2 6
5637 | O2 Delivery Index (Thermodilution) | - |  | Double | VitalSigns | IntelliVue MX800 | iDO2I
5659 | Noninvasive Venous Hemoglobin (SpHbv) | - |  | Double | VitalSigns | IntelliVue MX800 | SpHbv
5660 | Noninvasive Arterial Hemoglobin (SpHba) | - |  | Double | VitalSigns | IntelliVue MX800 | SpHba
5661 | Generic Pulse Rate | - |  | Double | VitalSigns | IntelliVue MX800 | iAvgPR
5680 | Corrected Total Hemoglobin | - |  | Double | VitalSigns | IntelliVue MX800 | cSpHb
5681 | Corrected Arterial Total Hemoglobin | - |  | Double | VitalSigns | IntelliVue MX800 | cSpHb
5682 | Corrected Venous Total Hemoglobin | - |  | Double | VitalSigns | IntelliVue MX800 | cSpHbv
5694 | Systemic Vascular Resistance (Thermodilution) | - |  | Double | VitalSigns | IntelliVue MX800 | iSVR
5695 | Systemic Vascular Resistance Index (Thermodilution) | - |  | Double | VitalSigns | IntelliVue MX800 | iSVRI
5707 | Pulse Rate (Abdominal) | - |  | Double | VitalSigns | IntelliVue MX800 | Pulse
5719 | ECG Lead V3R | - |  | Double | VitalSigns | IntelliVue MX800 | ECG: ST
5720 | ECG Lead V4R | - |  | Double | VitalSigns | IntelliVue MX800 | ECG: ST
5721 | ECG Lead V5R | - |  | Double | VitalSigns | IntelliVue MX800 | ECG: ST
5763 | Inspired NO | - |  | Double | VitalSigns | IntelliVue MX800 | NO
5767 | ECG Lead V9 | - |  | Double | VitalSigns | IntelliVue MX800 | ECG: ST
5774 | ECG Lead V7 | - |  | Double | VitalSigns | IntelliVue MX800 | ECG: ST
5775 | ECG Lead V8 | - |  | Double | VitalSigns | IntelliVue MX800 | ECG: ST
5807 | Apnea Ventilation I:E Ratio Setting | unitless | This variable is associated with ChannelID identifier. This variable is calculated from variable 4977 and 4978, which is sent as a string in the format "1:xx.xx". | String | Settings | IntelliVue MX800 | Calculated Variables
5817 | Assisted Respiration Rate | - |  | Double | VitalSigns | IntelliVue MX800 | Rrasst
5818 | Correlation Coefficient Squared | - |  | Double | VitalSigns | IntelliVue MX800 | r2
5837 | ECG Lead V6R | - |  | Double | VitalSigns | IntelliVue MX800 | ECG: ST
5873 | Citrate Solution Concentration | - |  | Double | Settings | IntelliVue MX800 | sCitCo
5874 | Calcium Solution Concentration | - |  | Double | Settings | IntelliVue MX800 | sCConc
5886 | Citrate Acid Concentration | - |  | Double | Settings | IntelliVue MX800 | sCSCAC
5887 | Citrate Load Setting | - |  | Double | Settings | IntelliVue MX800 | sCitLo
5949 | Backup Respiration Rate Setting | - |  | Double | Settings | IntelliVue MX800 | sRRbak
5958 | Inspired Expired Pressure Difference | - |  | Double | VitalSigns | IntelliVue MX800 | sPi-e
5959 | Backup Inspired Expired Pressure Difference Setting | - |  | Double | Settings | IntelliVue MX800 | sPi-eB
6124 | Femoral Arterial Pressure Mean | - |  | Double | VitalSigns | IntelliVue MX800 | IBP: FAP
6125 | Femoral Arterial Pressure Systolic | - |  | Double | VitalSigns | IntelliVue MX800 | IBP: FAP
6126 | Femoral Arterial Pressure Diastolic | - |  | Double | VitalSigns | IntelliVue MX800 | IBP: FAP
6281 | Tidal Volume Lower Alarm Limit | - |  | Double | Settings | IntelliVue MX800 | AlarmLimit: lowTV
6283 | Minute Volume Lower Alarm Limit | - |  | Double | Settings | IntelliVue MX800 | AlarmLimit: lowMV
6322 | ECG Lead V7 Wave | - |  | Double | Waveforms | IntelliVue MX800 | Waves: V7
6323 | ECG Lead V8 Wave | - |  | Double | Waveforms | IntelliVue MX800 | Waves: V8
6324 | ECG Lead V9 Wave | - |  | Double | Waveforms | IntelliVue MX800 | Waves: V9
6326 | Airway Aux Wave | - |  | Double | Waveforms | IntelliVue MX800 | Waves: AWPx
6331 | Fetal DECG Wave | - |  | Double | Waveforms | IntelliVue MX800 | Waves: DECG
6339 | Patient BSA Formula | unitless | Consult detailed help for more information on this variable | Double | PatientInfo | Common | Patient Demographics
6343 | Platelets (Thrombocyte Count) | - |  | Double | VitalSigns | IntelliVue MX800 | Plts
6388 | Endotracheal Tube Type | - |  | Double | DeviceInfo | IntelliVue MX800 | sTube
6390 | Minute Volume Upper Alarm Limit | - |  | Double | Settings | IntelliVue MX800 | AlarmLimit: highMV
6394 | Pulse Rate from Plethysmogram | - |  | Double | VitalSigns | Common | Pulse
6396 | ECG Ectopic Status | - | Consult detailed help for more information on this variable | Double | VitalSigns | Common | Ectopic Status (on IntelliVue)
6399 | ST Lead ES | - |  | Double | VitalSigns | IntelliVue MX800 | ECG: ST
6400 | ST Lead AS | - |  | Double | VitalSigns | IntelliVue MX800 | ECG: ST
6401 | ST Lead AI | - |  | Double | VitalSigns | IntelliVue MX800 | ECG: ST
6402 | ST Index | - |  | Double | VitalSigns | IntelliVue MX800 | ECG: Stindx
6409 | Pulse Pressure | - |  | Double | VitalSigns | IntelliVue MX800 | PP
6410 | Mean Transit Time (MTt) | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
6450 | Alveolar Dead Space | - |  | Double | VitalSigns | IntelliVue MX800 | Vdalv
6451 | Tympanic Temperature | - |  | Double | VitalSigns | IntelliVue MX800 | Ttymp
6452 | Injectate Volume (CO) | - |  | Double | VitalSigns | IntelliVue MX800 | InjVol
6453 | SpO2 (Telemetry) | - |  | Double | VitalSigns | IntelliVue MX800 | %SpO2T
6454 | Pulse Rate (Telemetry) | - |  | Double | VitalSigns | IntelliVue MX800 | PulseT
6455 | Intracranial Pressure ICP1 (Systolic) | - |  | Double | VitalSigns | IntelliVue MX800 | IBP: IC2
6456 | Intracranial Pressure ICP1 (Diastolic) | - |  | Double | VitalSigns | IntelliVue MX800 | IBP: IC1
6457 | Intracranial Pressure ICP1 (Mean) | - |  | Double | VitalSigns | IntelliVue MX800 | IBP: IC3
6458 | Intracranial Pressure ICP2 Systolic | - |  | Double | VitalSigns | IntelliVue MX800 | IBP: IC2
6459 | Intracranial Pressure ICP2 Diastolic | - |  | Double | VitalSigns | IntelliVue MX800 | IBP: IC2
6460 | Intracranial Pressure ICP2 Mean | - |  | Double | VitalSigns | IntelliVue MX800 | IBP: IC2
6461 | Brachial Arterial Pressure (BA) Systolic | - |  | Double | VitalSigns | IntelliVue MX800 | IBP: BAP
6462 | Brachial Arterial Pressure (BA) Diastolic | - |  | Double | VitalSigns | IntelliVue MX800 | IBP: BAP
6463 | Brachial Arterial Pressure (BA) Mean | - |  | Double | VitalSigns | IntelliVue MX800 | IBP: BAP
6464 | Vesical Temperature (Bladder) | - |  | Double | VitalSigns | IntelliVue MX800 | Tvesic
6465 | Cerebral Temperature | - |  | Double | VitalSigns | IntelliVue MX800 | Tcereb
6466 | ST Level for Lead I | - |  | Double | VitalSigns | IntelliVue MX800 | ECG: Rf-I
6467 | ST Level for Lead II | - |  | Double | VitalSigns | IntelliVue MX800 | ECG: Rf-II
6468 | ST Level for Lead III | - |  | Double | VitalSigns | IntelliVue MX800 | ECG: Rf-III
6469 | ST Level for Lead aVR | - |  | Double | VitalSigns | IntelliVue MX800 | ECG: Rf-aVR
6470 | ST Level for Lead aVL | - |  | Double | VitalSigns | IntelliVue MX800 | ECG: Rf-aVL
6471 | ST Level for Lead aVF | - |  | Double | VitalSigns | IntelliVue MX800 | ECG: Rf-aVF
6472 | ST Level for Lead V1 | - |  | Double | VitalSigns | IntelliVue MX800 | ECG: Rf-V1
6473 | ST Level for Lead V2 | - |  | Double | VitalSigns | IntelliVue MX800 | ECG: Rf-V2
6474 | ST Level for Lead V3 | - |  | Double | VitalSigns | IntelliVue MX800 | ECG: Rf-V3
6475 | ST Level for Lead V4 | - |  | Double | VitalSigns | IntelliVue MX800 | ECG: Rf-V4
6476 | ST Level for Lead V5 | - |  | Double | VitalSigns | IntelliVue MX800 | ECG: Rf-V5
6477 | ST Level for Lead V6 | - |  | Double | VitalSigns | IntelliVue MX800 | ECG: Rf-V6
6478 | Prothrombin Time (International Normalized Ratio) | - |  | Double | VitalSigns | IntelliVue MX800 | PT INR
6479 | PSV Tidal Volume | - | Tidal Volume in Pressure Support Ventilation Mode | Double | VitalSigns | IntelliVue MX800 | TVPSV
6480 | Negative Inspiratory Force Index | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
6481 | A-Line ARX Index | - |  | Double | VitalSigns | IntelliVue MX800 | EEG: AAI
6484 | Impedance Cardiography (ICG) Wave | - |  | Double | Waveforms | IntelliVue MX800 | Waves: ICG
6485 | Thoracic Fluid Content | - |  | Double | VitalSigns | IntelliVue MX800 | TFC
6486 | Thoracic Fluid Content Index | - |  | Double | VitalSigns | IntelliVue MX800 | TFI
6487 | Accelerated Cardiac Index | - |  | Double | VitalSigns | IntelliVue MX800 | ACI
6488 | Heart Contractility Index | - |  | Double | VitalSigns | IntelliVue MX800 | HI
6489 | O2 Saturation (Left) | - |  | Double | VitalSigns | IntelliVue MX800 | SO2 l
6490 | O2 Saturation (Right) | - |  | Double | VitalSigns | IntelliVue MX800 | SO2 r
6491 | Pulse Pressure Variation | - |  | Double | VitalSigns | IntelliVue MX800 | PPV
6492 | Downslope Time | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
6493 | Pulmonary Vascular Permeability Index | - | This is the device's measurement with label="NLS_NOM_PERM_VASC_PULM_INDEX" and physio_id="NOM_PERM_VASC_PULM_INDEX". See also variable 8791. | Double | VitalSigns | IntelliVue MX800 | PVPI
6502 | ART Pulse Rate | - | ART pulse rate | Double | VitalSigns | IntelliVue MX800 | Pulse
6543 | Umbilical Arterial Pressure Pulse Rate | - |  | Double | VitalSigns | IntelliVue MX800 | Pulse
6544 | Patient Visit Number | unitless | Patient encounter ID contains the ID of the current visit of the patient | String | PatientInfo | IntelliVue MX800 | Patient Demographics
6555 | Patient Height at Entered Time | - |  | Double | PatientInfo | Common | Patient Demographics
6557 | Patient Weight at Entered Time | - |  | Double | PatientInfo | Common | Patient Demographics
6567 | Patient Body Surface Area (BSA) | - |  | Double | PatientInfo | Common | Patient Demographics
6685 | Inspired O2 Setting (FiO2) | - |  | Double | Settings | IntelliVue MX800 | VENT: sFIO2
6702 | Inspiratory Component of I:E Ratio Setting | - |  | Double | Settings | IntelliVue MX800 | sI
6703 | Expiratory Component of I:E Ratio Setting | - |  | Double | Settings | IntelliVue MX800 | sE
6735 | Lung Elastance | - |  | Double | VitalSigns | IntelliVue MX800 | E
6739 | Backup Inspiratory Time Setting | - |  | Double | Settings | IntelliVue MX800 | sInTiB
6744 | Pressure Support Level Above Phigh | - |  | Double | Settings | IntelliVue MX800 | sdPin
6746 | Total Respiration Duration Setting | - |  | Double | Settings | IntelliVue MX800 | sBreaT
6772 | Heater Power | - |  | Double | DeviceInfo | IntelliVue MX800 | sPower
6861 | Capillary Refill Time | - |  | Double | VitalSigns | IntelliVue MX800 | CRT
6935 | Estimated Patient Circuit Leakage | - |  | Double | OtherMeasuredData | IntelliVue MX800 | LkPat
6936 | Estimated Ventilation Total Leakage | - |  | Double | OtherMeasuredData | IntelliVue MX800 | LkTot
6967 | Ventilation Mode Additional Information | - |  | Double | DeviceInfo | IntelliVue MX800 | sVSupp
6995 | Backup Plateau Time Setting | - |  | Double | Settings | IntelliVue MX800 | sPltTA
7000 | ECG Lead V3R Wave | - |  | Double | Waveforms | IntelliVue MX800 | Waves: V3R
7001 | ECG Lead V4R Wave | - |  | Double | Waveforms | IntelliVue MX800 | Waves: V4R
7002 | ECG Lead V5R Wave | - |  | Double | Waveforms | IntelliVue MX800 | Waves: V5R
7003 | ECG Lead V6R Wave | - |  | Double | Waveforms | IntelliVue MX800 | Waves: V6R
7019 | Signal Quality Index (BIS) | - | This variable is associated with ChannelID identifier. | Double | OtherMeasuredData | IntelliVue MX800 | SQI L/SQI R
7106 | Continuous NIBP Systolic | - |  | Double | VitalSigns | IntelliVue MX800 | cNBP SYS
7107 | Continuous NIBP Diastolic | - |  | Double | VitalSigns | IntelliVue MX800 | cNBP DIA
7108 | Continuous NIBP Mean | - |  | Double | VitalSigns | IntelliVue MX800 | cNBP MEAN
7118 | Next Intervention Time | - |  | Double | OtherMeasuredData | IntelliVue MX800 | InTtvTi
7315 | Total Power (CH3) | - |  | Double | VitalSigns | IntelliVue MX800 | TP3
7316 | Total Power (CH4) | - |  | Double | VitalSigns | IntelliVue MX800 | TP4
7335 | Total Power (Left) | - |  | Double | VitalSigns | IntelliVue MX800 | TP L
7336 | Total Power (Right) | - |  | Double | VitalSigns | IntelliVue MX800 | TP R
7414 | PCV I:E Ratio Setting | unitless | This variable is associated with ChannelID identifier. This variable is calculated from variable 4975 and 4976, which is sent as a string in the format "1:xx.xx". | String | Settings | IntelliVue MX800 | Calculated Variables
7434 | Inspiratory CO2 (FiCO2) | - |  | Double | VitalSigns | IntelliVue MX800 | VENT: FICO2
7532 | End Tidal NO2 | - |  | Double | VitalSigns | IntelliVue MX800 | NO2
7535 | Fresh Gas Total Flow | - |  | Double | OtherMeasuredData | IntelliVue MX800 | fgFlow
7545 | Bolus Volume Setting | - |  | Double | Settings | IntelliVue MX800 | sVBol
7557 | Generic NIBP Pressure | - | Available on IntelliVue devices with firmware revision G or later. | Double | VitalSigns | Common | NBP
7587 | Endotracheal Tube Diameter | - |  | Double | DeviceInfo | IntelliVue MX800 | sTubeD
7640 | Mandatory Respiration Rate | - |  | Double | VitalSigns | IntelliVue MX800 | VENT: MRR
7641 | Mandatory Minute Volume | - |  | Double | VitalSigns | IntelliVue MX800 | VENT: MMV
7642 | Mandatory Tidal Volume | - |  | Double | VitalSigns | IntelliVue MX800 | VENT: MTV
7644 | N2O Consumption | - |  | Double | OtherMeasuredData | IntelliVue MX800 | AccN2O
7765 | Device Power Up Time | s | Number of seconds since the Monitor device's last power-on. | Double | DeviceInfo | Common | System Information
7768 | O2 Consumption | - |  | Double | VitalSigns | IntelliVue MX800 | AccO2
7770 | Air Consumption | - |  | Double | VitalSigns | IntelliVue MX800 | AccAir
7773 | Arterial Blood Pressure (ABP) | - | Available on IntelliVue devices with firmware revision G or later. Arterial Blood Pressure (NOM_PRESS_BLD_ART_ABP). | Double | VitalSigns | Common | IBP: ABP
7777 | ABP Pulse Rate | - | ABP pulse rate | Double | VitalSigns | IntelliVue MX800 | Pulse
7797 | Signal Noise Ratio (CCO) | - |  | Double | VitalSigns | IntelliVue MX800 | SNR
7803 | Spontaneous Airway Respiration Rate | - |  | Double | VitalSigns | IntelliVue MX800 | SpAWRR
7804 | Synchronized Respiration Rate | - |  | Double | VitalSigns | IntelliVue MX800 | RRsync
7805 | Mechanical Respiration Rate | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
7806 | Daily Urine Output | - |  | Double | VitalSigns | IntelliVue MX800 | U/O
7809 | Activated Partial Thromboplastin Time Ratio | - |  | Double | VitalSigns | IntelliVue MX800 | PTTrat
7810 | Global Ejection Fraction | - | This is the device's measurement with label="NLS_NOM_FRACT_EJECT" and physio_id="NOM_FRACT_EJECT". See also 8790. | Double | VitalSigns | IntelliVue MX800 | GEF
7811 | Generic CO2 Parameter | - | This variable contains "PECO2, Partial expired CO2", which is measured from label="NLS_NOM_VENT_AWAY_CO2_EXP", physio_id="NOM_VENT_AWAY_CO2_EXP". | Double | VitalSigns | IntelliVue MX800 | VENT: PECO2
7823 | Central Venous O2 Saturation (ScvO2) | - |  | Double | VitalSigns | IntelliVue MX800 | ScvO2
7858 | Suppression Ratio (BIS) | - | This variable is associated with ChannelID identifier. | Double | VitalSigns | IntelliVue MX800 | SR L/SR R
7865 | Right Ventricular Stroke Work Index | - |  | Double | VitalSigns | IntelliVue MX800 | RVSWI
7870 | Base Excess of Extracellular Fluid | - |  | Double | VitalSigns | IntelliVue MX800 | BEecf
7878 | Rise Time Setting | - |  | Double | Settings | IntelliVue MX800 | sRmpTi
7879 | Time Duration at Low Pressure (Tlow) | - |  | Double | VitalSigns | IntelliVue MX800 | Th/TL
7880 | Time Duration at High Pressure (Thigh) | - |  | Double | VitalSigns | IntelliVue MX800 | Th/TL
7904 | Spontaneous Tidal Volume | - |  | Double | VitalSigns | IntelliVue MX800 | VENT: SpTV
7905 | Cardiac Troponin l | - |  | Double | VitalSigns | IntelliVue MX800 | cTnI
7906 | Cardiopulmonary Bypass Status | - |  | Double | VitalSigns | IntelliVue MX800 | CPB
7907 | Cardiac Brain Natriuretic Peptide | - |  | Double | VitalSigns | IntelliVue MX800 | BNP
7908 | Mean Dominant Frequency | - |  | Double | VitalSigns | IntelliVue MX800 | EEG: MDF
7983 | Filtration Current Volume | - |  | Double | VitalSigns | IntelliVue MX800 | Vufr/h
7986 | Calcium Flow Rate Setting | - |  | Double | Settings | IntelliVue MX800 | sCSR
7987 | Calcium Compensation Setting | - |  | Double | Settings | IntelliVue MX800 | sCComp
8189 | Cerebral State Index (CSI) | - |  | Double | VitalSigns | IntelliVue MX800 | CSI
8201 | EEG Amplitude | - | This variable is associated with ChannelID identifier. | Double | VitalSigns | IntelliVue MX800 | aEEG1/aEEG2/aEEG3/aEEG4
8339 | Patient Middle Names | unitless | The Middle Name attribute contains the middle name of the patient. | String | PatientInfo | IntelliVue MX800 | Patient Demographics
8350 | Ao Pulse Rate | - |  | Double | VitalSigns | IntelliVue MX800 | Pulse
8351 | PAP Pulse Rate | - |  | Double | VitalSigns | IntelliVue MX800 | Pulse
8352 | CVP Pulse Rate | - |  | Double | VitalSigns | IntelliVue MX800 | Pulse
8353 | RAP Pulse Rate | - |  | Double | VitalSigns | IntelliVue MX800 | Pulse
8354 | LAP Pulse Rate | - |  | Double | VitalSigns | IntelliVue MX800 | Pulse
8355 | ICP Pulse Rate | - |  | Double | VitalSigns | IntelliVue MX800 | Pulse
8356 | Umbilical Venous Pressure Pulse Rate | - |  | Double | VitalSigns | IntelliVue MX800 | Pulse
8357 | FAP Pulse Rate | - |  | Double | VitalSigns | IntelliVue MX800 | Pulse
8358 | BAP Pulse Rate | - |  | Double | VitalSigns | IntelliVue MX800 | Pulse
8359 | IC1 Pulse Rate | - |  | Double | VitalSigns | IntelliVue MX800 | Pulse
8360 | Intracranial Pressure ICP2 Pulse Rate | - |  | Double | VitalSigns | IntelliVue MX800 | Pulse
8361 | Pulse Rate (P1) | - |  | Double | VitalSigns | IntelliVue MX800 | Pulse
8362 | Pulse Rate (P2) | - |  | Double | VitalSigns | IntelliVue MX800 | Pulse
8363 | Pulse Rate (P3) | - |  | Double | VitalSigns | IntelliVue MX800 | Pulse
8364 | Pulse Rate (P4) | - |  | Double | VitalSigns | IntelliVue MX800 | Pulse
8365 | Pulse Rate from Plethysmogram (Left) | - |  | Double | VitalSigns | IntelliVue MX800 | Pulse
8367 | Pleth Wave (Right) | - |  | Double | Waveforms | IntelliVue MX800 | Waves: PLETHr
8368 | Pleth Wave (Left) | - |  | Double | Waveforms | IntelliVue MX800 | Waves: PLETHl
8369 | Calculated SO2 | - |  | Double | VitalSigns | IntelliVue MX800 | CALC: 'SO2
8370 | Airway Respiration Rate (Ohmeda) | - |  | Double | VitalSigns | IntelliVue MX800 | VENT: RRaw
8371 | Airway Respiration Rate Setting (Ohmeda) | - |  | Double | Settings | IntelliVue MX800 | VENT: sRRaw
8372 | Fresh Gas O2 Concentration | - |  | Double | OtherMeasuredData | IntelliVue MX800 | fgO2
8373 | Manually entered FIO2 or Flow | - | This variable contains measurement from label "NLS_NOM_CONC_AWAY_O2" and observed value "NOM_CONC_AWAY_O2". | Double | VitalSigns | IntelliVue MX800 | O2
8374 | Generic O2 Parameter | - | This is the device's measurement with label="NLS_NOM_SETT_CONC_AWAY_O2" and physio_id="NOM_CONC_AWAY_O2". | Double |  | IntelliVue MX800 | sO2
8375 | Fresh Gas O2 Flow Setting | - | Mixer parameter. | Double | Settings | IntelliVue MX800 | sfgO2
8376 | Airway Pressure Numeric Value (Ohmeda) | - |  | Double | VitalSigns | IntelliVue MX800 | AWP
8377 | Inspiratory Peak Flow | - |  | Double | VitalSigns | IntelliVue MX800 | PIF
8378 | Maximum Pressure During Breathing Cycle | - |  | Double | VitalSigns | IntelliVue MX800 | VENT: Pmax
8379 | Peak Pressure | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
8380 | Peak Inspiratory Flow Setting | - |  | Double | Settings | IntelliVue MX800 | sPIF
8381 | Airway Maximum Pressure Setting | - |  | Double | Settings | IntelliVue MX800 | VENT: sPmax
8382 | Peak Inspiratory Pressure Setting | - |  | Double | Settings | IntelliVue MX800 | sPIP
8383 | Setting: Inspiratory Pressure (Ohmeda) | - |  | Double | Settings | IntelliVue MX800 | Available only in revF or earlier
8384 | Mean Airway Pressure (Ohmeda) | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
8385 | EEG Wave (Left) | - |  | Double | Waveforms | IntelliVue MX800 | Waves: LT EEG
8386 | EEG Wave (Right) | - |  | Double | Waveforms | IntelliVue MX800 | Waves: RT EEG
8387 | % of Power in the Delta Band - 1st EEG | - |  | Double | VitalSigns | IntelliVue MX800 | EEG: Prcnt1
8388 | % of Power in the Delta Band - 2nd EEG | - |  | Double | VitalSigns | IntelliVue MX800 | EEG: Prcnt2
8389 | Percent Delta - Left Side | - |  | Double | VitalSigns | IntelliVue MX800 | EEG: LT %DL
8390 | Percent Delta - Right Side | - |  | Double | VitalSigns | IntelliVue MX800 | EEG: RT %DL
8391 | % of Power in the Theta Band (1st EEG) | - |  | Double | VitalSigns | IntelliVue MX800 | EEG: Prcnt1
8392 | % of Power in the Theta Band (2nd EEG) | - |  | Double | VitalSigns | IntelliVue MX800 | EEG: Prcnt2
8393 | Percent Theta - Left Side | - |  | Double | VitalSigns | IntelliVue MX800 | EEG: LT %TH
8394 | Percent Theta - Right Side | - |  | Double | VitalSigns | IntelliVue MX800 | EEG: RT %TH
8395 | Generic Calculation Constant | - | This variable contains measurement from Philips label "NLS_NOM_NUM_CALC_CONST" and observed value "NOM_METRIC_NOS". See also variable 8493. | Double | VitalSigns | IntelliVue MX800 | CompCt
8396 | Pressure Ventilation Control Pressure | - |  | Double | VitalSigns | IntelliVue MX800 | VENT: PVcP
8397 | Right Heart Ejection Fraction | - |  | Double | VitalSigns | IntelliVue MX800 | REF
8398 | White Blood Count (Leucocyte Count) | - |  | Double | VitalSigns | IntelliVue MX800 | WBC
8399 | Red Blood Count (Erithrocyte Count) | - |  | Double | VitalSigns | IntelliVue MX800 | RBC
8400 | Mean Corpuscular Volume | - |  | Double | VitalSigns | IntelliVue MX800 | MCV
8401 | Corpuscular Mean Hemoglobin Content | - |  | Double | VitalSigns | IntelliVue MX800 | MCH
8402 | Ferrum | - |  | Double | VitalSigns | IntelliVue MX800 | Fe
8403 | Fetal Hemoglobin | - |  | Double | VitalSigns | IntelliVue MX800 | HbF
8404 | MetHemoglobin | - |  | Double | VitalSigns | IntelliVue MX800 | Met-Hb
8405 | Airway Volume Wave | - |  | Double | Waveforms | IntelliVue MX800 | Waves: AWV
8406 | Expiratory Peak Flow | - |  | Double | VitalSigns | IntelliVue MX800 | PEF
8407 | Expiratory Peak Flow | - |  | Double | VitalSigns | IntelliVue MX800 | eeFlow
8408 | The V stands for Volume | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
8409 | The R stands for Rate | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
8410 | Adjusted PaCO2 (at Patient Temperature) | - |  | Double | VitalSigns | IntelliVue MX800 | &PaCO2
8411 | Incubator Heating Power | - | Power Required to set Air&Pat Temp in the Incubator | Double | OtherMeasuredData | IntelliVue MX800 | INCUBATOR: Power
8412 | Actual Patient Age | - | Actual patient age, provided by a module, measured in years. Please note that this Patient Age will be available only when the corresponding module is connected to the Monitor. Variable 169 contains the patient age as provided by the Monitor itself (calculated from the birth date entered on the Monitor device.) | Double | PatientInfo | IntelliVue MX800 | Age
8413 | The P stands for Pressure | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
8414 | PCV Apnea Control Pressure Setting | - |  | Double | Settings | IntelliVue MX800 | VENT: sAPVcP
8415 | Fresh Gas Air Flow Setting | - | Mixer parameter. | Double | Settings | IntelliVue MX800 | sfgAir
8416 | PEEP/CPAP Lower Alarm Limit | - |  | Double | Settings | IntelliVue MX800 | AlarmLimit: loPEEP
8417 | Tidal Volume Upper Alarm Limit | - |  | Double | Settings | IntelliVue MX800 | AlarmLimit: highTV
8418 | T-Tre Temperature difference between Temp and Trect | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
8419 | T-Tbl Temperature difference between Temp and Tblood | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
8420 | T-Tco Temperature difference between Temp and Tcore | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
8421 | T-Tsk Temperature difference between Temp and Tskin | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
8422 | T-Tes Temperature difference between Temp and Tesoph | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
8423 | T-Tna Temperature difference between Temp and Tnaso | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
8424 | T-Tar Temperature difference between Temp and Tart | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
8425 | T-Tve Temperature difference between Temp and Tven | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
8426 | T-Taw Temperature difference between Temp and Tairwy | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
8427 | Tre-T Temperature difference between Trect and Temp | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
8428 | Tre-bl Temperature difference between Trect and Tblood | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
8429 | Tre-sk Temperature difference between Trect and Tskin | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
8430 | Tre-aw Temperature difference between Trect and Tairwy | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
8431 | Temperature difference between Tblood and Temp | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
8433 | Tbl-aw Temperature difference between Tblood and Tairwy | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
8434 | Tco-T Temperature difference between Tcore and Temp | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
8435 | Tco-re Temperature difference between Tcore and Trect | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
8436 | Tco-bl Temperature difference between Tcore and Tblood | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
8437 | Tco-sk Temperature difference between Tcore and Tskin | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
8438 | Tco-es Temperature difference between Tcore and Tesoph | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
8439 | Tco-na Temperature difference between Tcore and Tnaso | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
8440 | Tco-aw Temperature difference between Tcore and Tairwy | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
8441 | Tsk-T Temperature difference between Tskin and Temp | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
8442 | Tsk-bl Temperature difference between Tskin and Tblood | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
8443 | Tsk-aw Temperature difference between Tskin and Tairwy | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
8444 | Tes-T Temperature difference between Tesoph and Temp | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
8445 | Tes-re Temperature difference between Tesoph and Trect | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
8446 | Tes-bl Temperature difference between Tesoph and Tblood | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
8447 | Tes-sk Temperature difference between Tesoph and Tskin | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
8448 | Tes-na Temperature difference between Tesoph and Tnaso | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
8449 | Tes-aw Temperature difference between Tesoph and Tairwy | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
8450 | Tna-T Temperature difference between Tnaso and Temp | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
8451 | Tna-re Temperature difference between Tnaso and Trect | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
8452 | Tna-bl Temperature difference between Tnaso and Tblood | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
8453 | Tna-sk Temperature difference between Tnaso and Tskin | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
8454 | Tna-aw Temperature difference between Tnaso and Tairwy | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
8455 | Tar-T Temperature difference between Tart and Temp | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
8456 | Tar-re Temperature difference between Tart and Trect | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
8457 | Tar-bl Temperature difference between Tart and Tblood | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
8458 | Tar-co Temperature difference between Tart and Tcore | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
8459 | Tar-sk Temperature difference between Tart and Tskin | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
8460 | Tar-es Temperature difference between Tart and Tesoph | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
8461 | Tar-na Temperature difference between Tart and Tnaso | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
8462 | Tar-aw Temperature difference between Tart and Tairwy | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
8463 | Tve-T Temperature difference between Tven and Temp | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
8464 | Tve-re Temperature difference between Tven and Trect | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
8465 | Tve-bl Temperature difference between Tven and Tblood | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
8466 | Tve-co Temperature difference between Tven and Tcore | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
8467 | Tve-sk Temperature difference between Tven and Tskin | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
8468 | Tve-es Temperature difference between Tven and Tesoph | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
8469 | Tve-na Temperature difference between Tven and Tnaso | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
8470 | MAC Concentration (Awake) | - |  | Double | VitalSigns | IntelliVue MX800 | MAC
8471 | Tve-ar Temperature difference between Tven and Tart | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
8472 | Tve-aw Temperature difference between Tven and Tairwy | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
8473 | Taw-T Temperature difference between Tairwy and Temp | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
8474 | Taw-re Temperature difference between Tairwy and Trect | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
8475 | Taw-bl Temperature difference between Tairwy and Tblood | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
8476 | Taw-co Temperature difference between Tairwy and Tcore | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
8477 | Taw-sk Temperature difference between Tairwy and Tskin | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
8478 | Taw-es Temperature difference between Tairwy and Tesoph | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
8479 | Taw-na Temperature difference between Tairwy and Tnaso | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
8480 | Taw-ar Temperature difference between Tairwy and Tart | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
8481 | Taw-ve Temperature difference between Tairwy and Tven | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
8483 | Generic Agent 1 Wave | - |  | Double | Waveforms | IntelliVue MX800 | Waves: AGT1
8484 | End Tidal Generic Agent 1 | - |  | Double | VitalSigns | IntelliVue MX800 | AGT1
8485 | Inspired Generic Agent 1 | - |  | Double | VitalSigns | IntelliVue MX800 | AGT1
8486 | Generic Agent 2 Wave | - |  | Double | Waveforms | IntelliVue MX800 | Waves: AGT2
8487 | End Tidal Generic Agent 2 | - |  | Double | VitalSigns | IntelliVue MX800 | AGT2
8488 | Inspired Generic Agent 2 | - |  | Double | VitalSigns | IntelliVue MX800 | AGT2
8489 | Adjusted Arterial Blood pH Value | - |  | Double | VitalSigns | IntelliVue MX800 | &pHa
8490 | Adjusted Venous Blood pH Value | - |  | Double | VitalSigns | IntelliVue MX800 | &pHv
8491 | Adjusted PvCO2 (at Patient Temperature) | - |  | Double | VitalSigns | IntelliVue MX800 | &PvCO2
8492 | Adjusted PvO2 (at Patient Temperature) | - |  | Double | VitalSigns | IntelliVue MX800 | &PvO2
8493 | Generic Calculation Constant | - | This variable contains parameter from Philips label "NLS_NOM_NUM_CATHETER_CONST" and observed value "NOM_NUM_CATHETER_CONST". See also variable 8395. | Double | VitalSigns | IntelliVue MX800 | CathCt
8501 | ECG Lead V Wave | - |  | Double | Waveforms | IntelliVue MX800 | Waves: V
8502 | ECG Lead MCL Wave | - |  | Double | Waveforms | IntelliVue MX800 | Waves: MCL
8503 | ECG Lead MCL1 Wave | - |  | Double | Waveforms | IntelliVue MX800 | Waves: MCL1
8510 | MAC Concentration | - |  | Double | VitalSigns | IntelliVue MX800 | MAC
8511 | Pleth Wave (Telemetry) | - |  | Double | Waveforms | IntelliVue MX800 | Waves: PlethT
8513 | NIBP Systolic (Telemetry) | - |  | Double | VitalSigns | IntelliVue MX800 | NBP_T SYS
8514 | NIBP Diastolic (Telemetry) | - |  | Double | VitalSigns | IntelliVue MX800 | NBP_T DIA
8515 | NIBP Mean (Telemetry) | - |  | Double | VitalSigns | IntelliVue MX800 | NBP_T MEAN
8517 | Generic Pressure 1 (P1) Systolic | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
8518 | Generic Pressure 1 (P1) Diastolic | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
8519 | Generic Pressure 1 (P1) Mean | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
8521 | Generic Pressure 2 (P2) Systolic | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
8522 | Generic Pressure 2 (P2) Diastolic | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
8523 | Generic Pressure 2 (P2) Mean | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
8525 | Generic Pressure 3 (P3) Systolic | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
8526 | Generic Pressure 3 (P3) Diastolic | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
8527 | Generic Pressure 3 (P3) Mean | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
8529 | Generic Pressure 4 (P4) Systolic | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
8530 | Generic Pressure 4 (P4) Diastolic | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
8531 | Generic Pressure 4 (P4) Mean | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
8535 | Incubator Humidity | - |  | Double | OtherMeasuredData | IntelliVue MX800 | INCUBATOR: Hum
8536 | Predictive Oral Temperature | - |  | Double | VitalSigns | IntelliVue MX800 | pToral
8537 | Predictive Rectal Temperature | - |  | Double | VitalSigns | IntelliVue MX800 | pTrect
8538 | Predictive Axillary Temperature | - |  | Double | VitalSigns | IntelliVue MX800 | pTaxil
8539 | Perfusion (Telemetry) | - |  | Double | VitalSigns | IntelliVue MX800 | Perf T
8541 | Right to Left Heart Shunt | - |  | Double | VitalSigns | IntelliVue MX800 | RLShnt
8542 | Adjusted PcO2 (at Patient Temperature) | - |  | Double | VitalSigns | IntelliVue MX800 | &PCO2
8543 | Adjusted Arterial PaO2 (at Patient Temperature) | - |  | Double | VitalSigns | IntelliVue MX800 | &PaO2
8544 | Pulse Rate from Plethysmogram (Right) | - |  | Double | VitalSigns | IntelliVue MX800 | Pulse
8549 | Patient O2 Consumption per Body Surface | - |  | Double | VitalSigns | IntelliVue MX800 | VO2m²
8550 | Patient O2 Consumption per Body Weight | - |  | Double | VitalSigns | IntelliVue MX800 | VO2kg
8560 | Ideal Patient Weight | - |  | Double | PatientInfo | IntelliVue MX800 | IBW
8562 | End Flow Setting | - |  | Double | Settings | IntelliVue MX800 | sEndFl
8564 | PSV Backup Delay Setting | - |  | Double | Settings | IntelliVue MX800 | sPSVbd
8640 | Technical Alarm Internal Source ID | unitless | Consult detailed help for more information on this variable | Double (Multiset) | Alarms | Common | Alarms
8641 | Technical Alarm Source Type | unitless |  | Double (Multiset) | Alarms | Common | Alarms
8642 | Device Technical Alarm | unitless | Consult detailed help for more information on this variable | Double (Multiset) | Alarms | Common | Alarms
8643 | Patient Alarm Internal Source ID | unitless | Consult detailed help for more information on this variable | Double (Multiset) | Alarms | Common | Alarms
8644 | Patient Alarm Source Type | unitless |  | Double (Multiset) | Alarms | Common | Alarms
8645 | Patient Alarm | unitless | Consult detailed help for more information on this variable | Double (Multiset) | Alarms | Common | Alarms
8656 | Accoustic Respiration Rate Wave | - |  | Double | Waveforms | IntelliVue MX800 | Waves: acRESP
8703 | Blood Temperature Wave | - |  | Double | Waveforms | IntelliVue MX800 | Waves: Tblood
8714 | Non-specific Pressure 1 Wave | - |  | Double | Waveforms | IntelliVue MX800 | Waves: P1
8715 | Generic Invasive Blood Pressure P1 | - | Available on IntelliVue devices. | Double | VitalSigns | IntelliVue MX800 | IBP: P1
8716 | Generic Invasive Blood Pressure P1 Systolic | - |  | Double | VitalSigns | IntelliVue MX800 | IBP: P1
8717 | Generic Invasive Blood Pressure P1 Diastolic | - |  | Double | VitalSigns | IntelliVue MX800 | IBP: P1
8718 | Generic Invasive Blood Pressure P1 Mean | - |  | Double | VitalSigns | IntelliVue MX800 | IBP: P1
8719 | Non-specific Pressure 2 Wave | - |  | Double | Waveforms | IntelliVue MX800 | Waves: P2
8720 | Generic Invasive Blood Pressure P2 | - | Available on IntelliVue devices. | Double | VitalSigns | IntelliVue MX800 | IBP: P2
8721 | Generic Invasive Blood Pressure P2 Systolic | - |  | Double | VitalSigns | IntelliVue MX800 | IBP: P2
8722 | Generic Invasive Blood Pressure P2 Diastolic | - |  | Double | VitalSigns | IntelliVue MX800 | IBP: P2
8723 | Generic Invasive Blood Pressure P2 Mean | - |  | Double | VitalSigns | IntelliVue MX800 | IBP: P2
8724 | Non-specific Pressure 3 Wave | - |  | Double | Waveforms | IntelliVue MX800 | Waves: P3
8725 | Generic Invasive Blood Pressure P3 | - | Available on IntelliVue devices. | Double | VitalSigns | IntelliVue MX800 | IBP: P3
8726 | Generic Invasive Blood Pressure P3 Systolic | - |  | Double | VitalSigns | IntelliVue MX800 | IBP: P3
8727 | Generic Invasive Blood Pressure P3 Diastolic | - |  | Double | VitalSigns | IntelliVue MX800 | IBP: P3
8728 | Generic Invasive Blood Pressure P3 Mean | - |  | Double | VitalSigns | IntelliVue MX800 | IBP: P3
8729 | Non-specific Pressure 4 Wave | - |  | Double | Waveforms | IntelliVue MX800 | Waves: P4
8730 | Generic Invasive Blood Pressure P4 | - | Available on IntelliVue devices. | Double | VitalSigns | IntelliVue MX800 | IBP: P4
8731 | Generic Invasive Blood Pressure P4 Systolic | - |  | Double | VitalSigns | IntelliVue MX800 | IBP: P4
8732 | Generic Invasive Blood Pressure P4 Diastolic | - |  | Double | VitalSigns | IntelliVue MX800 | IBP: P4
8733 | Generic Invasive Blood Pressure P4 Mean | - |  | Double | VitalSigns | IntelliVue MX800 | IBP: P4
8734 | Non-specific Pressure 5 Wave | - |  | Double | Waveforms | IntelliVue MX800 | Waves: P_5
8735 | Generic Invasive Blood Pressure P5 | - | Available on IntelliVue devices. | Double | VitalSigns | IntelliVue MX800 | IBP: P5
8736 | Generic Invasive Blood Pressure P5 Systolic | - |  | Double | VitalSigns | IntelliVue MX800 | IBP: P5
8737 | Generic Invasive Blood Pressure P5 Diastolic | - |  | Double | VitalSigns | IntelliVue MX800 | IBP: P5
8738 | Generic Invasive Blood Pressure P5 Mean | - |  | Double | VitalSigns | IntelliVue MX800 | IBP: P5
8739 | Non-specific Pressure 6 Wave | - |  | Double | Waveforms | IntelliVue MX800 | Waves: P_6
8740 | Generic Invasive Blood Pressure P6 | - | Available on IntelliVue devices. | Double | VitalSigns | IntelliVue MX800 | IBP: P6
8741 | Generic Invasive Blood Pressure P6 Systolic | - |  | Double | VitalSigns | IntelliVue MX800 | IBP: P6
8742 | Generic Invasive Blood Pressure P6 Diastolic | - |  | Double | VitalSigns | IntelliVue MX800 | IBP: P6
8743 | Generic Invasive Blood Pressure P6 Mean | - |  | Double | VitalSigns | IntelliVue MX800 | IBP: P6
8744 | Non-specific Pressure 7 Wave | - |  | Double | Waveforms | IntelliVue MX800 | Waves: P_7
8745 | Generic Invasive Blood Pressure P7 | - | Available on IntelliVue devices. | Double | VitalSigns | IntelliVue MX800 | IBP: P7
8746 | Generic Invasive Blood Pressure P7 Systolic | - |  | Double | VitalSigns | IntelliVue MX800 | IBP: P7
8747 | Generic Invasive Blood Pressure P7 Diastolic | - |  | Double | VitalSigns | IntelliVue MX800 | IBP: P7
8748 | Generic Invasive Blood Pressure P7 Mean | - |  | Double | VitalSigns | IntelliVue MX800 | IBP: P7
8749 | Non-specific Pressure 8 Wave | - |  | Double | Waveforms | IntelliVue MX800 | Waves: P_8
8750 | Generic Invasive Blood Pressure P8 | - | Available on IntelliVue devices. | Double | VitalSigns | IntelliVue MX800 | IBP: P8
8751 | Generic Invasive Blood Pressure P8 Systolic | - |  | Double | VitalSigns | IntelliVue MX800 | IBP: P8
8752 | Generic Invasive Blood Pressure P8 Diastolic | - |  | Double | VitalSigns | IntelliVue MX800 | IBP: P8
8753 | Generic Invasive Blood Pressure P8 Mean | - |  | Double | VitalSigns | IntelliVue MX800 | IBP: P8
8754 | Intra-Uterine Pressure (IUP) Wave | - |  | Double | Waveforms | IntelliVue MX800 | Waves: IUP
8755 | Intra-Uterine Pressure | - |  | Double | VitalSigns | IntelliVue MX800 | IBP: IUP
8756 | Intra-Uterine Pressure Systolic | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
8757 | Intra-Uterine Pressure Diastolic | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
8758 | Intra-Uterine Pressure Mean | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
8760 | Unspecified Blood Pressure | - |  | Double | VitalSigns | IntelliVue MX800 | IBP: P
8761 | Unspecified Blood Pressure Systolic | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
8762 | Unspecified Blood Pressure Diastolic | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
8763 | Unspecified Blood Pressure Mean | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
8788 | Airway Flow Wave | - |  | Double | Waveforms | IntelliVue MX800 | Waves: AWF
8789 | Auxiliary Wave/Parameter | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
8790 | Global Ejection Fraction (2) | - | This is the device's measurement with label="NLS_NOM_EMFC_GEF" and physio_id="NOM_METRIC_NOS". See also 7810. | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
8791 | Pulmonary Vascular Permeability Index (2) | - | This is the device's measurement with label="NLS_NOM_EMFC_PVPI" and physio_id="NOM_METRIC_NOS". See also variable 6493. | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
8793 | Pressure Drop | - |  | Double | OtherMeasuredData | IntelliVue MX800 | DP
8803 | Replacement Pre-delivered Rate Setting | - |  | Double | Settings | IntelliVue MX800 | sRPrFl
8804 | Replacement Post-delivered Rate Setting | - |  | Double | Settings | IntelliVue MX800 | sRPoFl
8805 | Pre-blood Flow Rate Setting | - |  | Double | Settings | IntelliVue MX800 | sPBPFl
8806 | Patient Weight Balance Setting | - |  | Double | Settings | IntelliVue MX800 | sWTbal
8808 | Anticoagulant Bolus Interval Setting | - |  | Double | Settings | IntelliVue MX800 | sBolIn
8809 | Anticoagulant Infusion Rate Setting | - |  | Double | Settings | IntelliVue MX800 | sHepRt
8810 | Ultrafiltration Ratio Setting | - |  | Double | Settings | IntelliVue MX800 | sUFR
8811 | Therapy Elapsed Time | - |  | Double | OtherMeasuredData | IntelliVue MX800 | RunTi
8812 | Airway Volume | - |  | Double | VitalSigns | IntelliVue MX800 | Available only in revF or earlier
8813 | Processed Cumulated Blood Volume | - |  | Double | VitalSigns | IntelliVue MX800 | BVolPr
8814 | Replacement Pre-delivered Cumulated Volume | - |  | Double | VitalSigns | IntelliVue MX800 | RpreC
8815 | Replacement Post-delivered Cumulated Volume | - | This variable is associated with ChannelID identifier. | Double | VitalSigns | IntelliVue MX800 | Vdpo/h / RpostC
8816 | Pre-blood Cumulated Volume | - | This variable is associated with ChannelID identifier. | Double | VitalSigns | IntelliVue MX800 | VPBP/h / PreBC
8835 | Syringe Speed | - |  | Double | OtherMeasuredData | IntelliVue MX800 | Vsyr/h
8851 | Maximum Volume Gain Setting | - |  | Double | Settings | IntelliVue MX800 | sMaxG
8853 | Patient Fluid Balance | - |  | Double | VitalSigns | IntelliVue MX800 | Wtcum
8857 | Arterial Blood Pressure (ABP) Wave | - |  | Double | Waveforms | IntelliVue MX800 | Waves: ABP
8858 | Arterial Blood Pressure (ART) Wave | - |  | Double | Waveforms | IntelliVue MX800 | Waves: ART
8859 | Pulmonary Arterial Pressure (PAP) Wave | - |  | Double | Waveforms | IntelliVue MX800 | Waves: PAP
8860 | Central Venous Pressure (CVP) Wave | - |  | Double | Waveforms | IntelliVue MX800 | Waves: CVP
8861 | Right Atrial Pressure (RAP) Wave | - |  | Double | Waveforms | IntelliVue MX800 | Waves: RAP
8862 | Left Atrial Pressure (LAP) Wave | - |  | Double | Waveforms | IntelliVue MX800 | Waves: LAP
8863 | Intra-cranial Pressure (ICP) Wave | - |  | Double | Waveforms | IntelliVue MX800 | Waves: ICP
8864 | Umbilical Arterial Pressure (UAP) Wave | - |  | Double | Waveforms | IntelliVue MX800 | Waves: UAP
8865 | Umbilical Venous Pressure (UVP) Wave | - |  | Double | Waveforms | IntelliVue MX800 | Waves: UVP
8866 | Airway CO2 Level Waveform | - |  | Double | Waveforms | IntelliVue MX800 | Waves: CO2
8867 | Airway O2 Wave | - |  | Double | Waveforms | IntelliVue MX800 | Waves: O2
8868 | Airway N2 Wave | - |  | Double | Waveforms | IntelliVue MX800 | Waves: N2
8869 | Generic Agent Wave (2nd) | - |  | Double | Waveforms | IntelliVue MX800 | Waves: AGTs
8870 | Impedance Respiration Wave | - |  | Double | Waveforms | IntelliVue MX800 | Waves: Resp
8871 | Brachial Arterial Pressure (BAP) Wave | - |  | Double | Waveforms | IntelliVue MX800 | Waves: BAP
8872 | Intracranial Pressure 1 (IC1) Wave | - |  | Double | Waveforms | IntelliVue MX800 | Waves: IC1
8873 | Intracranial Pressure 2 (IC2) Wave | - |  | Double | Waveforms | IntelliVue MX800 | Waves: IC2
8874 | Femoral Arterial Pressure (FAP) Wave | - |  | Double | Waveforms | IntelliVue MX800 | Waves: FAP
8875 | Vector ECG Wave | - |  | Double | Waveforms | IntelliVue MX800 | Waves: vECG
8889 | Transcutaneous Heater Power | - |  | Double | OtherMeasuredData | IntelliVue MX800 | HeatPa
8939 | BIS Standard Deviation | - | This variable is associated with ChannelID identifier. | Double | VitalSigns | IntelliVue MX800 | sdBISl/sdBISr
8940 | EMG Standard Deviation (BIS) | - | This variable is associated with ChannelID identifier. | Double | VitalSigns | IntelliVue MX800 | sdEMGl/sdEMGr
8941 | Asymmetric Index | - | This variable is associated with ChannelID identifier. | Double | VitalSigns | IntelliVue MX800 | ASYM/ASYM L/ASYM R
9037 | O2 Consumption Index Estimate | - |  | Double | VitalSigns | IntelliVue MX800 | iVO2Ie
9110 | Aortic Blood Pressure (Ao) | - | Available on IntelliVue devices with firmware revision G or later. Arterial Blood Pressure in the Aorta (NOM_PRESS_BLD_AORT). | Double | VitalSigns | Common | IBP: Ao
9111 | Umbilical Arterial Pressure (UAP) | - | Available on IntelliVue devices with firmware revision G or later. | Double | VitalSigns | Common | IBP: UAP
9112 | Umbilical Venous Pressure (UVP) | - | Available on IntelliVue devices with firmware revision G or later. | Double | VitalSigns | Common | IBP: UVP
9154 | Humidifier Activation Status | - |  | Double | DeviceInfo | IntelliVue MX800 | sHmCtl
9155 | Nebulizer Time Setting | - |  | Double | Settings | IntelliVue MX800 | sNebTi
9165 | Edi Diaphragm Electrical Activity Waveform | - |  | Double | Waveforms | IntelliVue MX800 | Waves: EDI
9166 | Peak Diaphragm Electrical Activity (Edi) | - |  | Double | VitalSigns | IntelliVue MX800 | Edimax
9167 | Minimum Diaphragm Electrical Activity (Edi) | - |  | Double | VitalSigns | IntelliVue MX800 | Edimin
9168 | % Edi Trigger In NAVA | - |  | Double | VitalSigns | IntelliVue MX800 | %EdiTr
9169 | % Edi Cycle Off In NAVA | - |  | Double | VitalSigns | IntelliVue MX800 | %EdiCy
9170 | NAVA Level Setting | - |  | Double | Settings | IntelliVue MX800 | sNAVA
9171 | Edi Trigger Setting | - |  | Double | Settings | IntelliVue MX800 | sEdiTr
9196 | Patient Body Mass Index (BMI) | - |  | Double | PatientInfo | IntelliVue MX800 | BMI
9270 | Maximum Positive Auxiliary Pressure | - |  | Double | VitalSigns | IntelliVue MX800 | Pxpeak
9271 | Mean Positive Auxiliary Pressure | - |  | Double | VitalSigns | IntelliVue MX800 | Pxmean
9272 | Minimum Positive Auxiliary Pressure | - |  | Double | VitalSigns | IntelliVue MX800 | Pxmin
9286 | Patient CO2 Production Indexed By Body Surface | - |  | Double | VitalSigns | IntelliVue MX800 | VCO2m²
9287 | Patient CO2 Production Indexed By Body Weight | - |  | Double | VitalSigns | IntelliVue MX800 | VCO2kg
9288 | Tidal Volume Indexed By Body Weight | - |  | Double | VitalSigns | IntelliVue MX800 | TV/kg
9314 | Functional Residual Capacity | - |  | Double | VitalSigns | IntelliVue MX800 | FRC
9315 | Corrected QT Interval (QTc) | - |  | Double | VitalSigns | IntelliVue MX800 | ECG: QTc
9317 | Airway N2 Concentration | - | Measurement from VueLink | Double | VitalSigns | IntelliVue MX800 | N2
9318 | Airway Agent Concentration | - | Measurement from VueLink. This variable is associated with a SourceID identifier. | Double | VitalSigns | IntelliVue MX800 | AGT
9321 | Respiration Rate (Spirometry) | - |  | Double | VitalSigns | IntelliVue MX800 | RRspir
9323 | QT Heart Rate (QT-HR) | - | QT heart rate | Double | VitalSigns | IntelliVue MX800 | ECG: QT-HR
9324 | Delta QTc | - |  | Double | VitalSigns | IntelliVue MX800 | ECG: DeltaQTc
9329 | Humidification Type Setting | - |  | Double | Settings | IntelliVue MX800 | sHumid
9336 | Normalized Rapid Shallow Breathing Index | - |  | Double | VitalSigns | IntelliVue MX800 | RSBInm
9337 | Percent of Spontaneous Inspiratory Time | - |  | Double | VitalSigns | IntelliVue MX800 | Tin/Tt
9408 | Energy Balance Calculated by Metabolics | - |  | Double | VitalSigns | IntelliVue MX800 | dE
9437 | Alveolar Volume | - |  | Double | VitalSigns | IntelliVue MX800 | Valv
9438 | Apparatus Dead Space | - |  | Double | VitalSigns | IntelliVue MX800 | appVd
9439 | Pleth (Pre-Ductal) | - | Pulse Rate from Plethysmogram (pre ductal) | Double | VitalSigns | IntelliVue MX800 | Pulse
9440 | Pleth (Post-Ductal) | - | Pulse Rate from Plethysmogram (post ductal) | Double | VitalSigns | IntelliVue MX800 | Pulse
9441 | Relative Perfusion (Pre-Ductal) | - |  | Double | VitalSigns | IntelliVue MX800 | PerfPr
9442 | Relative Perfusion (Post-Ductal) | - |  | Double | VitalSigns | IntelliVue MX800 | PerfPo
9444 | Cardioplegia Main Pump Time of Current Plegia Session | - |  | Double | VitalSigns | IntelliVue MX800 | CpOnTi
9445 | Cardioplegia Main Pump Time Since Stop of Plegia Session | - |  | Double | VitalSigns | IntelliVue MX800 | CpOffT
9446 | Cardioplegia Main Pump Total Plegia Time | - |  | Double | VitalSigns | IntelliVue MX800 | CpPlTi
9447 | Cardioplegia Slave Pump Time Since Start of Current Plegia Session | - |  | Double | VitalSigns | IntelliVue MX800 | CsOnTi
9448 | Cardioplegia Slave Pump Time Since Stop of Plegia Session | - |  | Double | VitalSigns | IntelliVue MX800 | CsOffT
9449 | Cardioplegia Slave Pump Total Plegia Time | - |  | Double | VitalSigns | IntelliVue MX800 | CsPlTi
9453 | PAV-based Lung Compliance | - |  | Double | VitalSigns | IntelliVue MX800 | Cpav
9454 | PAV-based Patient Resistance | - |  | Double | VitalSigns | IntelliVue MX800 | Rpav
9455 | PAV-based Total Resistance Estimate | - |  | Double | VitalSigns | IntelliVue MX800 | Rtot
9456 | PAV-based Lung Elastance | - |  | Double | VitalSigns | IntelliVue MX800 | Epav
9467 | Backup Inspiratory Pressure Setting | - |  | Double | Settings | IntelliVue MX800 | sPIbak
9476 | Percent of Leakage | - |  | Double | OtherMeasuredData | IntelliVue MX800 | %Leak
9485 | Maximum Inspiratory Time Setting | - |  | Double | Settings | IntelliVue MX800 | sTimax
9585 | Plateau Time | - |  | Double | VitalSigns | IntelliVue MX800 | PlatTi
9586 | QT Baseline | - |  | Double | VitalSigns | IntelliVue MX800 | ECG: QT Bsl
9587 | QT Heart Rate Baseline | - |  | Double | VitalSigns | IntelliVue MX800 | ECG: QTHRBl
9588 | High Density Lipoprotein | - |  | Double | VitalSigns | IntelliVue MX800 | HDL
9589 | Low Density Lipoprotein | - |  | Double | VitalSigns | IntelliVue MX800 | LDL
9590 | C-reactive Protein | - |  | Double | VitalSigns | IntelliVue MX800 | CRP
9591 | Urine Hemoglobin | - |  | Double | VitalSigns | IntelliVue MX800 | UrHb
9595 | Noninvasive Total Hemoglobin (SpHb) | - |  | Double | VitalSigns | IntelliVue MX800 | SpHb
9596 | Noninvasive Oxygen Content (SpOC) | - |  | Double | VitalSigns | IntelliVue MX800 | SpOC
9597 | Noninvasive Carboxyhemoglobin (SpCO) | - |  | Double | VitalSigns | IntelliVue MX800 | SpCO
9598 | Noninvasive Methemoglobin (SpMet) | - |  | Double | VitalSigns | IntelliVue MX800 | SpMet
9600 | Pleth Variability Index (PVI) | - |  | Double | VitalSigns | IntelliVue MX800 | PVI
9603 | Regional O2 Saturation (rSO2) | - |  | Double | VitalSigns | Common
9605 | rSO2 Area Under Curve (AUC) | - | This variable is associated with ChannelID identifier. | Double | VitalSigns | IntelliVue MX800 | AUC_1/AUC_2/AUC_3/AUC_4/AUC_5/AUC_6
9610 | Generic O2 Saturation (3) | - |  | Double | VitalSigns | IntelliVue MX800 | SO2 3
9611 | Generic O2 Saturation (4) | - |  | Double | VitalSigns | IntelliVue MX800 | SO2 4
9612 | Tissue O2 Saturation (StO2) | - |  | Double | VitalSigns | IntelliVue MX800 | StO2
9613 | Nitrogen Balance (NsLoss) | - |  | Double | VitalSigns | IntelliVue MX800 | NsLoss
52523 | Inspired O2 Concentration | % | See variable 52527 for partial pressure | Double | VitalSigns | IntelliVue MX800 | O2
52527 | Inspired O2 Partial Pressure | mmHg | This variable is calculated from variable 52523 and will output inspired O2 partial pressure. | Double | VitalSigns | IntelliVue MX800 | O2

## PiCCO2
ID | Label | Unit | Description | Type | Group | Device | Source
1 | Heart Rate | 1/min |  | Short | VitalSigns | Common | Continuous Parameter
14 | SpO2 | % |  | Double | VitalSigns | PiCCO2 | Continuous Parameter
414 | Cardiac Output | L/min |  | Double | VitalSigns | Common | Discontinuous Mean Parameter
519 | Central Venous Pressure (CVP) | mmHg | This variable transmits either a continuous measurement or a manual input value. | Short | VitalSigns | Common | Continuous Parameter
585 | Patient Category | unitless | Consult detailed help for more information on this variable | Short | PatientInfo | PiCCO2 | Patient Setting
649 | Cardiac Index | L/min/m2 |  | Double | VitalSigns | PiCCO2 | Discontinuous Mean Parameter
650 | Stroke Volume | mL | Beat to beat stroke volume. | Double | VitalSigns | Common | Continuous Parameter
654 | Systemic Vascular Resistance | dyn·s/cm5 |  | Short | VitalSigns | Common | Continuous Parameter
756 | Patient Height | cm |  | Short | PatientInfo | Common | Patient Setting
757 | Patient Weight | kg |  | Double | PatientInfo | Common | Patient Setting
768 | Generic Error Code | unitless | TD measurement error, "0" is transmitted if no error occurs. | Short | Alarms | Common | Error
1246 | Hematocrit | % |  | Double | VitalSigns | PiCCO2 | ScvO2 Calibration
1434 | Patient Body Surface Area | m2 |  | Double | PatientInfo | PiCCO2 | Patient Calculation
1435 | Systemic Vascular Resistance Index | dyn·s·m2/cm5 |  | Short | VitalSigns | PiCCO2 | Continuous Parameter
1468 | O2 Consumption (VO2) | mL/min |  | Short | VitalSigns | PiCCO2 | Continuous Parameter
1472 | O2 Consumption Index | mL/min/m2 |  | Short | VitalSigns | PiCCO2 | Continuous Parameter
1710 | Cardiac Power Index (CCO) | Not Available |  | Double | VitalSigns | PiCCO2 | Continuous Parameter
1711 | Cardiac Power Output (CCO) | W |  | Double | VitalSigns | PiCCO2 | Continuous Parameter
1720 | ICG Measurement Counter | unitless | Counter for signalizing update of ICG measurement. | Short | OtherMeasuredData | PiCCO2 | Discontinuous ICG
1721 | Plasma Disappearance Rate | Not Available |  | Double | VitalSigns | PiCCO2 | Discontinuous ICG
1722 | ICG Retention Rate | % |  | Double | VitalSigns | PiCCO2 | Discontinuous ICG
1730 | Cardiac Index ((Thermodilution) | L/min/m2 |  | Double | VitalSigns | PiCCO2 | Discontinuous TD
1731 | Cardiac Output (Thermodilution) | L/min |  | Double | VitalSigns | PiCCO2 | Discontinuous TD
1732 | Global Ejection Fraction (Thermodilution) | % |  | Short | VitalSigns | PiCCO2 | Discontinuous TD
1733 | Global End-Diastolic Volum (Thermodilution) | mL |  | Short | VitalSigns | PiCCO2 | Discontinuous TD
1734 | Global End-Diastolic Volum Index (Thermodilution) | mL/m2 |  | Short | VitalSigns | PiCCO2 | Discontinuous TD
1735 | Pulmonary Vascular Permeability Index (Thermodilution) | unitless |  | Double | VitalSigns | PiCCO2 | Discontinuous TD
1821 | Generic ICG Error | unitless | ICG measurement error, "0" is transmitted if no error occurs. | Short | Alarms | PiCCO2 | Error
1929 | Patient ID | unitless | 9 digits maximum, only contain numbers from 0 to 9. | String | PatientInfo | Common | Patient Setting
1930 | Patient Name | unitless | Allowed characters include A..Z, 0..9, ' ', '.', and '-'. The device sends a single character each 400 ms, so it takes time to transmit the full patient name. | String | PatientInfo | PiCCO2 | Patient Setting
1943 | Predicted Patient Body Surface Area | m2 |  | Double | PatientInfo | PiCCO2 | Patient Calculation
1944 | Predicted Patient Weight | kg |  | Double | PatientInfo | PiCCO2 | Patient Calculation
2255 | Device Date/Time | unitless | This variable contains the current time of the device in the form of "MM/dd/yyyy - hh:mm:ss". It is also sent in the "DeviceTime" attribute of other variables. | String | DeviceInfo | PiCCO2 | Device Information
2259 | Arterial BP Systolic | mmHg |  | Double | VitalSigns | Common | Continuous Parameter
2260 | Arterial BP Diastolic | mmHg |  | Double | VitalSigns | Common | Continuous Parameter
2261 | Arterial BP Mean | mmHg |  | Double | VitalSigns | Common | Continuous Parameter
2341 | Patient Blood Temperature | °C | A offset of 30°C is already taken into account by the DDI and consequently, the user receives the blood temperature real value, without having to interpret it. | Double | VitalSigns | Common | Discontinuous TD
2342 | Injectate Temperature | °C | A offset of 15°C is already taken into account by the DDI and consequently, the user receives the injectate temperature real value, without having to interpret it. | Double | OtherMeasuredData | Common | Discontinuous TD
2452 | O2 Delivery (DO2) | mL/min |  | Short | VitalSigns | PiCCO2 | Continuous Parameter
2462 | SvO2 Alarm | unitless | High/Low Alarm of ScvO2 Consult detailed help for more information on this variable | Short | Alarms | PiCCO2 | Alarm
2819 | SvO2 Calibration Status | unitless | ScvO2Cal: Point of time of calibration. "1" once during “Bloodsample drawn”, otherwise "0". | Boolean | DeviceInfo | PiCCO2 | ScvO2 Calibration
2879 | Device Software Revision | unitless | Release PiCCO software version, composed by one letter (type of protocol: V for usual protocol, D for demo version and X for unreleased version) and from one to three digits seperated by dots (version number). For example: D5.1 | String | DeviceInfo | Common | Device Information
2915 | Ejection Fraction | % | Global Ejection Fraction | Short | VitalSigns | PiCCO2 | Discontinuous Mean Parameter
3184 | Alternate Patient ID | unitless | Alternate display of Patient ID. Allowed characters include A..Z, 0..9, ' ', '.', and '-'. The device sends a single character each 400 ms, so it takes time to transmit the full patient ID. | String | PatientInfo | PiCCO2 | Patient Setting
3789 | Arterial O2 Saturation (SaO2) | % | Arterial oxygenation saturation (input value for display of parameters DO2 and VO2). | Double | VitalSigns | PiCCO2 | ScvO2 Calibration
3877 | Device Protocol Revision | unitless | Release PiCCO2 software version, composed by one letter (type of protocol: V for usual protocol, D for demo version and X for unreleased version) and from one to three digits seperated by dots (version number). For example: D5.1 | String | DeviceInfo | PiCCO2 | Device Information
4532 | Stroke Volume Variation | % |  | Short | VitalSigns | Common | Continuous Parameter
4575 | O2 Delivery Index (DO2I) | mL/min/m2 |  | Short | VitalSigns | PiCCO2 | Continuous Parameter
4865 | Cardiac Function Index (Thermodilution) | 1/min |  | Double | VitalSigns | PiCCO2 | Discontinuous TD
4866 | Intrathoracic Blood Volume (Thermodilution) | mL |  | Short | VitalSigns | PiCCO2 | Discontinuous TD
4867 | Intrathoracic Blood Volume Index (Thermodilution) | mL/m2 |  | Short | VitalSigns | PiCCO2 | Discontinuous TD
4868 | Extravascular Lung Water (Thermodilution) | mL |  | Short | VitalSigns | PiCCO2 | Discontinuous TD
4869 | Extravascular Lung Water Index (Thermodilution) | mL/kg |  | Short | VitalSigns | PiCCO2 | Discontinuous TD
5072 | Intrathoracic Blood Volume (ITBV) | mL |  | Short | VitalSigns | Common | Discontinuous Mean Parameter
5073 | Intrathoracic Blood Volume Index (ITBI) | mL/m2 |  | Short | VitalSigns | PiCCO2 | Discontinuous Mean Parameter
5074 | Extravascular Lung Water | mL |  | Short | VitalSigns | Common | Discontinuous Mean Parameter
5075 | Extravascular Lung Water Index | mL/kg |  | Short | VitalSigns | PiCCO2 | Discontinuous Mean Parameter
5076 | Global End Diastolic Volume | mL |  | Short | VitalSigns | Common | Discontinuous Mean Parameter
5077 | Global End Diastolic Volume Index | mL/m2 |  | Short | VitalSigns | PiCCO2 | Discontinuous Mean Parameter
5078 | Cardiac Function Index | 1/min |  | Double | VitalSigns | PiCCO2 | Discontinuous Mean Parameter
5642 | SvO2 Calibration Data | unitless | ScvO2Cal: ScvO2 Calibration input. | Double | DeviceInfo | PiCCO2 | ScvO2 Calibration
5684 | Continuous Cardiac Output (CCO) Alarm | unitless | High/Low Alarm of PCCO Consult detailed help for more information on this variable | Short | Alarms | PiCCO2 | Alarm
6403 | Pulse Contour Cardiac Output | L/min |  | Double | VitalSigns | Common | Continuous Parameter
6404 | Thermodilution Measurement Status | unitless | Consult detailed help for more information on this variable | Short | OtherMeasuredData | Common | Discontinuous TD
6407 | Index Of Left Ventricular Contractility (dPmx) | mmHg/s |  | Short | VitalSigns | Common | Continuous Parameter
6410 | Mean Transit Time (MTt) | s |  | Double | VitalSigns | Common | Discontinuous TD
6411 | Down Slope Time | s |  | Double | VitalSigns | Common | Discontinuous TD
6422 | Temperature Difference (Thermodilution) | °C |  | Double | VitalSigns | PiCCO2 | Discontinuous TD
6424 | Injectate Temperature (Thermodilution) | °C |  | Double | OtherMeasuredData | PiCCO2 | Discontinuous TD
6428 | PICCO Catheter | unitless | Range 0-14 Consult detailed help for more information on this variable | Short | DeviceInfo | Common | Device Information
6434 | Pulse Pressure Variation | % |  | Short | VitalSigns | Common | Continuous Parameter
6493 | Pulmonary Vascular Permeability Index | unitless |  | Double | VitalSigns | PiCCO2 | Discontinuous Mean Parameter
6533 | Patient Gender | unitless | Consult detailed help for more information on this variable | Short | PatientInfo | PiCCO2 | Patient Setting
6689 | Pulse Contour Cardiac Index | L/min/m2 |  | Double | VitalSigns | PiCCO2 | Continuous Parameter
6690 | Thermodilution Measurement Counter | unitless | Counter for signalizing when all discontinuous thermodilution parameters have been updated. | Short | OtherMeasuredData | PiCCO2 | Discontinuous TD
6692 | Right to Left Heart Shunt Fraction | % |  | Short | VitalSigns | PiCCO2 | Discontinuous TD
6694 | Arterial Blood Pressure Wave | mmHg |  | Double | Waveforms | Common | Continuous Parameter
6695 | Injectate Volume Setting | mL |  | Double | Settings | Common | Discontinuous TD
6696 | Stroke Volume Min | mL |  | Double | VitalSigns | Common | Continuous Parameter
6697 | Stroke Volume Max | mL |  | Double | VitalSigns | Common | Continuous Parameter
6698 | Stroke Volume Index | mL/m2 |  | Double | VitalSigns | PiCCO2 | Continuous Parameter
6699 | Stroke Volume Mean | mL |  | Double | VitalSigns | Common | Continuous Parameter
7823 | Central Venous O2 Saturation (ScvO2) | % |  | Double | VitalSigns | PiCCO2 | Continuous Parameter
7941 | SpO2 Alarm | unitless | High/Low Alarm of SpO2 Consult detailed help for more information on this variable | Short | Alarms | PiCCO2 | Alarm

## V60
ID | Label | Unit | Description | Type | Group | Device | Source
542 | Maximum Tidal Volume Limit | - | Possible units: mL, L | Double | Settings | Common
955 | Battery Running Alarm | unitless | Consult detailed help for more information on this variable | Long | Alarms | Common
2032 | Volume Not Delivered Alarm | unitless | Consult detailed help for more information on this variable | Long | Alarms | Common
2445 | Patient Rebreathing Alarm | unitless | Consult detailed help for more information on this variable | Long | Alarms | Common
2595 | General Alarm Status | unitless | Contains every alarm code that is currently active | Long (Multiset) | Alarms | Common
Consult detailed help for more information on this variable
2784 | ASB Ramp Time Setting | - | Possible units: msec, sec, min | Double | Settings | Common
3147 | Maximum Lung Elastance Setting | - | Possible units: cmH2O/L, mBar/L | Double | Settings | Common
3203 | Maximum Airway Resistance Setting | - | Possible units: cmH2O/L/s, mBar/L/s | Double | Settings | Common
3592 | Low Airway Flow Alarm | unitless | Consult detailed help for more information on this variable | Long | Alarms | V60 (PVOI)
3640 | Pressure Alarm Delay | - |  | Double | DeviceInfo | Common
4468 | Patient Alarm Global State | unitless | Alarm disabled status | Long | Alarms | V60 (PVOI)
Consult detailed help for more information on this variable
4957 | Trigger Sensitivity Setting | - |  | Double | Settings | Common
6288 | Percent of Patient Triggered Breaths | % |  | Double | OtherMeasuredData | Common
6935 | Estimated Patient Circuit Leakage | - | Possible units: mL/min, L/min | Double | OtherMeasuredData | Common
7835 | Ventilation Comfort Mode Flex Setting | - |  | Double | Settings | Common
8153 | IPAP Minimum Pressure Setting | - | Possible units: cmH2O, hPa, kPa, mBar, mmHg | Double | Settings | Common
8261 | Patient Expiratory Port Type | unitless | Consult detailed help for more information on this variable | Long | DeviceInfo | Common
8590 | Vt Not Achieved Alarm | unitless | Consult detailed help for more information on this variable | Long | Alarms | Common
8646 | Patient Interface Type | unitless | Consult detailed help for more information on this variable | Long | DeviceInfo | Common
8653 | Tidal Volume Exceeded Alarm | unitless | Consult detailed help for more information on this variable | Long | Alarms | Common
9882 | Check Maximum Pressure Alarm | unitless | Consult detailed help for more information on this variable | Long | Alarms | Common
9955 | Maximum IPAP Setting | - | Possible units: cmH2O, hPa, kPa, mBar, mmHg | Double | Settings | Common
31065 | Low Ventilator Total Respiration Rate Alarm | unitless | Consult detailed help for more information on this variable | Long | Alarms | Common
32211 | Low Expired Tidal Volume Alarm | unitless | Consult detailed help for more information on this variable | Long | Alarms | Common
32221 | Low Expired Minute Volume Alarm | unitless | Consult detailed help for more information on this variable | Long | Alarms | Common
32311 | Low Peak Inspiratory Pressure Alarm | unitless | Consult detailed help for more information on this variable | Long | Alarms | Common
50203 | Device Model | unitless |  | String | DeviceInfo | Common
50212 | Device Software Revision | unitless |  | String | DeviceInfo | Common
51065 | Ventilator Total Respiration Rate | 1/min |  | Double | VitalSigns | Common
52127 | Inspiration Time Percentage | % |  | Double | VitalSigns | Common
52211 | Expired Tidal Volume | mL |  | Double | VitalSigns | Common
52221 | Expired Minute Volume | L/min |  | Double | VitalSigns | Common
52311 | Peak Inspiratory Pressure | cmH2O |  | Double | VitalSigns | Common
52911 | Airway Flow Waveform | L/min |  | Double | Waveforms | Common
52912 | Airway Volume Waveform | mL |  | Double | Waveforms | Common
52913 | Airway Pressure Waveform | cmH2O |  | Double | Waveforms | Common
54011 | Ventilation Standby State | unitless |  | Boolean | VitalSigns | Common
54013 | Ventilation Mode | unitless | Consult detailed help for more information on this variable | String | Settings | Common
54021 | Ventilation Mode Primary | unitless | Consult detailed help for more information on this variable | String | Settings | Common
54022 | Ventilation Mode Options | unitless | Consult detailed help for more information on this variable | String | Settings | Common
54051 | Inspired O2 (FiO2) Setting | % |  | Double | Settings | Common
54061 | Ventilation Leakage Minute Volume | L/min |  | Double | VitalSigns | Common
54111 | Mandatory Respiratory Rate Setting | 1/min |  | Double | Settings | Common
54121 | Inspiratory Time Setting | s |  | Double | Settings | Common
54123 | lnspiratory Rise Time Setting | s |  | Double | Settings | Common
54151 | Inspired Tidal Volume Setting | mL |  | Double | Settings | Common
54172 | CPAP Setting | cmH2O |  | Double | Settings | Common
54231 | Inspiratory Flow Termination Setting | L/min | Auto-Trak sensitivity: set flow for cycling to exhalation | Double | Settings | Common
54351 | Proportional Support Level Setting | % |  | Double | Settings | Common
54381 | EPAP Setting | cmH2O |  | Double | Settings | Common
54382 | IPAP Setting | cmH2O |  | Double | Settings | Common
54624 | Fresh Gas Total Flow Rate Setting | L/min |  | Double | Settings | V60 (PVOI)
71065 | High Ventilator Total Respiration Rate Alarm | unitless | Consult detailed help for more information on this variable | Long | Alarms | Common
72211 | High Expired Tidal Volume Alarm | unitless | Consult detailed help for more information on this variable | Long | Alarms | Common
72311 | High Peak Inspiratory Pressure Alarm | unitless | Consult detailed help for more information on this variable | Long | Alarms | Common
90714 | Battery Low Alarm | unitless | This variable is associated with a ChannelID identifier. | Long | Alarms | Common
Consult detailed help for more information on this variable
90719 | Main Power Disconnected Alarm | unitless | Consult detailed help for more information on this variable | Long | Alarms | Common
92001 | Ventilator Patient Disconnected Alarm | unitless | Consult detailed help for more information on this variable | Long | Alarms | Common
92011 | Ventilator Inoperative Alarm | unitless | Consult detailed help for more information on this variable | Long | Alarms | Common
92021 | Ventilator Obstruction Alarm | unitless | Consult detailed help for more information on this variable | Long | Alarms | Common
92405 | Pressure Sensor Disconnected Alarm | unitless | Consult detailed help for more information on this variable | Long | Alarms | Common
92612 | O2 Supply Alarm | unitless | Consult detailed help for more information on this variable | Long | Alarms | Common
92631 | O2 Supply High Alarm | unitless | Consult detailed help for more information on this variable | Long | Alarms | Common
92632 | O2 Supply Low Alarm | unitless | Consult detailed help for more information on this variable | Long | Alarms | Common

## Monnal T60
ID | Label | Unit | Description | Type | Group | Device | Source
20 | End Tidal CO2 (EtCO2) | - | This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. Possible units: cmH2O, mbar, bar, hPa, kPa, mmHg, % | Double | VitalSigns | Monnal T60
22 | Respiration Rate | - | This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. Possible units: 1/min | Double | VitalSigns | Common
544 | Minimum Respiration Rate Threshold Setting | - | Possible units: 1/min | Double | Settings | Common
584 | Ventilation Mode | unitless | Consult detailed help for more information on this variable | String | DeviceInfo | Common
635 | Inspired O2 (FiO2) | - | This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. Possible units: % | Double | VitalSigns | Common
754 | Device Name | unitless | ASCII string up to 24 characters long. | String | DeviceInfo | Common
784 | Inspiration Time Percent Setting | - | Possible units: % | Double | Settings | Common
955 | Battery Running Alarm | unitless |  | Boolean | Alarms | Common
1187 | Low Inspired Tidal Volume Alarm | unitless |  | Boolean | Alarms | Common
1188 | Low Expired Tidal Volume Alarm | unitless |  | Boolean | Alarms | Common
1189 | Positive End Expiratory Pressure (PEEP) | - | This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. Possible units: cmH2O, mbar, bar, hPa, kPa, mmHg | Double | VitalSigns | Common
1228 | High Inspired Tidal Volume Alarm | unitless |  | Boolean | Alarms | Common
1321 | Tidal Volume Setting | - | Possible units: mL, L | Double | Settings | Common
1324 | Sigh Interval Setting | - | Possible units: ms, sec, min, h | Double | Settings | Common
1327 | Sigh Pressure Limit Setting | - | Possible units: cmH2O, mbar, bar, hPa, kPa, mmHg | Double | Settings | Common
1328 | Apnea Tidal Volume Setting | - | Possible units: mL, L | Double | Settings | Common
1329 | Apnea Respiration Rate Setting | - | Possible units: 1/min | Double | Settings | Common
1370 | Battery Low Alarm | unitless |  | Boolean | Alarms | Common
1383 | Inspiratory Flow Trigger Setting | - | Possible units: mL/min, mL/h, L/s, L/min, L/h | Double | Settings | Common
1391 | Apnea Time Setting | - | Possible units: ms, sec, min, h | Double | Settings | Common
1415 | Mean Airway Pressure | - | This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. Possible units: cmH2O, mbar, bar, hPa, kPa, mmHg | Double | VitalSigns | Monnal T60
1538 | O2 Sensor Alarm | unitless |  | Boolean | Alarms | Common
1657 | Expiratory Trigger Level Setting | - | Possible units: % | Double | Settings | Common
1891 | Device Power Supply Status | unitless | Value is Sector, Internal Battery, or External Battery | String | DeviceInfo | Common
1913 | Device Serial Number | unitless | ASCII string up to 24 characters long. | String | DeviceInfo | Common
2088 | SIMV Respiration Rate Setting | - | Possible units: 1/min | Double | Settings | Common
2092 | Inspired Minute Volume | - | This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. Possible units: mL, L | Double | VitalSigns | Common
2102 | Inspiratory Time Setting | - | Possible units: ms, sec, min, h | Double | Settings | Common
2104 | Positive End Expiratory Pressure (PEEP) Setting | - | Possible units: cmH2O, mbar, bar, hPa, kPa, mmHg | Double | Settings | Common
2110 | Inspired O2 (FiO2) Setting | - | Possible units: % | Double | Settings | Monnal T60
2120 | Battery Alarm | unitless | This alarm is raised if the internal battery is discharged. | Boolean | Alarms | Common
2324 | Inspired Tidal Volume | - | This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. Possible units: mL, L | Double | VitalSigns | Common
2325 | Expired Tidal Volume | - | This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. Possible units: mL, L | Double | VitalSigns | Common
2326 | Expired Minute Volume | - | This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. Possible units: mL, L | Double | VitalSigns | Common
2445 | Patient Rebreathing Alarm | unitless |  | Boolean | Alarms | Monnal T60
2546 | Low Expired Minute Volume Alarm | unitless |  | Boolean | Alarms | Common
2568 | Airway Peak Pressure Alarm | unitless |  | Boolean | Alarms | Common
2573 | Ventilator Control Mode | unitless | The variable indicates if the ventilator is in Non-Invasive Ventilation. ON means non-invasive ventilation and OFF means invasive ventilation. | String | DeviceInfo | Common
2582 | Target Respiration Rate | - | Possible units: 1/min | Double | Settings | Monnal T60
2583 | I:E Ratio Setting | unitless |  | Double | Settings | Common
2590 | Generic Device Alarm | unitless | This is a multiset variable containing the text messages for all active alarms. Associated with variable 6164 for the matching alarm codes. | String (Multiset) | Alarms | Common
2759 | Support Pressure Setting | - | Possible units: cmH2O, mbar, bar, hPa, kPa, mmHg | Double | Settings | Common
2879 | Device Software Revision | unitless | ASCII string up to 24 characters long. | String | DeviceInfo | Common
2960 | Tidal Volume Target | - | Possible units: mL, L | Double | Settings | Common
2964 | I:E Ratio | unitless |  | Double | VitalSigns | Common
2968 | Plateau Airway Pressure | - | This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. Possible units: cmH2O, mbar, bar, hPa, kPa, mmHg | Double | VitalSigns | Common
2976 | Patient Disconnect Alarm | unitless |  | Boolean | Alarms | Common
3853 | Patient Type | unitless | Value is ADULT, CHILD, INFANT, or BABY. | String | PatientInfo | Common
3858 | Volume Not Achieved Alarm | unitless |  | Boolean | Alarms | Common
3961 | Plateau Time Setting | - | Possible units: ms, sec, min, h | Double | Settings | Monnal T60
4004 | Breathing System Leakage | - | This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. Possible units: % | Double | OtherMeasuredData | Common
4367 | Respiration Rate Setting | - | Possible units: 1/min | Double | Settings | Common
4371 | Sigh Tidal Volume Setting | - | Possible units: mL, L | Double | Settings | Common
4391 | Apnea Ventilation Alarm | unitless |  | Boolean | Alarms | Common
4992 | Maximum Inspired Pressure Setting | - | Possible units: cmH2O, mbar, bar, hPa, kPa, mmHg | Double | Settings | Common
5010 | Inspiratory Pressure Setting | - | Possible units: cmH2O, mbar, bar, hPa, kPa, mmHg | Double | Settings | Common
5359 | Airway Peak Flow Setting | - | Possible units: mL/min, mL/h, L/s, L/min, L/h | Double | Settings | Monnal T60
5712 | Low EtCO2 Alarm | unitless |  | Boolean | Alarms | Monnal T60
5713 | High EtCO2 Alarm | unitless |  | Boolean | Alarms | Monnal T60
5946 | Waveform Shape Setting | unitless | Consult detailed help for more information on this variable | String | Settings | Common
5961 | PEEP/CPAP Alarm | unitless |  | Boolean | Alarms | Common
6153 | ET Tube Blocked Alarm | unitless |  | Boolean | Alarms | Common
6164 | General Alarm Code | unitless | This is a multiset variable containing the alarm codes for all active alarms. Associated with variable 2590 for the matching alarm text messages. | Byte (Multiset) | Alarms | Common
Consult detailed help for more information on this variable
6296 | O2 Supply Alarm | unitless |  | Boolean | Alarms | Monnal T60
6376 | High Leak Alarm | unitless |  | Boolean | Alarms | Monnal T60
6379 | Ratio of Inspiration Time to Total Breath Cycle | - | This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. Possible units: % | Double | VitalSigns | Common
6891 | Expiratory Flow Sensor Failure Alarm | unitless |  | Boolean | Alarms | Common
7033 | Patient Demand High Alarm | unitless |  | Boolean | Alarms | Common
7038 | High Inspired Minute Volume Alarm | unitless |  | Boolean | Alarms | Common
7039 | Low Inspired Minute Volume Alarm | unitless |  | Boolean | Alarms | Common
8012 | Plateau Airway Pressure Alarm | unitless |  | Boolean | Alarms | Monnal T60
8301 | Nebulizer Error | unitless |  | Boolean | Alarms | Monnal T60
8379 | Peak Pressure | - | This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. Possible units: cmH2O, mbar, bar, hPa, kPa, mmHg | Double | VitalSigns | Monnal T60
8579 | Plimit Reached Alarm | unitless |  | Boolean | Alarms | Monnal T60
8589 | High Expired Tidal Volume Alarm | unitless |  | Boolean | Alarms | Common
8596 | Low FiO2 Alarm | unitless |  | Boolean | Alarms | Common
8597 | High FiO2 Alarm | unitless |  | Boolean | Alarms | Common
8598 | Low Respiration Rate Alarm | unitless |  | Boolean | Alarms | Common
8599 | High Respiration Rate Alarm | unitless |  | Boolean | Alarms | Common
8600 | CO2 Apnea Alarm | unitless |  | Boolean | Alarms | Monnal T60
8606 | Battery Failure Alarm | unitless |  | Boolean | Alarms | Common
8614 | Ventilator Failure Alarm | unitless |  | Boolean | Alarms | Common
9293 | Pressure Rise Time Setting | cmH2O/s |  | Double | Settings | Common
9485 | Maximum Inspiratory Time Setting | - | Possible units: ms, sec, min, h | Double | Settings | Common
9584 | High Expired Minute Volume Alarm | unitless |  | Boolean | Alarms | Common
9836 | CO2 Sensor Failure Alarm | unitless |  | Boolean | Alarms | Monnal T60

## Xenios Console
ID | Label | Unit | Description | Type | Group | Device | Source
1033 | Air Bubble Detection State | unitless | Consult detailed help for more information on this variable | Byte | DeviceInfo | Common
1869 | Module State | unitless | This variable is 16 bits long. | UShort | DeviceInfo | Common
Consult detailed help for more information on this variable
1891 | Device Power Supply Status | unitless | This variable is 24 bits long. | UShort | DeviceInfo | Common
Consult detailed help for more information on this variable
2132 | AC Power Supply Failure Alarm | unitless | This variable is 32 bits long. | ULong | Alarms | Common
Consult detailed help for more information on this variable
2590 | Generic Device Alarm | unitless | This variable is 40 bits long. | ULong | Alarms | Common
Consult detailed help for more information on this variable
2899 | Flow Sensor Type | unitless | Consult detailed help for more information on this variable | UByte | DeviceInfo | Common
2907 | Patient Temperature | °C | This variable is associated with LowerAlarmLimit and UpperAlarmLimit attributes. This variable is associated with ChannelID identifier. | Double | VitalSigns | Common
3052 | Device Screen State | unitless | Possible values : | UByte | DeviceInfo | Common
0 : GUI Unlocked
1 : GUI Locked
3267 | Pulsatile Pulse Percentage Setting | % | This variable is for pulsatile mode. | UByte | Settings | Common
3288 | Generic Pump Rotation Speed | rpm | This variable is associated with LowerAlarmLimit and UpperAlarmLimit attributes. | Short | VitalSigns | Common
3335 | Pulsatile Simulated Frequency Setting | 1/min | This variable is for pulsatile mode. | UByte | Settings | Common
3636 | Device Alarm Status | unitless | This variable is 8 bits long. | UByte | Alarms | Common
Consult detailed help for more information on this variable
3691 | HLM System Control Module Technical Alarm | unitless | This variable is 24 bits long. | ULong | Alarms | Common
Consult detailed help for more information on this variable
3696 | Generic Pump Running Status | unitless | This variable is 32 bits long. | ULong | DeviceInfo | Common
Consult detailed help for more information on this variable
3701 | Generic Pump Flow Rate Setting | mL/min |  | Short | Settings | Common
3702 | Generic Pump Rotation Speed Setting | rpm |  | Short | Settings | Common
3705 | Generic Pump Flow Alarm | unitless | This variable is 16 bits long. | UShort | Alarms | Common
Consult detailed help for more information on this variable
5233 | Blood Level Measurement Status | unitless | Consult detailed help for more information on this variable | Byte | OtherMeasuredData | Common
5855 | Blood Flow Rate | mL/min | This variable is associated with LowerAlarmLimit and UpperAlarmLimit attributes. | Short | VitalSigns | Common
7244 | Generic Pressure Line Value | mmHg | This variable is associated with LowerAlarmLimit and UpperAlarmLimit attributes. This variable is associated with ChannelID identifier. | Short | VitalSigns | Common
7945 | Pulsatile Pump Speed Amplitude Setting | rpm | This variable is for pulsatile mode. | Short | Settings | Common
9266 | Generic ECG Alarm | unitless | This variable is 8 bits long. | UByte | Alarms | Common
Consult detailed help for more information on this variable