## Sheet1
| Label | Description
1 | Heart Rate | The MeasurementTime of BP variables are only updated when a new measurement is performed. In between measurements, the last valid values are sent repeatedly.
2 |  | This variable is associated with SourceID identifier.
4 | NIBP Systolic | The MeasurementTime of BP variables are only updated when a new measurement is performed. In between measurements, the last valid values are sent repeatedly.
5 | NIBP Diastolic | The MeasurementTime of BP variables are only updated when a new measurement is performed. In between measurements, the last valid values are sent repeatedly.
6 | Dialysate Inlet Pressure
7 | Dialysate Outlet Pressure
8 | Stroke Volume
9 | Patient Weight
10 | Therapy Running Status | Possible values:
11 |  | 0: No treatment running
12 |  | 32: Treatment running
14 | Blood Flow Rate Setting
15 | Fluid Removal Volume Setting
16 | Transmembrane Pressure
17 | Fluid Removal Actual Rate
18 | Relative Blood Volume Change
19 | Generic Venous Line Pressure
20 | Equivalent Dialysate Sodium (Na)
21 | Desired Equivalent Dialysate Sodium (Na)
22 | Venous Pump Flow
23 | Arterial/Main Pump Flow
24 | Replacement Cumulated Volume
25 | Fluid Removal Cumulated Volume
26 | Transmembrane Pressure Setting
27 | Desired Blood Volume Change Percentage
28 | Plasma Sodium (Na)
29 | Dilution Ratio (Qf/Qb)
30 | Therapy Remaining Time
31 | Anticoagulant Cumulated Volume
32 | Delivered Treatment Effectiveness (Kt/V)
33 | Blood Flow Rate
34 | Dialysate Flow Rate
35 | Effective Blood Volume Processed (Kt)
36 | Dialysate Temperature
37 | Replacement Cumulated Bolus Volume
38 | Dialysate pH
39 | Ionic Clearance | Ionic Dialysance
40 | Ionic Mass Balance
41 | Bicarbonate Conductivity (8.4%)
42 | Actual Anticoagulant Rate
43 | Substitution Fluid And Blood Ratio (Qi/Qb)
44 | Generic Arterial Line Pressure
45 | Dialysate Conductivity
46 | Plasma Conductivity
47 | Therapy Elapsed Time
48 | Processed Cumulated Blood Volume
49 | Therapy Operation Status | Consult detailed help for more information on this variable
50 | Replacement Flow Rate
51 | Heart Rate
52 | Cardiac Output | This variable is available only if patient category 'adult' or 'paediatric' is selected.
53 | Patient Category | A: adult C: paediatric L: linear
54 | Cardiac Index | This variable is available only if patient category 'adult' or 'paediatric' is selected.
55 | Stroke Volume | This variable is available only if patient category 'adult' or 'paediatric' is selected.
56 | Stroke Volume Index | This variable is available only if patient category 'adult' or 'paediatric' is selected.
57 | Device Date/Time
58 | Blood Minute Distance
59 | Blood Stroke Distance
60 | Flow Time To Peak
61 | Blood Mean Acceleration
62 | Corrected Flow Time
63 | Peak Velocity
65 | End Tidal CO2 (EtCO2) | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Evita V300, Fabius Family. This variable is associated with UpperAlarmLimit attribute, which is available for Perseus A500, Atlan, Evita Infinity V500, Babylog VN500, Savina 300, Primus, Primus IE, Apollo, Zeus IE, Oxylog 3000 plus, Evita V300. This variable is associated with LowerAlarmLimit attribute, which is available for Perseus A500, Atlan, Evita Infinity V500, Babylog VN500, Savina 300, Primus, Primus IE, Apollo, Zeus IE, Oxylog 3000 plus, Evita V300. Possible units: mmHg, %, kPa. This variable is in Multiset. See variables 2355 and 6551 for concentration and partial pressure respectively. For Evita Infinity V500, Babylog VN500: This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. end tidal CO2 etCO2 Note: etCO2 values and alarm limits are transmitted according to the configuration on the device screen (mmHg or Vol%). If kPa is configured be aware that the Medibus alarm limits will be transmitted with mmHg.
66 | Tidal Volume | For Evita 4, Evita 2 dura, Evita XL: Expiratory tidal volume. This variable is associated with UpperAlarmLimit attribute, which is also sent in separate variable 2763. Values below 1 L will be transmitted in units of mL, otherwise in units of L. Device firmware version 04.00 and higher: Limit is not transmitted, if alarm is switched off. For Savina: This variable is associated with UpperAlarmLimit attribute.Device SW version 02.00 and higher: Limit is not transmitted, if alarm is switched off. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300, Fabius Family. This variable is associated with UpperAlarmLimit attribute, which is available for Perseus A500, Evita Infinity V500, Savina 300, Evita V300. This variable is associated with LowerAlarmLimit attribute, which is available for Evita Infinity V500, Evita V300. For Evita Infinity V500, Babylog VN500: This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. For Zeus, Zeus IE, Primus, Primus IE, Apollo, Pallas: Exp. Tidal Volume.
67 | Dynamic Compliance | For Evita 4, Evita 2 dura, Evita XL: Values below 100 will be transmitted with one decimal place, otherwise without a decimal place. For Savina: Available since device SW version 02.00. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Evita V300.
68 | CO2 Concentration Waveform | For Evita 4, Evita 2 dura, Evita XL: Possible units: mmHg, kPa, and % For Evita 2: Possible units: mmHg, kPa, and % Firmware version 2.00 or higher. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Evita V300. Possible units: mmHg, %, kPa. Unit "mmHg" is available for Perseus A500, Atlan, Evita Infinity V500, Babylog VN500, Savina 300, Primus, Primus IE, Apollo, Zeus IE, Oxylog 3000 plus, Evita V300. Unit "kPa" and "%" are available for Perseus A500, Atlan, Savina 300, Primus, Primus IE, Apollo, Zeus IE, Oxylog 3000 plus. See variables 8866 and 7872 for concentration and partial pressure respectively. For Evita Infinity V500, Babylog VN500: Possible units: mmHg, %. Expiratory CO2 concentration. For Zeus, Zeus IE, Primus, Primus IE, Apollo, Pallas, Vamos, Vamos plus: Possible units: mmHg, %, kPa.
69 | Airway Flow Wave | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300, Fabius Family. For Evita Infinity V500, Babylog VN500, Carina, Zeus, Zeus IE, Primus, Primus IE, Apollo, Pallas: Flow (insp./exp.)
70 | Alarm Data | This variable is same as 3878 but sent as a string, which is an array containing all alarm messages and separated by delimiter "¦¦". This variable, along with variables 588 and 4131, are marked as obsolete because they use either"¦¦" or "," in the array string as delimiter, which are often not compatible with customers' applications. Variables 6164, 3878, and 2596 should be used instead. The HL7 output can contain an '&' character to indicate a LOW alarm. In order to avoid issues, the '&' character will be replaced by a '?' character in the DDI output.
71 | Resistance | For Evita 4, Evita 2 dura, Evita XL: Values below 100 will be transmitted with one decimal place, otherwise without a decimal place. For Savina: Available since device SW version 02.00. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Zeus IE, Evita V300.
72 | Ventilation Mode | The ventilation mode is reported in variable 584 and 2756. Variable 2756 contains mode ID numeric value, while variable 584 contains string representation which is the complete ventilation mode information concatenated from basic ventilation mode together with available options such as "Mode VC-CMV /AF". See variable 2756 for full list of supported ventilation mode.
73 | Alarm Message | This variable is same as 6164 but sent as a string, which is an array containing all alarm codes and separated by delimiter ",". This variable, along with variables 177 and 4131, are marked as obsolete because they use either"¦¦" or "," in the array string as delimiter, which are often not compatible with customers' applications. Variables 6164, 3878, and 2596 should be used instead.
74 | Inspired O2 (FiO2) | For Savina: This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. Available since device SW version 03.00. Limits are transmitted in LPO-Mode (Option) only. For Fabius GS, Fabius GS premium, Fabius Tiro, Fabius plus, Fabius plus XL, Fabius OS, Fabius MRI: This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Evita V300, Fabius Family. This variable is associated with UpperAlarmLimit attribute, which is available for Perseus A500, Atlan, Primus, Primus IE, Apollo, Zeus IE, Fabius Family. This variable is associated with LowerAlarmLimit attribute, which is available for Perseus A500, Atlan, Primus, Primus IE, Apollo, Zeus IE, Fabius Family. Inspiratory oxygen fraction, FiO2. For Zeus, Zeus IE: This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. For Primus, Primus IE, Apollo, Pallas: This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes.
75 | Device Name | This variable contains the device name defined by the manufacture. See variable 3591 for device model identification number.
76 | Inspiratory Component of I:E Ratio | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: Obsolete variable that may be removed in future release. Partial information of I:E ratio. For Evita Infinity V500, Savina: Partial information. Please use variable 1190 instead. Obsolete variable that may be removed in future release.
77 | Expiratory Component of I:E Ratio | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: Obsolete variable that may be removed in future release. Partial information of I:E ratio. For Evita Infinity V500, Savina: Partial information. Please use variable 1190 instead. Obsolete variable that may be removed in future release.
78 | Positive End Expiratory Pressure (PEEP) | For Fabius GS, Fabius GS premium, Fabius Tiro, Fabius plus, Fabius plus XL, Fabius OS, Fabius MRI: 1mbar = 1cmH2O. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300, Fabius Family.
79 | I:E Ratio | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Evita V300. Format: 1:x if E > I, x:1 if I > E, or 1:1 if I = E. The useless leading or ending “0” or “.” is not transmitted, and the maximum digits after the floating point is 2. For Evita Infinity V500, Babylog VN500: Format: 1:x if E > I, x:1 if I > E, or 1:1 if I = E. The useless leading or ending “0” or “.” is not transmitted, and the maximum digits after the floating point is 2.
80 | Message Text | This variable, along with variable 3909 are marked as obsolete because they use either"¦¦" or "," in the array string as delimiter, which are often not compatible with customers' applications. Variables 8366 and 8696 should be used instead.
81 | Tidal Volume | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: Obsolete variable that may be removed in future release. This variable is marked as deprecated in order to synchronize with other models. Please use variable 60 instead. For Evita Infinity V500, Babylog VN500: Deprecated in order to synchronize with other models. Please use variable 60 instead.
82 | Minute Volume | For Evita 4, Evita 2 dura, Evita XL: This variable is associated with UpperAlarmLimit attribute, which is also sent in separate variable 1400. This variable is associated with LowerAlarmLimit attribute, which is also sent in separate variable 1399. For device firmware version 04.00 and higher: Lower alarm limit is not transmitted, if alarm is switched off. Values may be transmitted with one or two decimal place, depended on firmware version. Note that the value sent by the device is less precise than what is displayed on device screen. For example, a displayed value of 5.12 is sent by the device and thus by DDI as 5.1. For Evita 2, Evita, Fabius GS, Fabius GS premium, Fabius Tiro, Fabius plus, Fabius plus XL, Fabius OS, Fabius MRI, Evita Infinity V500, Babylog VN500, Carina: This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. This variable is available only with a precision to the tenth. For example: for a minute volume value of 7.43, the device displays 7.43 L/min but the DDI will output it as 7.4 L/min. For Savina: This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300, Fabius Family. This variable is associated with UpperAlarmLimit attribute, which is available for Perseus A500, Atlan, Savina 300, Primus, Primus IE, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Fabius Family. This variable is associated with LowerAlarmLimit attribute, which is available for Perseus A500, Atlan, Savina 300, Primus, Primus IE, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Fabius Family. This variable is available only with a precision to the tenth. For example: for a minute volume value of 7.43, the device displays 7.43 L/min but the DDI will output it as 7.4 L/min. For Zeus, Zeus IE, Primus, Primus IE, Apollo, Pallas: This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. Respiratory Minute Volume This variable is available only with a precision to the tenth. For example: for a minute volume value of 7.43, the device displays 7.43 L/min but the DDI will output it as 7.4 L/min. For Evita Infinity V500 and Babylog VN500: This variable is only output if the device version is less than 2.30, otherwise variable 2326 is exported.
83 | Spontaneous Respiration Rate | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Zeus IE, Oxylog 3000 plus, Evita V300. This variable is associated with UpperAlarmLimit attribute, which is available for Oxylog 3000 plus.
84 | Static Compliance | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500 Evita Infinity V600, Evita Infinity V800. For Evita Infinity V500: Cstat (Low Flow) Cstat.
85 | Mandatory Respiration Rate Setting | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300, Fabius Family. Respiratory rate, RR. For Evita Infinity V500, Babylog VN500: Respiratory Rate RR.
86 | Tidal Volume Setting | For Evita 4, Evita 2 dura, Evita XL, Evita 2, Evita, Savina, Fabius GS, Fabius GS premium, Fabius Tiro, Fabius plus, Fabius plus XL, Fabius OS, Fabius MRI, Carina, Zeus, Zeus IE, Primus, Primus IE, Apollo, Pallas: Inspired tidal volume setting. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: Obsolete variable that may be removed in future release. Variable is in Multiset format. Please use separate variable 2771 and 5044 instead. For Evita Infinity V500, Babylog VN500: Variable 2771 and 5044 contains same information but in different units. This variable combines these two variables and sends whichever is significant. For Zeus, Zeus IE: Obsolete variable that may be removed in future release. Multiset variable. Please use variables 2771 and 5044 instead.
87 | Sigh Interval Setting | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. Time interval between sighs, Interval sigh.
88 | Sigh Number Setting | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. Number of sigh cycles, Cycles sigh.
89 | Apnea Tidal Volume Setting | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Oxylog 3000 plus, Carina, Evita V300. Tidal volume for apnea ventilation, Vtapn. Possible units: L, mL. Unit "L" is available for Evita Infinity V500, Babylog VN500, Savina 300, Oxylog 3000 plus, Carina, Evita V300. Unit "mL" is available for Evita Infinity V500, Babylog VN500, Evita V300. For Evita Infinity V500, Babylog VN500: Possible units: L, mL. VT Apnea VTapn.
90 | Apnea Respiration Rate Setting | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Oxylog 3000 plus, Carina, Evita V300. Respiratory rate for apnea ventilation, Rrapn.
91 | Pressure Support Level Above PEEP Setting | For Evita 4, Evita 2 dura, Evita XL: This variable is calculated from variables 1392 and 2104. Available only with Evita XL. For Medibus.X compatible device, Perseus A500, Atlan, Evita V300, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300, Fabius Family. Pressure amplitude above PEEP in Pressure Support, DPsupp. For Evita Infinity V500, Babylog VN500: ΔPsupp. For Carina: Assisted spontaneous breathing.
92 | Inspiratory Flow Setting | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300, Fabius Family.
93 | Inspiratory Flow Trigger Setting | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Evita V300, Fabius Family. Flow trigger threshold, Flow trigger. For Primus, Primus IE, Apollo, Pallas: Flow trigger.
94 | Intermittent PEEP Setting | For Fabius GS, Fabius GS premium, Fabius Tiro, Fabius plus, Fabius plus XL, Fabius OS, Fabius MRI: 1mbar = 1cmH2O. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Evita V300. Additional intermittent PEEP for sighs, PEEPsigh. For Evita Infinity V500, Babylog VN500: Interm. PEEP ΔintPEEP. For Zeus, Zeus IE: PEEP for all vent mode different from Man/Spont.
95 | APRV Low Pressure Setting | For Evita: BIPAP low pressure For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Evita V300. Lower pressure level during APRV, Plow. For Evita Infinity V500, Babylog VN500: APRV Low Pressure Plow.
96 | APRV Low Time Setting | For Evita: BIPAP low time For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Evita V300. Time of lower pressure level during APRV, Tlow. For Evita Infinity V500, Babylog VN500: APRV Low Time Tlow.
97 | APRV High Pressure Setting | For Evita: BIPAP high pressure For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Evita V300. Upper pressure level during APRV, Phigh. For Evita Infinity V500, Babylog VN500: APRV High Pressure Phigh.
98 | APRV High Time Setting | For Evita: BIPAP high time For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Evita V300. Time of upper pressure level during APRV, Thigh. For Evita Infinity V500, Babylog VN500: APRV High Time Thigh.
99 | Apnea Time Setting | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Oxylog 3000 plus, Carina, Evita V300. For Evita Infinity V500, Babylog VN500: Apnea Time Tapn. For Carina: Apnea alarm time.
100 | Maximum Inspired Pressure Setting | For Fabius GS, Fabius GS premium, Fabius Tiro, Fabius plus, Fabius plus XL, Fabius OS, Fabius MRI: 1mbar = 1cmH2O. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300, Fabius Family. Pressure limitation, Pmax. For Evita Infinity V500, Babylog VN500: Max Pressure Pmax. For Carina: Pressure limitation. For Primus, Primus IE, Apollo, Pallas: Max. insp. Airway pressure.
101 | Mean Airway Pressure | For Fabius GS, Fabius GS premium, Fabius Tiro, Fabius plus, Fabius plus XL, Fabius OS, Fabius MRI: This variable is associated with LowerAlarmLimit attribute. 1mbar = 1cmH2O. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300, Fabius Family. For Zeus, Zeus IE, Primus, Primus IE, Apollo, Pallas: Mean Breathing Pressure.
102 | Rapid Shallow Breathing Index | For Evita 4, Evita 2 dura, Evita XL: Firmware version 04.00 or higher For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Atlan, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Evita V300. Possible units: 1/L * min, 1/mL*min. Unit "1/L * min" is available for Evita Infinity V500, Babylog VN500, Savina 300, Evita V300. Unit "1/mL*min" is available for Evita Infinity V500, Babylog VN500, Evita V300 (added with SW 2.30 for Babylog VN500, added with SW 2.30 for Evita Infinity V500).
103 | Intrinsic PEEP | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300.
104 | Minimum Airway Pressure | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300.
105 | P0.1 Occlusion Pressure | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. Occlusion pressure in P0.1 maneuver. For Evita Infinity V500, Babylog VN500: Only available for Evita Infinity V500
106 | Lung Volume Trapped (TrpVol) | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. Volume trapped in the lung by intrinsic PEEP, Vtrap. For Evita Infinity V500: Trapped Volume Vtrap.
107 | Respired Volume Wave | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Evita V300. Respiratory volume since start of inspiration, Volume. For Evita Infinity V500, Babylog VN500: Volume since inspiration began.
108 | Ventilation Pressure Variability Setting | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800. For Evita Infinity V500, Babylog VN500: Supported through Medibus.X protocol.
109 | APRV Low Time Limitation | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. Maximum time of low pressure level during APRV/AutoRelease, Tlow max. For Evita Infinity V500, Babylog VN500: Tlow max.
110 | Spontaneous Mean Tidal Volume | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500 Evita Infinity V600, Evita Infinity V800. Spontaneous mean tidal volume, VTspon mean. For Evita Infinity V500, Babylog VN500: VTspon mean (Not available on Babylog VN500). Only available for Evita Infinity V500
111 | Spontaneous Inspired Mean Tidal Volume | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500 Evita Infinity V600, Evita Infinity V800. For Evita Infinity V500, Babylog VN500: VTispon mean (Not available on Babylog VN500). Only available for Evita Infinity V500
112 | Spontaneous Expired Mean Tidal Volume | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800. For Evita Infinity V500, Babylog VN500: VTespon mean (Not available on Babylog VN500). Only available for Evita Infinity V500
113 | SmartCare FiO2 Limitation | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. FiO2 in SmartCare - upper limit, SC-FiO2high For Evita Infinity V500, Babylog VN500: SC-Upper limit inspiratory O2 FiO2 max.
114 | SmartCare PEEP Limitation | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. PEEP in SmartCare - upper limit, SC-PEEPmax. For Evita Infinity V500, Babylog VN500: SC-Upper limit PEEP, PEEP max.
115 | SmartCare Spontaneous Respiration Rate Lower Limitation | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. Spontaneous respiratory rate in SmartCare - lower limit, SC-RRspon low. For Evita Infinity V500, Babylog VN500: SC-Lower limit respiratory rate in the guideline, RRspon low.
116 | SmartCare Spontaneous Respiration Rate Upper Limitation | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. Spontaneous respiratory rate in SmartCare - upper limit, SC-RRspon high. For Evita Infinity V500, Babylog VN500: SC-Upper limit respiratory rate in the guideline, RRspon high.
117 | SmartCare Tidal Volume Limitation | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. Tidal volume in SmartCare - lower limit, SC-VTlow. For Evita Infinity V500, Babylog VN500: SC-Lower limit tidal volume, VT low.
118 | SmartCare EtCO2 Limitation | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. End-tidal CO2 in SmartCare - upper limit, SC-etCO2high. For Evita Infinity V500, Babylog VN500: SC-Upper limit endtidal CO2 in the guideline etCO2 high.
119 | SmartCare Patient Height | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. Patient height in SmartCare, SC-Ht. For Evita Infinity V500: SC-Body height.
120 | Airway Resistance | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300.
121 | P01 Repetition Interval Setting | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. P0.1 repetition interval.
122 | Inspired Minute Volume | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Carina, Evita V300. This variable is associated with UpperAlarmLimit attribute, which is available for Carina. This variable is associated with LowerAlarmLimit attribute, which is available for Carina.
123 | Inspiratory Time Setting | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300, Fabius Family.
124 | Positive End Expiratory Pressure (PEEP) Setting | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300, Fabius Family. For Evita Infinity V500, Babylog VN500: PEEP / CPAP PEEP. For Zeus, Zeus IE: PEEP (CPAP). PEEP for mode Man/Spont.
125 | Inspired O2 (FiO2) Setting | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Zeus IE, Oxylog 3000 plus, Carina, Evita V300. For Primus, Primus IE, Apollo, Pallas: Not in device Apollo and Pallas
126 | Manual Breath Ventilation Pressure Setting | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300(added with SW 2.30 for Babylog VN500, added with SW 2.30 for Evita Infinity V500). Manual inspiratory pressure for Neonates during NIV - Non-Invasive Ventilation, PmanInsp.
127 | Manual Breath Maximum Time Limit | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300(added with SW 2.30 for Babylog VN500, added with SW 2.30 for Evita Infinity V500). Maximum time of pressure for manual inspiration pressure for Neonates during NIV - Non-Invasive Ventilation, TmanInsp.
128 | CO2 Tidal Production | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. CO2 production volume per breath, VTCO2. For Evita Infinity V500, Babylog VN500: Supported through Medibus.X protocol.
129 | CO2 Minute Production | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Evita V300. CO2 production, V'CO2.
130 | Device Date/Time | Format: YYYYMMDDhhmmss, where ss is always present but might be 00.
131 | Inspired Tidal Volume | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Carina, Evita V300. This variable is associated with UpperAlarmLimit attribute, which is available for Evita Infinity V500, Zeus IE, Evita V300. This variable is associated with LowerAlarmLimit attribute, which is available for Evita Infinity V500, Evita V300. For Primus, Primus IE, Apollo, Pallas: Insp. Tidal Volume.
132 | Expired Tidal Volume | For Evita XL: Neo expired tidal volume. Available only with firmware version 7.01 or higher. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Evita V300.
133 | Expired Minute Volume | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. This variable is associated with UpperAlarmLimit attribute, which is available for Evita Infinity V500, Babylog VN500, Evita V300. This variable is associated with LowerAlarmLimit attribute, which is available for Evita Infinity V500, Babylog VN500, Evita V300. For Evita Infinity V500, Babylog VN500: This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes.
134 | End Tidal CO2 Percentage | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Evita V300, Fabius Family. This variable is associated with UpperAlarmLimit attribute, which is available for Perseus A500, Atlan, Evita Infinity V500, Babylog VN500, Savina 300, Primus, Primus IE, Apollo, Zeus IE, Oxylog 3000 plus, Evita V300. This variable is associated with LowerAlarmLimit attribute, which is available for Perseus A500, Atlan, Evita Infinity V500, Babylog VN500, Savina 300, Primus, Primus IE, Apollo, Zeus IE, Oxylog 3000 plus, Evita V300. end tidal CO2 etCO2
135 | Fresh Gas O2 Flow Rate | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Evita V300. O2 therapy flow, Flow. For Evita Infinity V500, Babylog VN500: Flow O2 Therapy Flow.
136 | Ventilator Operation Mode | This variable contains ventilation operation mode in string format, which is also send as a numerical value in variable 2758. See detailed description of variable 2758 for the list of supported values. For Evita 4, Evita 2 dura, Evita XL: Firmware version 04.00 or higher For Evita 2: Variable label is not correct. Please use variable 3853 instead. Obsolete variable that may be removed in future release.
137 | Spontaneous Inspired Tidal Volume | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. For Zeus, Zeus IE: This variable is associated with UpperAlarmLimit attribute. Insp. Tidal volume.
138 | Airway Dead Space | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. Serial dead space volume, Vds. For Evita Infinity V500, Babylog VN500: Serial dead space Vds.
139 | Dynamic Compliance Over Last 20% Breath | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. Ratio of compliance during the last 20 % of inspiration over dynamic compliance, C20/Cdyn. For Evita Infinity V500, Babylog VN500: C20/Cdyn.
140 | Expired Spontaneous Tidal Volume | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300.
141 | Mechanical Expired Minute Volume | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Primus, Primus Infinity Empowered, Apollo, Evita V300. Mandatory expiratory minute volume, MVemand
142 | Spontaneous Expired Minute Volume | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Primus, Primus Infinity Empowered, Apollo, Evita V300.
143 | Patient Type ID | Consult detailed help for more information on this variable
144 | Airway Pressure Wave | For Fabius GS, Fabius GS premium, Fabius Tiro, Fabius plus, Fabius plus XL, Fabius OS, Fabius MRI: 1mbar = 1cmH2O. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300, Fabius Family. For Evita Infinity V500, Babylog VN500: Airway pressure Paw.
145 | Peak Airway Pressure | For Evita 4, Evita 2 dura, Evita XL: This variable is associated with UpperAlarmLimit attribute, which is also sent in separate variable 7040. For Evita 2, Evita, Savina: This variable is associated with UpperAlarmLimit attribute. For Fabius GS, Fabius GS premium, Fabius Tiro, Fabius plus, Fabius plus XL, Fabius OS, Fabius MRI: 1mbar = 1cmH2O. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300, Fabius Family. This variable is associated with UpperAlarmLimit attribute, which is available for Perseus A500, Atlan, Evita Infinity V500, Babylog VN500, Savina 300, Zeus IE, Oxylog 3000 plus, Carina, Evita V300, Fabius Family. This variable is associated with LowerAlarmLimit attribute, which is available for Zeus IE, Fabius Family. Peak inspiratory pressure, PIP. For Evita Infinity V500, Babylog VN500, Carina: This variable is associated with UpperAlarmLimit attribute. For Zeus, Zeus IE: This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. Peak Breathing Press. For Primus, Primus IE, Apollo, Pallas: Peak Breathing Pressure.
146 | Plateau Airway Pressure | For Fabius GS, Fabius GS premium, Fabius Tiro, Fabius plus, Fabius plus XL, Fabius OS, Fabius MRI: 1mbar = 1cmH2O. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Evita V300, Fabius Family. For Evita Infinity V500, Babylog VN500: Only available for Evita Infinity V500
147 | Spontaneous Tidal Volume | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Zeus IE, Evita V300. Possible units: L, mL. For Evita Infinity V500, Babylog VN500: Possible units: L, mL.
148 | I:E Ratio Setting | String format: 1:x if E > I, x:1 if I > E, or 1:1 if I = E. The useless leading or ending “0” or “.” is not transmitted, and the maximum digits after the floating point is 2. For Carina, Primus, Primus IE, Apollo, Pallas, Perseus A500: This variable is calculated by the DDI from respiration rate and inspiratory time settings. For Medibus.X compatible device: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300, Fabius Family. For Primus, Primus IE, Apollo, Perseus A500, Carina and Oxylog 3000 plus, it is calculated by the DDI from respiration rate and inspiratory time settings. For Atlan, Evita Infinity V500, Babylog VN500, Savina 300, Zeus IE, Evita V300, and Fabius Family, it is calculated by the DDI from inspiratory and expiratory components. For Evita, Evita 2, Evita 4, Evita 2 dura, and Evita XL, Fabius GS, Fabius GS premium, Fabius Tiro, Fabius plus, Fabius plus XL, Fabius OS, Fabius MRI, Evita Infinity V500, Babylog VN500, Zeus IE, Evita V300: This variable is calculated by the DDI from inspiratory and expiratory components of I:E ratio sent by the devices.
149 | Alarm Severity Level | This variable is in multiset format. This variable contains alarm priority level, variable 6164 contains alarm codes, and variable 3878 contains alarm messages. They need to be interpreted together. One alarm code may have one or more priority values and one or more alarm messages associated with it. The possible alarms are depended on device firmware version. Please see device user manuals for details. See variables 177 and 588 and 4131 for string format.
150 | Ventilator General Status | Ventilator Standby mode status.
152 |  | Consult detailed help for more information on this variable
153 | Ventilation Mode ID | This variable contains ID of basic ventilation mode ID without options. See also variable 584.
155 |  | Consult detailed help for more information on this variable
156 | Ventilator Operation Mode ID | For Evita 4, Evita 2 dura, Evita XL: Firmware version 04.00 or higher For Evita 2: Variable label is not correct. Please use variable 2467 instead. Obsolete variable that may be removed in future release.
158 |  | Consult detailed help for more information on this variable
159 | Tidal Volume Setting | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300, Fabius Family. Inspiratory tidal volume, Vti. For Evita Infinity V500, Babylog VN500: See also variable 1321 and 5044. For Zeus, Zeus IE: This variable contains same information as variable 1321, except that the value of different units (L or mL) are sent in separate variables. See also variable 5044.
160 | ASB Ramp Time Setting | For Evita 2: Firmware version 2.00 or higher. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300. Pressure rise time, Slope. For Carina: ASB ramp. For Primus, Primus IE, Apollo, Pallas: Slope time (Ramp).
161 | PS Activation State | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300, Fabius Family. A value of 1 indicates that Pressure Support (PS) is added to the current ventilation mode. A value of 0 indicates that Pressure Support (PS) is not added to the current ventilation mode. Text message displayed on device screen: "/PS"
162 | VS Activation State | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. A value of 1 indicates that Volume support (VS) is added to the current ventilation mode. A value of 0 indicates that Volume support (VS) is not added to the current ventilation mode. Text messages displayed on device screen: "/VS"
163 | APRV High Pressure (Phigh) | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. Upper pressure level during APRV, Phigh.
164 | APRV Low Pressure (Plow) | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. Lower pressure level during APRV, Plow.
165 | Inspired Pressure Setting | For Evita 4, Evita 2 dura, Evita XL: Available only with EvitaXL firmware version 7.00 or higher. For Fabius GS, Fabius GS premium, Fabius Tiro, Fabius plus, Fabius plus XL, Fabius OS, Fabius MRI: 1mbar = 1cmH2O. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300, Fabius Family.
166 | Device Average Gas Flow | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. Average device flow, FlowDev.
167 | Auto Flow Active Status | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300. A value of 1 indicates that AutoFlow (AF) is added to the current ventilation mode. A value of 0 indicates that AutoFlow (AF) is not added to the current ventilation mode. Text message displayed on device screen: "/AF"
168 | ATC Activation Status | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Evita V300. A value of 1 indicates that automatic tube compensation (ATC) is added to the current ventilation mode. A value of 0 indicates that automatic tube compensation (ATC) is not added to the current ventilation mode. Text message displayed on device screen: "/ATC" For Savina 300 (MEDIBUS.X): added with SW 5.00
169 | Device Number | Consult detailed help for more information on this variable
170 | Device Communications Status | Value "stop" sent when a request to stop the communication is received from the device
171 | Patient Type | This variable contains patient type in string format, which is also send as a numerical value in variable 2467. See detailed description of variable 2467 for the list of supported values.
172 | Generic Alarm Message | This variable is in multiset format. This variable contains alarm messages, variable 2596 contains alarm priority level, variable 6164 contains alarm codes. They need to be interpreted together. One alarm code may have one or more priority values and one or more alarm messages associated with it. The possible alarms are dependent on device firmware version. Please see device user manuals for details. See variables 177 and 588 and 4131 for string format. The HL7 output can contain an '&' character to indicate a LOW alarm. In order to avoid issues, the '&' character will be replaced by a '?' character in the DDI output.
173 | Message Code | This variable, along with variable 1197 are marked as obsolete because they use either"¦¦" or "," in the array string as delimiter, which are often not compatible with customers' applications. Variables 8366 and 8696 should be used instead.
174 | Inspired Mandatory Tidal Volume | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300.
175 | Spontaneous Fraction of Minute Volume | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. Spontaneous breathing portion of minute volume, % Mvspon. For Evita Infinity V500, Babylog VN500: Spontaneous Fraction Min. Vol. %MVspon (MVspon : MVtotal).
176 | Negative Inspiratory Force | For Evita 4, Evita 2 dura, Evita XL: The value is not transmitted, if it is displayed on Evita 4 / Evita 2 dura as "[less than] –45 mbar". Firmware version 04.00 or higher For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300.
177 | Auto Release Activation Status | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. A value of 1 indicates that AutoRelease is added to the current ventilation mode. A value of 0 indicates that AutoRelease is not added to the current ventilation mode. Text messages displayed on device screen: "/AutoRelease"
178 | Generic Respiration Rate | For Evita 4, Evita 2 dura, Evita XL: This variable is associated with UpperAlarmLimit attribute, which is also sent in separate variable 1339. Total Respiratory Rate (RRtotal). For Evita 2, Evita, Savina, Carina: This variable is associated with UpperAlarmLimit attribute. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300, Fabius Family. This variable is associated with UpperAlarmLimit attribute, which is available for Evita Infinity V500, Babylog VN500, Savina 300, Carina, Evita V300. Respiratory rate, RR. For Evita Infinity V500, Babylog VN500: This variable is associated with UpperAlarmLimit attribute. Respiratory Rate (RRtotal) RR.
179 | Minute Volume Alarm Delay | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. Delay for minute volume high alarm, MVhigh delay. This variable is associated with a ChannelID identifier.
180 | Flow Assist Level Setting | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. Flow assist during PPS ventilation, Flow Assist. For Evita Infinity V500, Babylog VN500: Flow Assist.
181 | Volume Assist Level Setting | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. Volume assist during PPS ventilation, Vol. Assist. For Evita Infinity V500, Babylog VN500: Vol. Assist.
182 | Disconnection Alarm Delay Time | For Evita 4, Evita 2 dura, Evita XL: Firmware version 4.00 or higher. For Savina: Tdeconnect. Available since Device SW version 02.00. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Carina, Evita V300. Disconnection alarm time, Tdiscon. For Evita Infinity V500, Babylog VN500: Tdisconnect. For Carina: Disconnection detection timeout.
183 | Medibus Revision
184 | Device Revision | Format: mm.nn
185 | Alarm Priority | This variable is same as 2596 but sent as a string, which is an array containing all alarm priorities and separated by "," delimiter. This variable, along with variables 177 and 588, are marked as obsolete because they use either"¦¦" or "," in the array string as delimiter, which are often not compatible with customers' applications. Variables 6164, 3878, and 2596 should be used instead.
186 | VG Activation Status | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Carina, Evita V300. A value of 1 indicates that Volume Guarantee (VG) is added to the current ventilation mode. A value of 0 indicates that Volume Guarantee (VG) is not added to the current ventilation mode. Text message displayed on device screen: "/VG"
187 | Inspiratory Flow Waveform | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300.
188 | Spontaneous Inspiration Time | For Evita 4, Evita 2 dura, Evita XL: Available only with EvitaXL firmware version 7.02 or higher. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300.
189 | Lung Time Constant | For Evita 4, Evita 2 dura, Evita XL: Available only with EvitaXL firmware version 7.00 or higher. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300.
190 | Airway Correlation Coefficient | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. Correlation factor, r2. For Evita Infinity V500, Babylog VN500: r².
191 | Sigh Activation Status | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. A value of 1 indicates that the sigh function has been added to the current ventilation mode. A value of 0 indicates that the sigh function has not been added to the current ventilation mode. Text message displayed on device screen: "/Sigh"
192 | Inspired Tidal Volume Setting | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Zeus IE, Evita V300. Inspiratory tidal volume, Vti. For Evita Infinity V500, Babylog VN500: See also variable 1321 and 2771. For Zeus, Zeus IE: This variable contains same information as variable 1321, except that the value of different units (L or mL) are sent in separate variables. See also variable 2771.
193 | Dead Space to Tidal Volume Ratio | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. Ratio of serial dead space volume to expiratory tidal volume, Vds/VTe. For Evita Infinity V500, Babylog VN500: Dead space ventilation Vds/VTe.
194 | Expired Mandatory Tidal Volume | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300.
195 | Expiratory Time Constant (RCexp) | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300.
196 | EtCO2 Phase III Increment | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. Possible units: mmHg/L, kPa/L, Vol%/L. Increase of measured CO2 value in phase III of the capnogram, CO2 slope.
197 | Spontaneous I:E Ratio | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. Format: 1:x if E > I, x:1 if I > E, or 1:1 if I = E. The useless leading or ending “0” or “.” is not transmitted, and the maximum digits after the floating point is 2. For Evita Infinity V500, Babylog VN500: Format: 1:x if E > I, x:1 if I > E, or 1:1 if I = E. The useless leading or ending “0” or “.” is not transmitted, and the maximum digits after the floating point is 2.
198 | General Alarm Code | This variable is in multiset format. This variable contains alarm codes, variable 3878, alarm messages, variable 2596 contains alarm priority level. They need to be interpreted together. One alarm code may have one or more priority values and one or more alarm messages associated with it. The possible alarms are depended on device firmware version. Please see device user manuals for details. See variables 177 and 588 and 4131 for string format.
200 |  | Consult detailed help for more information on this variable
201 | Endotracheal Tube Size | For Evita 4, Evita 2 dura, Evita XL: Possible units: "0.1mm" in firmware versions 7.00 and 7.01, "mm" otherwise. Available only with EvitaXL firmware version 7.00 or higher. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Evita V300. Inner diameter of the endotracheal/tracheostoma tube. For Evita Infinity V500, Babylog VN500: ATC Tube Diameter.
202 | End Tidal CO2 Pressure | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Evita V300, Fabius Family. This variable is associated with UpperAlarmLimit attribute, which is available for Perseus A500, Atlan, Evita Infinity V500, Babylog VN500, Savina 300, Primus, Primus IE, Apollo, Zeus IE, Oxylog 3000 plus, Evita V300. This variable is associated with LowerAlarmLimit attribute, which is available for Perseus A500, Atlan, Evita Infinity V500, Babylog VN500, Savina 300, Primus, Primus IE, Apollo, Zeus IE, Oxylog 3000 plus, Evita V300. Possible units: mmHg, kPa. end tidal CO2 etCO2
203 | Expiratory Time Setting | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Zeus IE, Evita V300.
204 | Inspiratory Component of I:E Ratio Setting | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Zeus IE, Evita V300, Fabius Family, Atlan For Evita Infinity V500: Partial information. Please use variable 2583 instead. Obsolete variable that may be removed in future release.
205 | Expiratory Component of I:E Ratio Setting | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800,Savina 300, Zeus IE, Evita V300, Fabius Family, Atlan For Evita Infinity V500: Partial information. Please use variable 2583 instead. Obsolete variable that may be removed in future release.
206 | Lung Elastance | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300.
207 | Rapid Shallow Breathing Index (SBI) | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: Obsolete variable that may be removed in future release. Variable is in Multiset format. This variable is same as 1422 but DDI buffer and send all values from different data packages together.
208 | Generic Ventilation Mode | Consult detailed help for more information on this variable
209 | Ventilation Mode Primary | Consult detailed help for more information on this variable
210 | Ventilation Mode Options | Consult detailed help for more information on this variable
211 | Variable PS Activation Status | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800. A value of 1 indicates that Variable PS (VPS) has been added to the current ventilation mode. A value of 0 indicates that Variable PS (VPS) has not been added to the current ventilation mode. Text message displayed on device screen: "VariablePS"
212 | Patient Operation Mode | Possible values: 1 for Invasive, 2 for Noninvasive
213 | Artificial Airway Compensation Status | For Evita 4, Evita 2 dura, Evita XL: Tube Compensation. Available only with EvitaXL firmware version 7.00 or higher. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. Compensation rate during ATC, % comp. For Evita Infinity V500, Babylog VN500: ATC Compensation Compens.
214 | Mandatory Respiration Rate | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Zeus IE, Evita V300.
215 | Mandatory Tidal Volume | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Babylog VN500, Babylog VN600, Babylog VN800, Zeus IE, Evita V300(added with SW 2.30 for Babylog VN500, added with SW 2.30 for Evita Infinity V500). Possible units: L, mL. Unit "L" available for Evita Infinity V500, Babylog VN500, Evita V300 (added with SW 2.30 for Babylog VN500, added with SW 2.30 for Evita Infinity V500). Unit "mL" available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Zeus IE, Evita V300.
216 | CO2 Partial Pressure Waveform | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Evita V300. Possible units: mmHg, kPa. Unit "mmHg" is available for Perseus A500, Atlan, Evita Infinity V500, Babylog VN500, Savina 300, Primus, Primus IE, Apollo, Zeus IE, Oxylog 3000 plus, Evita V300. Unit "kPa" is available for Perseus A500, Atlan, Savina 300, Primus, Primus IE, Apollo, Zeus IE, Oxylog 3000 plus.
217 | Time Duration at Low Pressure (Tlow) | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. Effective expiratory time during APRV/AutoRelease, Tlow. For Evita Infinity V500, Babylog VN500: Tlow (resulting Tlow from expiratory flow cycling).
218 | Spontaneous Tidal Volume
219 | Generic Message Code | This variable is in multiset format. This variable, along with variable 8696, contains text message code and alarm text respectively. They need to be interpreted together.
220 | Generic Message Text | This variable is in multiset format. This variable, along with variable 8366, contains text message code and alarm text respectively. They need to be interpreted together.
221 | End Inspiratory Pressure | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. End-inspiratory pressure, EIP.
222 | Tracheal Pressure Waveform | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Evita V300. For Evita Infinity V500, Babylog VN500: Ptrach.
223 | Tidal Volume Indexed By Body Weight | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. Tidal volume per kg body weight, VT/Wt. The leakage is determined by the difference between the delivered flow on the inspiratory side and the measured flow on the expiratory side, which is displayed as the leakage minute volume MVleak and relative as % leak (MVleak to MV). For Evita Infinity V500, Babylog VN500: VT/kg BW.
224 | Mandatory Triggered Respiration Rate | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300(added with SW 2.42 for Evita Infinity V500, Evita V300 and Babylog VN500). Respiratory rate of triggered, mandatory breaths, RRtrig For Evita 4, Evita 2 dura, Evita XL: Available only with Evita XL firmware version 7.00 or higher.
225 | Volume of Leakage (MVLeak) | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Carina, Evita V300. Leakage minute volume, Mvleak. The leakage is determined by the difference between the delivered flow on the inspiratory side and the measured flow on the expiratory side, which is displayed as the leakage minute volume MVleak and relative as % leak (MVleak to MV). For Evita Infinity V500, Babylog VN500: Mvleak. The leakage is determined by the difference between the delivered flow on the inspiratory side and the measured flow on the expiratory side, which is displayed as the leakage minute volume MVleak and relative as % leak (MVleak to MV). For Carina: Leakage minute volume. For Evita XL: Leakage minute volume. Available since device SW version 7.02.
226 | Percent of Leakage | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Evita V300. Leakage minute volume in % of inspiratory minute volume, % leak Leakage rel. % leak - Infinity V500 determines the difference between the delivered flow on the inspiratory side and the measured flow on the expiratory side. This difference provides a measure of the amount of leakage and is displayed by Infinity V500 as the leakage minute volume MVleak and relative as % leak (MVleak to MV). For Evita Infinity V500, Babylog VN500: Leakage rel. % leak.
227 | Inspiratory Component of Spontaneous I:E Ratio | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: Obsolete variable that may be removed in future release. Partial information of I:E ratio. For Evita Infinity V500: Partial information. Please use variable 5815 instead. Obsolete variable that may be removed in future release.
228 | Expiratory Component of Spontaneous I:E Ratio | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: Obsolete variable that may be removed in future release. Partial information of I:E ratio. For Evita Infinity V500: Partial information. Please use variable 5815 instead. Obsolete variable that may be removed in future release.
229 | SmartCare Target Delta Support Pressure | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. Dpressure support goal in SmartCare, SC-DPsupp goal. For Evita Infinity V500: SC-Target ΔPsupp ΔPsupp goal.
230 | SmartCare Rated Delta Support Pressure | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. Rated Dpressure support by internal controller in SmartCare, SC-DPsupp rated. For Evita Infinity V500: SC-Rated ΔPsupp SC-Δpsupp.
231 | SmartCare Duration of Patient Session | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. Duration of session (hours and minutes) in SmartCare, SC-duration. Variable is in Multiset format (need combine the first value in hours with the second value in minutes to obtain the exact duration). Please use variable 4346 instead. For Evita Infinity V500: Multiset variable (need combine the first value in hours with the second value in minutes to obtain the exact duration).
232 | SmartCare Spontaneous Respiration Rate | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. Spontaneous respiratory rate in SmartCare, SC-Rrspon. For Evita Infinity V500: SC-Rrspon.
233 | SmartCare Tidal Volume | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. Tidal volume in SmartCare, SC-VT. For Evita Infinity V500: SC-VT.
234 | SmartCare EtCO2 | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. End-tidal CO2 concentration in SmartCare, SC-etCO2. For Evita Infinity V500: SC-etCO2.
235 | Maximum Inspiratory Time Setting | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Carina, Evita V300. Maximal inspiratory time for flow cycled strokes, Timax. For Carina: Maximal inspiration time for CPAP, ASB.
236 | Inspiratory Flow Termination Level Setting | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Savina 300, Oxylog 3000 plus, Evita V300. Inspiratory termination criterion based on peak inspiratory flow, Insp. term. For Evita Infinity V500: Inspiratory termination.
237 | Expiratory Flow Termination Setting | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. Expiratory termination criterion during APRV/AutoRelease, Exp. term. For Evita Infinity V500, Babylog VN500: Expiratory termination.
238 | SmartCare Body Weight | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. Ideal body weight in SmartCare, SC-Wt. For Evita Infinity V500: SC-Body weight.
239 | SmartCare Start of Night Rest | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. Start of night rest in SmartCare, SC-startrest. For Evita Infinity V500: SC-Start of night rest.
240 | SmartCare End of Night Rest | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. End of night rest in SmartCare, SC-endrest. For Evita Infinity V500: SC-End of night rest.
241 | Device Alarm Inactivation State
243 | End Tidal CO2 (EtCO2) | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Evita V300, Fabius Family. This variable is associated with UpperAlarmLimit attribute, which is available for Perseus A500, Atlan, Evita Infinity V500, Babylog VN500, Savina 300, Primus, Primus IE, Apollo, Zeus IE, Oxylog 3000 plus, Evita V300. This variable is associated with LowerAlarmLimit attribute, which is available for Perseus A500, Atlan, Evita Infinity V500, Babylog VN500, Savina 300, Primus, Primus IE, Apollo, Zeus IE, Oxylog 3000 plus, Evita V300. Possible units: mmHg, %, kPa. This variable is in Multiset. See variables 2355 and 6551 for concentration and partial pressure respectively. For Evita Infinity V500, Babylog VN500: This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. end tidal CO2 etCO2 Note: etCO2 values and alarm limits are transmitted according to the configuration on the device screen (mmHg or Vol%). If kPa is configured be aware that the Medibus alarm limits will be transmitted with mmHg.
244 | Tidal Volume | For Evita 4, Evita 2 dura, Evita XL: Expiratory tidal volume. This variable is associated with UpperAlarmLimit attribute, which is also sent in separate variable 2763. Values below 1 L will be transmitted in units of mL, otherwise in units of L. Device firmware version 04.00 and higher: Limit is not transmitted, if alarm is switched off. For Savina: This variable is associated with UpperAlarmLimit attribute.Device SW version 02.00 and higher: Limit is not transmitted, if alarm is switched off. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300, Fabius Family. This variable is associated with UpperAlarmLimit attribute, which is available for Perseus A500, Evita Infinity V500, Savina 300, Evita V300. This variable is associated with LowerAlarmLimit attribute, which is available for Evita Infinity V500, Evita V300. For Evita Infinity V500, Babylog VN500: This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. For Zeus, Zeus IE, Primus, Primus IE, Apollo, Pallas: Exp. Tidal Volume.
245 | Dynamic Compliance | For Evita 4, Evita 2 dura, Evita XL: Values below 100 will be transmitted with one decimal place, otherwise without a decimal place. For Savina: Available since device SW version 02.00. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Evita V300.
246 | CO2 Concentration Waveform | For Evita 4, Evita 2 dura, Evita XL: Possible units: mmHg, kPa, and % For Evita 2: Possible units: mmHg, kPa, and % Firmware version 2.00 or higher. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Evita V300. Possible units: mmHg, %, kPa. Unit "mmHg" is available for Perseus A500, Atlan, Evita Infinity V500, Babylog VN500, Savina 300, Primus, Primus IE, Apollo, Zeus IE, Oxylog 3000 plus, Evita V300. Unit "kPa" and "%" are available for Perseus A500, Atlan, Savina 300, Primus, Primus IE, Apollo, Zeus IE, Oxylog 3000 plus. See variables 8866 and 7872 for concentration and partial pressure respectively. For Evita Infinity V500, Babylog VN500: Possible units: mmHg, %. Expiratory CO2 concentration. For Zeus, Zeus IE, Primus, Primus IE, Apollo, Pallas, Vamos, Vamos plus: Possible units: mmHg, %, kPa.
247 | Airway Flow Wave | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300, Fabius Family. For Evita Infinity V500, Babylog VN500, Carina, Zeus, Zeus IE, Primus, Primus IE, Apollo, Pallas: Flow (insp./exp.)
248 | Resistance | For Evita 4, Evita 2 dura, Evita XL: Values below 100 will be transmitted with one decimal place, otherwise without a decimal place. For Savina: Available since device SW version 02.00. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Zeus IE, Evita V300.
249 | Ventilation Mode | The ventilation mode is reported in variable 584 and 2756. Variable 2756 contains mode ID numeric value, while variable 584 contains string representation which is the complete ventilation mode information concatenated from basic ventilation mode together with available options such as "Mode VC-CMV /AF". See variable 2756 for full list of supported ventilation mode.
250 | Inspired O2 (FiO2) | For Savina: This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. Available since device SW version 03.00. Limits are transmitted in LPO-Mode (Option) only. For Fabius GS, Fabius GS premium, Fabius Tiro, Fabius plus, Fabius plus XL, Fabius OS, Fabius MRI: This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Evita V300, Fabius Family. This variable is associated with UpperAlarmLimit attribute, which is available for Perseus A500, Atlan, Primus, Primus IE, Apollo, Zeus IE, Fabius Family. This variable is associated with LowerAlarmLimit attribute, which is available for Perseus A500, Atlan, Primus, Primus IE, Apollo, Zeus IE, Fabius Family. Inspiratory oxygen fraction, FiO2. For Zeus, Zeus IE: This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. For Primus, Primus IE, Apollo, Pallas: This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes.
251 | Device Name | This variable contains the device name defined by the manufacture. See variable 3591 for device model identification number.
252 | Inspiratory Component of I:E Ratio | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: Obsolete variable that may be removed in future release. Partial information of I:E ratio. For Evita Infinity V500, Savina: Partial information. Please use variable 1190 instead. Obsolete variable that may be removed in future release.
253 | Expiratory Component of I:E Ratio | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: Obsolete variable that may be removed in future release. Partial information of I:E ratio. For Evita Infinity V500, Savina: Partial information. Please use variable 1190 instead. Obsolete variable that may be removed in future release.
254 | Positive End Expiratory Pressure (PEEP) | For Fabius GS, Fabius GS premium, Fabius Tiro, Fabius plus, Fabius plus XL, Fabius OS, Fabius MRI: 1mbar = 1cmH2O. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300, Fabius Family.
255 | I:E Ratio | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Evita V300. Format: 1:x if E > I, x:1 if I > E, or 1:1 if I = E. The useless leading or ending “0” or “.” is not transmitted, and the maximum digits after the floating point is 2. For Evita Infinity V500, Babylog VN500: Format: 1:x if E > I, x:1 if I > E, or 1:1 if I = E. The useless leading or ending “0” or “.” is not transmitted, and the maximum digits after the floating point is 2.
256 | Minute Volume | For Evita 4, Evita 2 dura, Evita XL: This variable is associated with UpperAlarmLimit attribute, which is also sent in separate variable 1400. This variable is associated with LowerAlarmLimit attribute, which is also sent in separate variable 1399. For device firmware version 04.00 and higher: Lower alarm limit is not transmitted, if alarm is switched off. Values may be transmitted with one or two decimal place, depended on firmware version. Note that the value sent by the device is less precise than what is displayed on device screen. For example, a displayed value of 5.12 is sent by the device and thus by DDI as 5.1. For Evita 2, Evita, Fabius GS, Fabius GS premium, Fabius Tiro, Fabius plus, Fabius plus XL, Fabius OS, Fabius MRI, Evita Infinity V500, Babylog VN500, Carina: This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. This variable is available only with a precision to the tenth. For example: for a minute volume value of 7.43, the device displays 7.43 L/min but the DDI will output it as 7.4 L/min. For Savina: This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300, Fabius Family. This variable is associated with UpperAlarmLimit attribute, which is available for Perseus A500, Atlan, Savina 300, Primus, Primus IE, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Fabius Family. This variable is associated with LowerAlarmLimit attribute, which is available for Perseus A500, Atlan, Savina 300, Primus, Primus IE, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Fabius Family. This variable is available only with a precision to the tenth. For example: for a minute volume value of 7.43, the device displays 7.43 L/min but the DDI will output it as 7.4 L/min. For Zeus, Zeus IE, Primus, Primus IE, Apollo, Pallas: This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. Respiratory Minute Volume This variable is available only with a precision to the tenth. For example: for a minute volume value of 7.43, the device displays 7.43 L/min but the DDI will output it as 7.4 L/min. For Evita Infinity V500 and Babylog VN500: This variable is only output if the device version is less than 2.30, otherwise variable 2326 is exported.
257 | Spontaneous Respiration Rate | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Zeus IE, Oxylog 3000 plus, Evita V300. This variable is associated with UpperAlarmLimit attribute, which is available for Oxylog 3000 plus.
258 | Static Compliance | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500 Evita Infinity V600, Evita Infinity V800. For Evita Infinity V500: Cstat (Low Flow) Cstat.
259 | Mandatory Respiration Rate Setting | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300, Fabius Family. Respiratory rate, RR. For Evita Infinity V500, Babylog VN500: Respiratory Rate RR.
260 | Sigh Interval Setting | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. Time interval between sighs, Interval sigh.
261 | Sigh Number Setting | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. Number of sigh cycles, Cycles sigh.
262 | Apnea Tidal Volume Setting | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Oxylog 3000 plus, Carina, Evita V300. Tidal volume for apnea ventilation, Vtapn. Possible units: L, mL. Unit "L" is available for Evita Infinity V500, Babylog VN500, Savina 300, Oxylog 3000 plus, Carina, Evita V300. Unit "mL" is available for Evita Infinity V500, Babylog VN500, Evita V300. For Evita Infinity V500, Babylog VN500: Possible units: L, mL. VT Apnea VTapn.
263 | Apnea Respiration Rate Setting | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Oxylog 3000 plus, Carina, Evita V300. Respiratory rate for apnea ventilation, Rrapn.
264 | Pressure Support Level Above PEEP Setting | For Evita 4, Evita 2 dura, Evita XL: This variable is calculated from variables 1392 and 2104. Available only with Evita XL. For Medibus.X compatible device, Perseus A500, Atlan, Evita V300, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300, Fabius Family. Pressure amplitude above PEEP in Pressure Support, DPsupp. For Evita Infinity V500, Babylog VN500: ΔPsupp. For Carina: Assisted spontaneous breathing.
265 | Inspiratory Flow Setting | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300, Fabius Family.
266 | Inspiratory Flow Trigger Setting | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Evita V300, Fabius Family. Flow trigger threshold, Flow trigger. For Primus, Primus IE, Apollo, Pallas: Flow trigger.
267 | Intermittent PEEP Setting | For Fabius GS, Fabius GS premium, Fabius Tiro, Fabius plus, Fabius plus XL, Fabius OS, Fabius MRI: 1mbar = 1cmH2O. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Evita V300. Additional intermittent PEEP for sighs, PEEPsigh. For Evita Infinity V500, Babylog VN500: Interm. PEEP ΔintPEEP. For Zeus, Zeus IE: PEEP for all vent mode different from Man/Spont.
268 | APRV Low Pressure Setting | For Evita: BIPAP low pressure For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Evita V300. Lower pressure level during APRV, Plow. For Evita Infinity V500, Babylog VN500: APRV Low Pressure Plow.
269 | APRV Low Time Setting | For Evita: BIPAP low time For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Evita V300. Time of lower pressure level during APRV, Tlow. For Evita Infinity V500, Babylog VN500: APRV Low Time Tlow.
270 | APRV High Pressure Setting | For Evita: BIPAP high pressure For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Evita V300. Upper pressure level during APRV, Phigh. For Evita Infinity V500, Babylog VN500: APRV High Pressure Phigh.
271 | APRV High Time Setting | For Evita: BIPAP high time For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Evita V300. Time of upper pressure level during APRV, Thigh. For Evita Infinity V500, Babylog VN500: APRV High Time Thigh.
272 | Apnea Time Setting | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Oxylog 3000 plus, Carina, Evita V300. For Evita Infinity V500, Babylog VN500: Apnea Time Tapn. For Carina: Apnea alarm time.
273 | Maximum Inspired Pressure Setting | For Fabius GS, Fabius GS premium, Fabius Tiro, Fabius plus, Fabius plus XL, Fabius OS, Fabius MRI: 1mbar = 1cmH2O. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300, Fabius Family. Pressure limitation, Pmax. For Evita Infinity V500, Babylog VN500: Max Pressure Pmax. For Carina: Pressure limitation. For Primus, Primus IE, Apollo, Pallas: Max. insp. Airway pressure.
274 | Mean Airway Pressure | For Fabius GS, Fabius GS premium, Fabius Tiro, Fabius plus, Fabius plus XL, Fabius OS, Fabius MRI: This variable is associated with LowerAlarmLimit attribute. 1mbar = 1cmH2O. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300, Fabius Family. For Zeus, Zeus IE, Primus, Primus IE, Apollo, Pallas: Mean Breathing Pressure.
275 | Rapid Shallow Breathing Index | For Evita 4, Evita 2 dura, Evita XL: Firmware version 04.00 or higher For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Atlan, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Evita V300. Possible units: 1/L * min, 1/mL*min. Unit "1/L * min" is available for Evita Infinity V500, Babylog VN500, Savina 300, Evita V300. Unit "1/mL*min" is available for Evita Infinity V500, Babylog VN500, Evita V300 (added with SW 2.30 for Babylog VN500, added with SW 2.30 for Evita Infinity V500).
276 | Intrinsic PEEP | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300.
277 | Minimum Airway Pressure | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300.
278 | P0.1 Occlusion Pressure | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. Occlusion pressure in P0.1 maneuver. For Evita Infinity V500, Babylog VN500: Only available for Evita Infinity V500
279 | Lung Volume Trapped (TrpVol) | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. Volume trapped in the lung by intrinsic PEEP, Vtrap. For Evita Infinity V500: Trapped Volume Vtrap.
280 | Respired Volume Wave | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Evita V300. Respiratory volume since start of inspiration, Volume. For Evita Infinity V500, Babylog VN500: Volume since inspiration began.
281 | Ventilation Pressure Variability Setting | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800. For Evita Infinity V500, Babylog VN500: Supported through Medibus.X protocol.
282 | APRV Low Time Limitation | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. Maximum time of low pressure level during APRV/AutoRelease, Tlow max. For Evita Infinity V500, Babylog VN500: Tlow max.
283 | Spontaneous Mean Tidal Volume | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500 Evita Infinity V600, Evita Infinity V800. Spontaneous mean tidal volume, VTspon mean. For Evita Infinity V500, Babylog VN500: VTspon mean (Not available on Babylog VN500). Only available for Evita Infinity V500
284 | Spontaneous Inspired Mean Tidal Volume | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500 Evita Infinity V600, Evita Infinity V800. For Evita Infinity V500, Babylog VN500: VTispon mean (Not available on Babylog VN500). Only available for Evita Infinity V500
285 | Spontaneous Expired Mean Tidal Volume | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800. For Evita Infinity V500, Babylog VN500: VTespon mean (Not available on Babylog VN500). Only available for Evita Infinity V500
286 | SmartCare FiO2 Limitation | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. FiO2 in SmartCare - upper limit, SC-FiO2high For Evita Infinity V500, Babylog VN500: SC-Upper limit inspiratory O2 FiO2 max.
287 | SmartCare PEEP Limitation | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. PEEP in SmartCare - upper limit, SC-PEEPmax. For Evita Infinity V500, Babylog VN500: SC-Upper limit PEEP, PEEP max.
288 | SmartCare Spontaneous Respiration Rate Lower Limitation | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. Spontaneous respiratory rate in SmartCare - lower limit, SC-RRspon low. For Evita Infinity V500, Babylog VN500: SC-Lower limit respiratory rate in the guideline, RRspon low.
289 | SmartCare Spontaneous Respiration Rate Upper Limitation | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. Spontaneous respiratory rate in SmartCare - upper limit, SC-RRspon high. For Evita Infinity V500, Babylog VN500: SC-Upper limit respiratory rate in the guideline, RRspon high.
290 | SmartCare Tidal Volume Limitation | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. Tidal volume in SmartCare - lower limit, SC-VTlow. For Evita Infinity V500, Babylog VN500: SC-Lower limit tidal volume, VT low.
291 | SmartCare EtCO2 Limitation | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. End-tidal CO2 in SmartCare - upper limit, SC-etCO2high. For Evita Infinity V500, Babylog VN500: SC-Upper limit endtidal CO2 in the guideline etCO2 high.
292 | SmartCare Patient Height | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. Patient height in SmartCare, SC-Ht. For Evita Infinity V500: SC-Body height.
293 | Airway Resistance | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300.
294 | P01 Repetition Interval Setting | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. P0.1 repetition interval.
295 | Inspired Minute Volume | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Carina, Evita V300. This variable is associated with UpperAlarmLimit attribute, which is available for Carina. This variable is associated with LowerAlarmLimit attribute, which is available for Carina.
296 | Inspiratory Time Setting | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300, Fabius Family.
297 | Positive End Expiratory Pressure (PEEP) Setting | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300, Fabius Family. For Evita Infinity V500, Babylog VN500: PEEP / CPAP PEEP. For Zeus, Zeus IE: PEEP (CPAP). PEEP for mode Man/Spont.
298 | Inspired O2 (FiO2) Setting | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Zeus IE, Oxylog 3000 plus, Carina, Evita V300. For Primus, Primus IE, Apollo, Pallas: Not in device Apollo and Pallas
299 | Manual Breath Ventilation Pressure Setting | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300(added with SW 2.30 for Babylog VN500, added with SW 2.30 for Evita Infinity V500). Manual inspiratory pressure for Neonates during NIV - Non-Invasive Ventilation, PmanInsp.
300 | Manual Breath Maximum Time Limit | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300(added with SW 2.30 for Babylog VN500, added with SW 2.30 for Evita Infinity V500). Maximum time of pressure for manual inspiration pressure for Neonates during NIV - Non-Invasive Ventilation, TmanInsp.
301 | CO2 Tidal Production | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. CO2 production volume per breath, VTCO2. For Evita Infinity V500, Babylog VN500: Supported through Medibus.X protocol.
302 | CO2 Minute Production | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Evita V300. CO2 production, V'CO2.
303 | Device Date/Time | Format: YYYYMMDDhhmmss, where ss is always present but might be 00.
304 | Inspired Tidal Volume | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Carina, Evita V300. This variable is associated with UpperAlarmLimit attribute, which is available for Evita Infinity V500, Zeus IE, Evita V300. This variable is associated with LowerAlarmLimit attribute, which is available for Evita Infinity V500, Evita V300. For Primus, Primus IE, Apollo, Pallas: Insp. Tidal Volume.
305 | Expired Tidal Volume | For Evita XL: Neo expired tidal volume. Available only with firmware version 7.01 or higher. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Evita V300.
306 | Expired Minute Volume | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. This variable is associated with UpperAlarmLimit attribute, which is available for Evita Infinity V500, Babylog VN500, Evita V300. This variable is associated with LowerAlarmLimit attribute, which is available for Evita Infinity V500, Babylog VN500, Evita V300. For Evita Infinity V500, Babylog VN500: This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes.
307 | End Tidal CO2 Percentage | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Evita V300, Fabius Family. This variable is associated with UpperAlarmLimit attribute, which is available for Perseus A500, Atlan, Evita Infinity V500, Babylog VN500, Savina 300, Primus, Primus IE, Apollo, Zeus IE, Oxylog 3000 plus, Evita V300. This variable is associated with LowerAlarmLimit attribute, which is available for Perseus A500, Atlan, Evita Infinity V500, Babylog VN500, Savina 300, Primus, Primus IE, Apollo, Zeus IE, Oxylog 3000 plus, Evita V300. end tidal CO2 etCO2
308 | Fresh Gas O2 Flow Rate | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Evita V300. O2 therapy flow, Flow. For Evita Infinity V500, Babylog VN500: Flow O2 Therapy Flow.
309 | Ventilator Operation Mode | This variable contains ventilation operation mode in string format, which is also send as a numerical value in variable 2758. See detailed description of variable 2758 for the list of supported values. For Evita 4, Evita 2 dura, Evita XL: Firmware version 04.00 or higher For Evita 2: Variable label is not correct. Please use variable 3853 instead. Obsolete variable that may be removed in future release.
310 | Spontaneous Inspired Tidal Volume | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. For Zeus, Zeus IE: This variable is associated with UpperAlarmLimit attribute. Insp. Tidal volume.
311 | Airway Dead Space | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. Serial dead space volume, Vds. For Evita Infinity V500, Babylog VN500: Serial dead space Vds.
312 | Dynamic Compliance Over Last 20% Breath | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. Ratio of compliance during the last 20 % of inspiration over dynamic compliance, C20/Cdyn. For Evita Infinity V500, Babylog VN500: C20/Cdyn.
313 | Expired Spontaneous Tidal Volume | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300.
314 | Mechanical Expired Minute Volume | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Primus, Primus Infinity Empowered, Apollo, Evita V300. Mandatory expiratory minute volume, MVemand
315 | Spontaneous Expired Minute Volume | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Primus, Primus Infinity Empowered, Apollo, Evita V300.
316 | Patient Type ID | Consult detailed help for more information on this variable
317 | Airway Pressure Wave | For Fabius GS, Fabius GS premium, Fabius Tiro, Fabius plus, Fabius plus XL, Fabius OS, Fabius MRI: 1mbar = 1cmH2O. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300, Fabius Family. For Evita Infinity V500, Babylog VN500: Airway pressure Paw.
318 | Peak Airway Pressure | For Evita 4, Evita 2 dura, Evita XL: This variable is associated with UpperAlarmLimit attribute, which is also sent in separate variable 7040. For Evita 2, Evita, Savina: This variable is associated with UpperAlarmLimit attribute. For Fabius GS, Fabius GS premium, Fabius Tiro, Fabius plus, Fabius plus XL, Fabius OS, Fabius MRI: 1mbar = 1cmH2O. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300, Fabius Family. This variable is associated with UpperAlarmLimit attribute, which is available for Perseus A500, Atlan, Evita Infinity V500, Babylog VN500, Savina 300, Zeus IE, Oxylog 3000 plus, Carina, Evita V300, Fabius Family. This variable is associated with LowerAlarmLimit attribute, which is available for Zeus IE, Fabius Family. Peak inspiratory pressure, PIP. For Evita Infinity V500, Babylog VN500, Carina: This variable is associated with UpperAlarmLimit attribute. For Zeus, Zeus IE: This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. Peak Breathing Press. For Primus, Primus IE, Apollo, Pallas: Peak Breathing Pressure.
319 | Plateau Airway Pressure | For Fabius GS, Fabius GS premium, Fabius Tiro, Fabius plus, Fabius plus XL, Fabius OS, Fabius MRI: 1mbar = 1cmH2O. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Evita V300, Fabius Family. For Evita Infinity V500, Babylog VN500: Only available for Evita Infinity V500
320 | Spontaneous Tidal Volume | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Zeus IE, Evita V300. Possible units: L, mL. For Evita Infinity V500, Babylog VN500: Possible units: L, mL.
321 | I:E Ratio Setting | String format: 1:x if E > I, x:1 if I > E, or 1:1 if I = E. The useless leading or ending “0” or “.” is not transmitted, and the maximum digits after the floating point is 2. For Carina, Primus, Primus IE, Apollo, Pallas, Perseus A500: This variable is calculated by the DDI from respiration rate and inspiratory time settings. For Medibus.X compatible device: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300, Fabius Family. For Primus, Primus IE, Apollo, Perseus A500, Carina and Oxylog 3000 plus, it is calculated by the DDI from respiration rate and inspiratory time settings. For Atlan, Evita Infinity V500, Babylog VN500, Savina 300, Zeus IE, Evita V300, and Fabius Family, it is calculated by the DDI from inspiratory and expiratory components. For Evita, Evita 2, Evita 4, Evita 2 dura, and Evita XL, Fabius GS, Fabius GS premium, Fabius Tiro, Fabius plus, Fabius plus XL, Fabius OS, Fabius MRI, Evita Infinity V500, Babylog VN500, Zeus IE, Evita V300: This variable is calculated by the DDI from inspiratory and expiratory components of I:E ratio sent by the devices.
322 | Alarm Severity Level | This variable is in multiset format. This variable contains alarm priority level, variable 6164 contains alarm codes, and variable 3878 contains alarm messages. They need to be interpreted together. One alarm code may have one or more priority values and one or more alarm messages associated with it. The possible alarms are depended on device firmware version. Please see device user manuals for details. See variables 177 and 588 and 4131 for string format.
323 | Ventilator General Status | Ventilator Standby mode status.
325 |  | Consult detailed help for more information on this variable
326 | Ventilation Mode ID | This variable contains ID of basic ventilation mode ID without options. See also variable 584.
328 |  | Consult detailed help for more information on this variable
329 | Ventilator Operation Mode ID | For Evita 4, Evita 2 dura, Evita XL: Firmware version 04.00 or higher For Evita 2: Variable label is not correct. Please use variable 2467 instead. Obsolete variable that may be removed in future release.
331 |  | Consult detailed help for more information on this variable
332 | Tidal Volume Setting | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300, Fabius Family. Inspiratory tidal volume, Vti. For Evita Infinity V500, Babylog VN500: See also variable 1321 and 5044. For Zeus, Zeus IE: This variable contains same information as variable 1321, except that the value of different units (L or mL) are sent in separate variables. See also variable 5044.
333 | ASB Ramp Time Setting | For Evita 2: Firmware version 2.00 or higher. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300. Pressure rise time, Slope. For Carina: ASB ramp. For Primus, Primus IE, Apollo, Pallas: Slope time (Ramp).
334 | PS Activation State | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300, Fabius Family. A value of 1 indicates that Pressure Support (PS) is added to the current ventilation mode. A value of 0 indicates that Pressure Support (PS) is not added to the current ventilation mode. Text message displayed on device screen: "/PS"
335 | VS Activation State | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. A value of 1 indicates that Volume support (VS) is added to the current ventilation mode. A value of 0 indicates that Volume support (VS) is not added to the current ventilation mode. Text messages displayed on device screen: "/VS"
336 | APRV High Pressure (Phigh) | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. Upper pressure level during APRV, Phigh.
337 | APRV Low Pressure (Plow) | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. Lower pressure level during APRV, Plow.
338 | Inspired Pressure Setting | For Evita 4, Evita 2 dura, Evita XL: Available only with EvitaXL firmware version 7.00 or higher. For Fabius GS, Fabius GS premium, Fabius Tiro, Fabius plus, Fabius plus XL, Fabius OS, Fabius MRI: 1mbar = 1cmH2O. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300, Fabius Family.
339 | Device Average Gas Flow | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. Average device flow, FlowDev.
340 | Auto Flow Active Status | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300. A value of 1 indicates that AutoFlow (AF) is added to the current ventilation mode. A value of 0 indicates that AutoFlow (AF) is not added to the current ventilation mode. Text message displayed on device screen: "/AF"
341 | ATC Activation Status | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Evita V300. A value of 1 indicates that automatic tube compensation (ATC) is added to the current ventilation mode. A value of 0 indicates that automatic tube compensation (ATC) is not added to the current ventilation mode. Text message displayed on device screen: "/ATC" For Savina 300 (MEDIBUS.X): added with SW 5.00
342 | Device Number | Consult detailed help for more information on this variable
343 | Device Communications Status | Value "stop" sent when a request to stop the communication is received from the device
344 | Patient Type | This variable contains patient type in string format, which is also send as a numerical value in variable 2467. See detailed description of variable 2467 for the list of supported values.
345 | Generic Alarm Message | This variable is in multiset format. This variable contains alarm messages, variable 2596 contains alarm priority level, variable 6164 contains alarm codes. They need to be interpreted together. One alarm code may have one or more priority values and one or more alarm messages associated with it. The possible alarms are dependent on device firmware version. Please see device user manuals for details. See variables 177 and 588 and 4131 for string format. The HL7 output can contain an '&' character to indicate a LOW alarm. In order to avoid issues, the '&' character will be replaced by a '?' character in the DDI output.
346 | Inspired Mandatory Tidal Volume | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300.
347 | Spontaneous Fraction of Minute Volume | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. Spontaneous breathing portion of minute volume, % Mvspon. For Evita Infinity V500, Babylog VN500: Spontaneous Fraction Min. Vol. %MVspon (MVspon : MVtotal).
348 | Negative Inspiratory Force | For Evita 4, Evita 2 dura, Evita XL: The value is not transmitted, if it is displayed on Evita 4 / Evita 2 dura as "[less than] –45 mbar". Firmware version 04.00 or higher For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300.
349 | Auto Release Activation Status | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. A value of 1 indicates that AutoRelease is added to the current ventilation mode. A value of 0 indicates that AutoRelease is not added to the current ventilation mode. Text messages displayed on device screen: "/AutoRelease"
350 | Generic Respiration Rate | For Evita 4, Evita 2 dura, Evita XL: This variable is associated with UpperAlarmLimit attribute, which is also sent in separate variable 1339. Total Respiratory Rate (RRtotal). For Evita 2, Evita, Savina, Carina: This variable is associated with UpperAlarmLimit attribute. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300, Fabius Family. This variable is associated with UpperAlarmLimit attribute, which is available for Evita Infinity V500, Babylog VN500, Savina 300, Carina, Evita V300. Respiratory rate, RR. For Evita Infinity V500, Babylog VN500: This variable is associated with UpperAlarmLimit attribute. Respiratory Rate (RRtotal) RR.
351 | Minute Volume Alarm Delay | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. Delay for minute volume high alarm, MVhigh delay. This variable is associated with a ChannelID identifier.
352 | Flow Assist Level Setting | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. Flow assist during PPS ventilation, Flow Assist. For Evita Infinity V500, Babylog VN500: Flow Assist.
353 | Volume Assist Level Setting | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. Volume assist during PPS ventilation, Vol. Assist. For Evita Infinity V500, Babylog VN500: Vol. Assist.
354 | Disconnection Alarm Delay Time | For Evita 4, Evita 2 dura, Evita XL: Firmware version 4.00 or higher. For Savina: Tdeconnect. Available since Device SW version 02.00. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Carina, Evita V300. Disconnection alarm time, Tdiscon. For Evita Infinity V500, Babylog VN500: Tdisconnect. For Carina: Disconnection detection timeout.
355 | Medibus Revision
356 | Device Revision | Format: mm.nn
357 | VG Activation Status | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Carina, Evita V300. A value of 1 indicates that Volume Guarantee (VG) is added to the current ventilation mode. A value of 0 indicates that Volume Guarantee (VG) is not added to the current ventilation mode. Text message displayed on device screen: "/VG"
358 | Inspiratory Flow Waveform | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300.
359 | Spontaneous Inspiration Time | For Evita 4, Evita 2 dura, Evita XL: Available only with EvitaXL firmware version 7.02 or higher. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300.
360 | Lung Time Constant | For Evita 4, Evita 2 dura, Evita XL: Available only with EvitaXL firmware version 7.00 or higher. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300.
361 | Airway Correlation Coefficient | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. Correlation factor, r2. For Evita Infinity V500, Babylog VN500: r².
362 | Sigh Activation Status | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. A value of 1 indicates that the sigh function has been added to the current ventilation mode. A value of 0 indicates that the sigh function has not been added to the current ventilation mode. Text message displayed on device screen: "/Sigh"
363 | Inspired Tidal Volume Setting | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Zeus IE, Evita V300. Inspiratory tidal volume, Vti. For Evita Infinity V500, Babylog VN500: See also variable 1321 and 2771. For Zeus, Zeus IE: This variable contains same information as variable 1321, except that the value of different units (L or mL) are sent in separate variables. See also variable 2771.
364 | Dead Space to Tidal Volume Ratio | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. Ratio of serial dead space volume to expiratory tidal volume, Vds/VTe. For Evita Infinity V500, Babylog VN500: Dead space ventilation Vds/VTe.
365 | Expired Mandatory Tidal Volume | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300.
366 | Expiratory Time Constant (RCexp) | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300.
367 | EtCO2 Phase III Increment | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. Possible units: mmHg/L, kPa/L, Vol%/L. Increase of measured CO2 value in phase III of the capnogram, CO2 slope.
368 | Spontaneous I:E Ratio | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. Format: 1:x if E > I, x:1 if I > E, or 1:1 if I = E. The useless leading or ending “0” or “.” is not transmitted, and the maximum digits after the floating point is 2. For Evita Infinity V500, Babylog VN500: Format: 1:x if E > I, x:1 if I > E, or 1:1 if I = E. The useless leading or ending “0” or “.” is not transmitted, and the maximum digits after the floating point is 2.
369 | General Alarm Code | This variable is in multiset format. This variable contains alarm codes, variable 3878, alarm messages, variable 2596 contains alarm priority level. They need to be interpreted together. One alarm code may have one or more priority values and one or more alarm messages associated with it. The possible alarms are depended on device firmware version. Please see device user manuals for details. See variables 177 and 588 and 4131 for string format.
371 |  | Consult detailed help for more information on this variable
372 | Endotracheal Tube Size | For Evita 4, Evita 2 dura, Evita XL: Possible units: "0.1mm" in firmware versions 7.00 and 7.01, "mm" otherwise. Available only with EvitaXL firmware version 7.00 or higher. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Evita V300. Inner diameter of the endotracheal/tracheostoma tube. For Evita Infinity V500, Babylog VN500: ATC Tube Diameter.
373 | End Tidal CO2 Pressure | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Evita V300, Fabius Family. This variable is associated with UpperAlarmLimit attribute, which is available for Perseus A500, Atlan, Evita Infinity V500, Babylog VN500, Savina 300, Primus, Primus IE, Apollo, Zeus IE, Oxylog 3000 plus, Evita V300. This variable is associated with LowerAlarmLimit attribute, which is available for Perseus A500, Atlan, Evita Infinity V500, Babylog VN500, Savina 300, Primus, Primus IE, Apollo, Zeus IE, Oxylog 3000 plus, Evita V300. Possible units: mmHg, kPa. end tidal CO2 etCO2
374 | Expiratory Time Setting | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Zeus IE, Evita V300.
375 | Inspiratory Component of I:E Ratio Setting | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Zeus IE, Evita V300, Fabius Family, Atlan For Evita Infinity V500: Partial information. Please use variable 2583 instead. Obsolete variable that may be removed in future release.
376 | Expiratory Component of I:E Ratio Setting | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800,Savina 300, Zeus IE, Evita V300, Fabius Family, Atlan For Evita Infinity V500: Partial information. Please use variable 2583 instead. Obsolete variable that may be removed in future release.
377 | Lung Elastance | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300.
378 | Generic Ventilation Mode | Consult detailed help for more information on this variable
379 | Ventilation Mode Primary | Consult detailed help for more information on this variable
380 | Ventilation Mode Options | Consult detailed help for more information on this variable
381 | Variable PS Activation Status | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800. A value of 1 indicates that Variable PS (VPS) has been added to the current ventilation mode. A value of 0 indicates that Variable PS (VPS) has not been added to the current ventilation mode. Text message displayed on device screen: "VariablePS"
382 | Patient Operation Mode | Possible values: 1 for Invasive, 2 for Noninvasive
383 | Artificial Airway Compensation Status | For Evita 4, Evita 2 dura, Evita XL: Tube Compensation. Available only with EvitaXL firmware version 7.00 or higher. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. Compensation rate during ATC, % comp. For Evita Infinity V500, Babylog VN500: ATC Compensation Compens.
384 | Mandatory Respiration Rate | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Zeus IE, Evita V300.
385 | Mandatory Tidal Volume | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Babylog VN500, Babylog VN600, Babylog VN800, Zeus IE, Evita V300(added with SW 2.30 for Babylog VN500, added with SW 2.30 for Evita Infinity V500). Possible units: L, mL. Unit "L" available for Evita Infinity V500, Babylog VN500, Evita V300 (added with SW 2.30 for Babylog VN500, added with SW 2.30 for Evita Infinity V500). Unit "mL" available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Zeus IE, Evita V300.
386 | CO2 Partial Pressure Waveform | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Evita V300. Possible units: mmHg, kPa. Unit "mmHg" is available for Perseus A500, Atlan, Evita Infinity V500, Babylog VN500, Savina 300, Primus, Primus IE, Apollo, Zeus IE, Oxylog 3000 plus, Evita V300. Unit "kPa" is available for Perseus A500, Atlan, Savina 300, Primus, Primus IE, Apollo, Zeus IE, Oxylog 3000 plus.
387 | Time Duration at Low Pressure (Tlow) | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. Effective expiratory time during APRV/AutoRelease, Tlow. For Evita Infinity V500, Babylog VN500: Tlow (resulting Tlow from expiratory flow cycling).
388 | Spontaneous Tidal Volume
389 | Generic Message Code | This variable is in multiset format. This variable, along with variable 8696, contains text message code and alarm text respectively. They need to be interpreted together.
390 | Generic Message Text | This variable is in multiset format. This variable, along with variable 8366, contains text message code and alarm text respectively. They need to be interpreted together.
391 | End Inspiratory Pressure | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. End-inspiratory pressure, EIP.
392 | Tracheal Pressure Waveform | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Evita V300. For Evita Infinity V500, Babylog VN500: Ptrach.
393 | Tidal Volume Indexed By Body Weight | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. Tidal volume per kg body weight, VT/Wt. The leakage is determined by the difference between the delivered flow on the inspiratory side and the measured flow on the expiratory side, which is displayed as the leakage minute volume MVleak and relative as % leak (MVleak to MV). For Evita Infinity V500, Babylog VN500: VT/kg BW.
394 | Mandatory Triggered Respiration Rate | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300(added with SW 2.42 for Evita Infinity V500, Evita V300 and Babylog VN500). Respiratory rate of triggered, mandatory breaths, RRtrig For Evita 4, Evita 2 dura, Evita XL: Available only with Evita XL firmware version 7.00 or higher.
395 | Volume of Leakage (MVLeak) | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Carina, Evita V300. Leakage minute volume, Mvleak. The leakage is determined by the difference between the delivered flow on the inspiratory side and the measured flow on the expiratory side, which is displayed as the leakage minute volume MVleak and relative as % leak (MVleak to MV). For Evita Infinity V500, Babylog VN500: Mvleak. The leakage is determined by the difference between the delivered flow on the inspiratory side and the measured flow on the expiratory side, which is displayed as the leakage minute volume MVleak and relative as % leak (MVleak to MV). For Carina: Leakage minute volume. For Evita XL: Leakage minute volume. Available since device SW version 7.02.
396 | Percent of Leakage | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Evita V300. Leakage minute volume in % of inspiratory minute volume, % leak Leakage rel. % leak - Infinity V500 determines the difference between the delivered flow on the inspiratory side and the measured flow on the expiratory side. This difference provides a measure of the amount of leakage and is displayed by Infinity V500 as the leakage minute volume MVleak and relative as % leak (MVleak to MV). For Evita Infinity V500, Babylog VN500: Leakage rel. % leak.
397 | Inspiratory Component of Spontaneous I:E Ratio | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: Obsolete variable that may be removed in future release. Partial information of I:E ratio. For Evita Infinity V500: Partial information. Please use variable 5815 instead. Obsolete variable that may be removed in future release.
398 | Expiratory Component of Spontaneous I:E Ratio | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: Obsolete variable that may be removed in future release. Partial information of I:E ratio. For Evita Infinity V500: Partial information. Please use variable 5815 instead. Obsolete variable that may be removed in future release.
399 | SmartCare Target Delta Support Pressure | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. Dpressure support goal in SmartCare, SC-DPsupp goal. For Evita Infinity V500: SC-Target ΔPsupp ΔPsupp goal.
400 | SmartCare Rated Delta Support Pressure | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. Rated Dpressure support by internal controller in SmartCare, SC-DPsupp rated. For Evita Infinity V500: SC-Rated ΔPsupp SC-Δpsupp.
401 | SmartCare Duration of Patient Session | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. Duration of session (hours and minutes) in SmartCare, SC-duration. Variable is in Multiset format (need combine the first value in hours with the second value in minutes to obtain the exact duration). Please use variable 4346 instead. For Evita Infinity V500: Multiset variable (need combine the first value in hours with the second value in minutes to obtain the exact duration).
402 | SmartCare Spontaneous Respiration Rate | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. Spontaneous respiratory rate in SmartCare, SC-Rrspon. For Evita Infinity V500: SC-Rrspon.
403 | SmartCare Tidal Volume | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. Tidal volume in SmartCare, SC-VT. For Evita Infinity V500: SC-VT.
404 | SmartCare EtCO2 | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. End-tidal CO2 concentration in SmartCare, SC-etCO2. For Evita Infinity V500: SC-etCO2.
405 | Maximum Inspiratory Time Setting | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Carina, Evita V300. Maximal inspiratory time for flow cycled strokes, Timax. For Carina: Maximal inspiration time for CPAP, ASB.
406 | Inspiratory Flow Termination Level Setting | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Savina 300, Oxylog 3000 plus, Evita V300. Inspiratory termination criterion based on peak inspiratory flow, Insp. term. For Evita Infinity V500: Inspiratory termination.
407 | Expiratory Flow Termination Setting | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. Expiratory termination criterion during APRV/AutoRelease, Exp. term. For Evita Infinity V500, Babylog VN500: Expiratory termination.
408 | SmartCare Body Weight | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. Ideal body weight in SmartCare, SC-Wt. For Evita Infinity V500: SC-Body weight.
409 | SmartCare Start of Night Rest | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. Start of night rest in SmartCare, SC-startrest. For Evita Infinity V500: SC-Start of night rest.
410 | SmartCare End of Night Rest | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. End of night rest in SmartCare, SC-endrest. For Evita Infinity V500: SC-End of night rest.
411 | Device Alarm Inactivation State
413 | Infusion Pump Type | 0: Smart or Go 2: Vision or Go+
414 | Pump Infusion Status | Consult detailed help for more information on this variable
415 | Device Serial Number | Format: 8 characters string. This variable is buffered and sent repeatedly by the DDI with each data message.
416 | Instrument Code | Format: 6 characters string
417 | Device Software Revision | Format: vv.r.c vv = software version r = software revision c = software correction
418 | Battery Charge | 58112 => battery charged
419 | Device Alarm Status | The alarms code are output as decimal values. Consult detailed help for more information on this variable
420 | Alarm History | 16 last alarms
421 | Product Descritpion | 4 characters for product identification
422 | Current Delivery Rate | If the pump is OFF, the return rate is 0.
423 | Cumulated Infused Volume
424 | Current Volume Infused
425 | Error Response From Device | Error code sent by the pump when an incorrect transmission happens.
426 | Device Pre-alarm Status | The alarms code are output as decimal values. Consult detailed help for more information on this variable
427 | Device Type | This variable contains a fixed value "48 (0x30)" for Infusion Pump.
428 | Infusion Pump Type | Consult detailed help for more information on this variable
429 | Infusion Pump Serial Number | 8 characters string
430 | Infusion Pump Rack Index | The variable has a fixed value "1" because it is a standalone pump.
431 | Infusion Pump Slot Index | The variable has a fixed value "1" because it is a standalone pump.
432 | Generic Infusion Pump Mode | Consult detailed help for more information on this variable
433 | Actual Infusion Volumetric Rate
434 | Total Volume Infused | This variable is reset to "0" only if users do so manually on the pump panel. It is not automatically reset if the infusion has been interrupted and/or the pump has been switched off.
435 | Infusion Drug Name | This variable contains a fixed nonsignificant value "NA"
436 | Infusion Volumetric Rate Setting
437 | Infusion Event | Consult detailed help for more information on this variable
439 | Device Name | This variable is obsolete for MultiFiltrate and might be removed in a future version.
440 | Cumulated Therapy Time | Treatment time, running, total
441 | Anticoagulant Cumulated Continuous Infused Volume | Anticoagulation, continuous, total
442 | Anticoagulant Cumulated Bolus Infused Volume | Anticoagulation bolus, total
443 | Ultrafiltration/Blood Flow Ratio
444 | Arterial Pressure Upper Alarm Limit | This variable is also sent as alarm limit of variableID 7438.
445 | Arterial Pressure Lower Alarm Limit | This variable is also sent as alarm limit of variableID 7438.
446 | Device Serial Number
447 | Patient ID
448 | Device System Code
449 | Citrate Calcium Activation Status | Possible values:
450 |  | 1: CiCa requested
451 |  | 0: CiCa not requested.
452 | Therapy Pause Start Time | Format: YYYYMMDDhhmmss
453 | Device Status | Consult detailed help for more information on this variable
454 | Device Software Revision
455 | Therapy Pause End Time | Format: YYYYMMDDhhmmss
456 | Disposable Lifetime | The resolution of this variable is 10 min. This variable is sent from the device to the DDI on a regular basis, but the device will keep on sending the same value until a minimal change of 10 min is observed.
457 | Blood Flow Rate Setting
458 | Replacement Flow Rate Setting | Substitution rate setting
459 | Dialysate Flow Rate Setting
460 | Fluid Removal Volume Setting | UF goal setting
461 | Transmembrane Pressure Upper Alarm Limit | This variable is also sent as alarm limit of variableID 3029.
462 | Filter Pod Pressure | Filtrate pressure
463 |  | For Multifiltrate PRO: The resolution of this variable is 10 mmHg. This variable is sent from the device to the DDI on a regular basis, but the device will keep on sending the same value until a minimal change of 10 mmHg is observed.
464 | Transmembrane Pressure | This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes.
465 |  | For Multifiltrate PRO: The resolution of this variable is 10 mmHg. This variable is sent from the device to the DDI on a regular basis, but the device will keep on sending the same value until a minimal change of 10 mmHg is observed.
466 | Therapy Type | Consult detailed help for more information on this variable
467 | Generic Venous Line Pressure | This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes.
468 |  | For Multifiltrate PRO: The resolution of this variable is 10 mmHg. This variable is sent from the device to the DDI on a regular basis, but the device will keep on sending the same value until a minimal change of 10 mmHg is observed.
469 | Therapy Ending Time
470 | Therapy Starting Time | Date/time in the format of "YYYYMMDDhhmmss"
471 | Device Total Operation Time | Format: [hours]:[minutes]
472 | Replacement Cumulated Volume | Substituate-Total, total
473 | Fluid Removal Cumulated Volume | Effective total withdrawal, total
474 | Dialysis Filter Lifetime | Filter Lifetime
475 |  | For Multifiltrate PRO: The resolution of this variable is 10 min. This variable is sent from the device to the DDI on a regular basis, but the device will keep on sending the same value until a minimal change of 10 min is observed.
476 | Blood Sensed Status | 0=no blood sensed 1=blood sensed
477 | Fluid Removal Rate Setting | UF rate setting
478 | Dialysate Temperature Setting
479 | Plasma Volume Setting
480 | Citrate Dose Setting
481 | Calcium Dose Setting
482 | Citrate Flow Rate
483 | Calcium Flow Rate
484 | Citrate Cumulated Volume
485 | Calcium Cumulated Volume
486 | Filtration Fluid Cumulated Balance | DF-Total, total balance
487 | Replacement Cumulated Bolus Volume | Bolus-Substituate-Total, total
488 | Central Venous Pressure Lower Alarm Limit | This variable is also sent as alarm limit of variableID 3248.
489 | Central Venous Pressure Upper Alarm Limit | This variable is also sent as alarm limit of variableID 3248.
490 | General Alarm Code | Consult detailed help for more information on this variable
491 | Device Quick Test Status | T0 Test Status. Possible values: Start, End, E1, E2, E3.
492 | Generic Arterial Line Pressure | This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes.
493 |  | For Multifiltrate PRO: The resolution of this variable is 10 mmHg. This variable is sent from the device to the DDI on a regular basis, but the device will keep on sending the same value until a minimal change of 10 mmHg is observed.
494 | Anticoagulant Infusion Rate Setting | Heparin continuous rate
495 | Patient Weight Cumulated Value | Absolute balancing error
496 | Patient Sensor Value 1 | Weight of scale 1
497 |  | For Multifiltrate PRO: The resolution of this variable is 10 g. This variable is sent from the device to the DDI on a regular basis, but the device will keep on sending the same value until a minimal change of 10 g is observed. The value of "0" will always be sent regardless of the delta change.
498 | Patient Sensor Value 2 | Weight of scale 2
499 |  | For Multifiltrate PRO: The resolution of this variable is 10 g. This variable is sent from the device to the DDI on a regular basis, but the device will keep on sending the same value until a minimal change of 10 g is observed. The value of "0" will always be sent regardless of the delta change.
500 | Patient Sensor Value 3 | Weight of scale 3
501 |  | For Multifiltrate PRO: The resolution of this variable is 10 g. This variable is sent from the device to the DDI on a regular basis, but the device will keep on sending the same value until a minimal change of 10 g is observed. The value of "0" will always be sent regardless of the delta change.
502 | Patient Sensor Value 4 | Weight of scale 4
503 |  | For Multifiltrate PRO: The resolution of this variable is 10 g. This variable is sent from the device to the DDI on a regular basis, but the device will keep on sending the same value until a minimal change of 10 g is observed. The value of "0" will always be sent regardless of the delta change.
504 | Patient Fluid Balance | Balance, total
505 | Therapy Operation Status | Consult detailed help for more information on this variable
506 | Case ID
507 | Blood Flow Rate | Actual blood flow This variable contains value "0" when blood pump is stopped.
508 |  | For Multifiltrate PRO: The resolution of this variable is 10 mL/min. This variable is sent from the device to the DDI on a regular basis, but the device will keep on sending the same value until a minimal change of 10 mL/min is observed. The value of "0" will always be sent regardless of the delta change.
509 | Prefilter Pod Pressure | Pre-Hemofilter pressure
510 |  | For Multifiltrate PRO: This variable is associated with LowerAlarmLimit and UpperAlarmLimit attributes.
511 |  | The resolution of this variable is 10 mmHg. This variable is sent from the device to the DDI on a regular basis, but the device will keep on sending the same value until a minimal change of 10 mmHg is observed.
512 | Transmembrane Pressure Lower Alarm Limit | This variable is also sent as alarm limit of variableID 3029.
514 | Heart Rate | This variable is associated with LowerAlarmLimit and UpperAlarmLimit attributes.
515 | NIBP Systolic | This variable is associated with LowerAlarmLimit and UpperAlarmLimit attributes.
516 | NIBP Diastolic | This variable is associated with LowerAlarmLimit and UpperAlarmLimit attributes.
517 | NIBP Mean | This variable is associated with LowerAlarmLimit and UpperAlarmLimit attributes.
518 | Patient Delta Temperature
519 | Generic Dialysate Alarm | Value of "1" when alarm is active.
520 | Urea Clearance Alarm | Value of "1" when alarm is active.
521 | Blood Volume Alarm | Value of "1" when alarm is active.
522 | Mixed Dilution Alarm | Value of "1" when alarm is active.
523 | Single Needle Ratio Setting
524 | Target Treatement Effectiveness (Kt/V) Remaining Time
525 | Shunt Blood Flow Rate
526 | Recirculation Measurement Active Status | Possible values: 0=off, 1=on
527 | Isolated Ultrafiltration Therapy Time
528 | Isolated Ultrafiltration Rate
529 | Initial Hemoglobin
530 | Initial Protein Concentration
531 | Single Needle Time
532 | Blood Outlet Pressure
533 | Dialysate Inlet Pressure
534 | Dialysate Outlet Pressure
535 | Maximum Transmembrane Pressure
536 | Auto Dialysate Flow Enabled Status | Possible values: 0=disabled, 1=enabled
537 | Auto Dialysate Flow Factor Setting
538 | Auto AntiCoagulant Enabled Status | Possible values: 0=disabled, 1=enabled
539 | Auto Anticoagulant Bolus Enabled Status | Possible values: 0=off, 1=on
540 | Auto Anticoagulant Bolus Rate Enabled Status | Possible values: 0=disabled, 1=enabled
541 | Single Needle Auto Rate Enabled Status | Possible values: 0=disabled, 1=enabled
542 | Replacement Pump Running Status | Possible values: 0=off, 1=on
543 | Clearance Measurement Interval Setting
544 | Low Venous Line Pressure Alarm | Value of "1" when alarm is active.
545 | High Venous Line Pressure Alarm | Value of "1" when alarm is active.
546 | High Arterial Line Pressure Alarm | Value of "1" when alarm is active.
547 | Low Arterial Line Pressure Alarm | Value of "1" when alarm is active.
548 | Heparin Pump Alarm | Value of "1" when alarm is active.
549 | Patient Specific Filtration Ratio Setting
550 | Average Dialysate Temperature
551 | Cuff Inflation Pressure
552 | Bicarbonate Concentration (HCO3)
553 | Device Name
554 | Patient Weight
555 | Predilution Enable Status | Internal status for the mixed HDF control
557 |  | Consult detailed help for more information on this variable
558 | Clearance Measurement Enabled Status | Possible values: 0=off, 1=on
559 | Maximum Ultrafiltration Rate Setting
560 | Ultrafiltration Goal Positive Deviation Setting
561 | Ultrafiltration Goal Negative Deviation Setting
562 | Residual Blood Volume Active Status | Possible values: 0=off, 1=on
563 | Temperature Change Rate Setting
564 | Reinfusion Type Setting | Consult detailed help for more information on this variable
565 | Auto Reinfusion Active Status | Possible values: 0=off, 1=on
566 | Low Volume Treatment Active Status | Possible values: 0=off, 1=on
567 | Battery Running Alarm | Value of "1" when alarm is active.
568 | Cumulated Therapy Time
569 | Current Therapy Time
570 | Current Therapy ID
571 | Device IP Address
572 | Therapy Running Status | Consult detailed help for more information on this variable
573 | Manually Entered Hematocrit
574 | Anticoagulant Cumulated Bolus Infused Volume
575 | Average Anticoagulant Rate
576 | Single Needle Stroke Volume
577 | Average Urea Clearance
578 | Online Bolus Volume
579 | Online Bolus Rate Setting
580 | Patient Birth Date | Format: YYYYMMDD
581 | Hematocrit
582 | Hemoglobin (Hb)
583 | Concentrate Source | Consult detailed help for more information on this variable
584 | Maximum Fluid Removal Rate Setting
585 | Average Blood Flow Rate
586 | Reinfusion Volume
587 | Device Alarm Global Status | Format: Bit 2: 0 --> Alarm
588 | Blood Pump Alarm | Value of "1" when alarm is active.
589 | Device Serial Number
590 | Patient ID
591 | Therapy Options | Consult detailed help for more information on this variable
592 | NIBP Interval Time
593 | Relative Blood Volume Change Limit
594 | Generic Device Alarm | Format: list of maximum 16 alarm IDs separated with a "¦".
596 |  | Consult detailed help for more information on this variable
597 | Urea Distribution Volume Setting
598 | Auto Ultrafiltration Control Activation Status | Possible values: 0=off, 1=on
599 | Device Software Revision
600 | Patient First Name
601 | Blood Flow Rate Setting
602 | Replacement Flow Rate Setting
603 | Dialysate Flow Rate Setting
604 | Fluid Removal Volume Setting
605 | Transmembrane Pressure
606 | Patient Profile | Na profile number
607 | Therapy Profile
608 | Dialysate Sodium Start Concentration Setting
609 | Fluid Removal Actual Rate
610 | Average Replacement Flow Rate
611 | Minimum Relative Treated Blood Volume
612 | Auto Replacement Control Activation Status | Possible values: 0=disabled, 1=enabled
613 | Relative Blood Volume Change
614 | Replacement Volume Setting
615 | Generic Venous Line Pressure
616 | Total Replacement Cumulated Volume | Expected total substitution volume (pre + post) at the end of the treatment.
617 | Blood Circuit Pressure Alarm | Value of "1" when alarm is active.
618 | Recirculation Percentage
619 | Filtration Ratio
620 | Ultrafiltration Remaining Time | Remaining ISO UF time.
621 | Total Replacement Flow Rate
622 | Replacement Post-delivered Percentage
623 | Arterial Temperature
624 | Venous Temperature
625 | Device Type
626 | Therapy Ending Time | Format: YYYYMMDDhhmm
627 | Therapy Starting Time | Format: YYYYMMDDhhmm
628 | Fluid Removal Cumulated Volume
629 | Single Needle Max Pressure Setting
630 | Single Needle Min Pressure Setting
631 | Plasma Sodium (Na) | For 6008 only: This variable will contain the Online Clearance Monitoring plasma sodium concentration when sodium management is being used.
632 | Therapy Remaining Time | Remaining UF time including ISO UF time.
633 | Device Communication Mode | Consult detailed help for more information on this variable
634 | Dialyzer Name
635 | Dialysate Name
636 | Dialysate Average Flow Rate
637 | Dialysate Sodium (Na+) Concentration Setting
638 | Dialysate Bicarbonate (HCO3) Concentration Setting
639 | Urea Clearance
640 | Fluid Removal Average Rate
641 | Anticoagulant Stopping Interval Setting
642 | Target Treatement Effectiveness (Kt/V) Setting
643 | Initial Hematocrit
644 | Serum Protein
645 | Total Protein
646 | Blood Temperature Alarm | Value of "1" when alarm is active.
647 | Total Replacement Cumulated Volume | Cumulated total substitution volume (pre + post).
648 | Single Needle Blood Pump Rate
649 | Therapy Duration Setting
650 | Anticoagulant Cumulated Volume
651 | Delivered Treatment Effectiveness (Kt/V)
652 | Dialysate Flow Rate
653 | Disinfect Mode Active Status | Consult detailed help for more information on this variable
654 | Dialysis Mode Active Status | Consult detailed help for more information on this variable
655 | Effective Blood Volume Processed (Kt)
656 | Dialysate Temperature
657 | Fluid Removal Rate Setting
658 | Dialysate Temperature Setting
659 | Replacement Cumulated Bolus Volume
660 | Actual Anticoagulant Rate
661 | Therapy Processed Percentage
662 | Blood Plasma Inlet Pressure
663 | Temperature Module General Status | Possible values: 0=off, 1=on
664 | Last Desinfection Time | Format: YYYYMMDDhhmm
665 | Last Therapy End Time | Format: YYYYMMDDhhmm
666 | Dialysis Pump Stop Status | Consult detailed help for more information on this variable
667 | Isolated Ultrafiltration Volume
668 | Single Needle Stroke Volume | Average single needle stroke volume
669 | Initial Ultrafiltration Factor Setting
670 | Generic Arterial Line Pressure
671 | Dialysate Conductivity
672 | Dialysate Conductivity Alarm | Value of "1" when alarm is active.
673 | Bolus Manual Enable Status | Possible values: 0=disabled, 1=enabled
674 | Plasma Filtration Fraction Setting
675 | Patient Family Name
676 | Replacement Pre-delivered Rate Setting
677 | Replacement Post-delivered Rate Setting
678 | Anticoagulant Bolus Volume Setting
679 | Anticoagulant Infusion Rate Setting
680 | Processed Cumulated Blood Volume
681 | Replacement Post-delivered Cumulated Volume
682 | Therapy Operation Status | Consult detailed help for more information on this variable
683 | Device Manufacturer
684 | Blood Flow Rate
685 | Device Location
687 | Device ID | The globally unique identifier for a Bed entity as defined by Hill-Rom (e.g. “VersaCare-0x0001”).
688 | Device Name | The bed model name.
689 | Patient Weight | The last patient weight captured by a caregiver from the bedside displayed in kilograms. -9999.9 -- 9999.9
691 | Base Device ID | The globally unique identifier for the BedConnector entity associated with this Bed entity or an empty string if no association currently exists.
692 | Bed Head Angle Alarm | The current status for the head angle alarm system.
693 | Power Supply Status | The current status of the AC power system.
694 | Device Serial Number | The serial number of the bed communication controller. 0x000000000000 - 0xFFFFFFFFFFFF
696 | Patient ID | The globally unique identifier for the Patient entity associated with this Bed entity or an empty string if no association currently exists.
697 | Patient Name | Easily recognizable patient name that can be used by authorized personnel to locate the specific physical patient (e.g. j_doe…).
698 | Patient Location ID | The globally unique identifier for the Location entity associated with this Bed entity or an empty string if no association currently exists.
699 | Patient Location Text | Easily recognizable location name that can be used by authorized personnel to locate the specific physical location (e.g. Tower1_WingC_Floor2_Room103).
700 | Remote Alarm Application Mode | The current configuration mode of the SafeView alarm system.
701 | Remote Alarm Application Indicator Status | The current status of the SafeView alarm system indicators.
702 | Patient Weight Scale Activity | The last command to be successfully processed by the weigh scale system.
703 | Remote CIS Application Status | The current status for the NaviCare alerts switch.
704 | Remote CIS Application Indicator Status | The current status for the NaviCare alerts indicator.
705 | Nurse Call Indicator Status | The current status for the nurse call indicator.
706 | Nurse Answer Indicator Status | The current status for the nurse answer indicator.
707 | Bed Height Position | The bed’s physical articulated position.
708 | Bed Lockout Status
709 | Bed Head Rails Position | The physical positions of the head-end siderails.
710 | Bed Foot RailsPosition | The physical positions of the foot-end siderails.
711 | Bed Surface Support Mode | The current configuration mode of the patient support surface system.
712 | Bed Emergency CPR Status | The current status for the CPR switch.
713 | Bed Clearance Status | The current status for the Bed Cleaned switch.
714 | Bed Rotation Therapy Status | The current status for the rotation therapy system.
715 | Bed Percussion Therapy Status | The current status for the percussion therapy system.
716 | Bed Vibration Therapy Status | The current status for the vibration therapy system.
717 | Patient Environment Control Activity | The last command completed by the patient environment control system.
718 | Patient Environment Activity Cleared Status | The last command completed by the patient environment control system.
719 | Bed Head Angle Alarm Mode | The current mode configuration for the head angle alarm system.
720 | Bed Head Angle Alarm Silence Status | The current mode configuration for the audible annunciation portion of the head angle alarm system.
721 | Bed Clearance Indicator Status | The current status for the bed cleaned indicator.
722 | SmartBed Patient Movement State | The current direction of patient activity as detected by the patient positioning monitor (PPM) alarm system.
723 | Battery Module Status | The current status of the battery power system.
724 | Device Software Revision | The software revision of the bed communication controller. 0.00 -- 9.99
726 | Device Communications Status | The current connection state between the Bed entity and the Bed Data Server application.
728 | Device Protocol Mode | The connectivity protocol supported by the bed communication controller.
729 | Service Message | The current status of the service required monitor system.
730 | Nurse Call Status | The current status for the nurse call switch.
731 | Bed Head Angle | The articulated angle of the bed’s head section in degrees. -360.0 -- 360.0
732 | Current Patient Weight | The last patient weight captured by the dynamic weight scale system displayed in kilograms. -9999.9 -- 9999.9
733 | Bed Brake Status | The current status for the brake switch.
734 | SmartBed Working Mode | The current configuration mode of the patient positioning monitor (PPM) alarm system.
736 | SmartBed General Alarm | The current status of the patient positioning monitor (PPM) alarm system.
737 | SmartBed Patient Position State | The current level of patient activity as detected by the patient positioning monitor (PPM) alarm system.
739 | Heart Rate | This variable is associated with a SourceID identifier. For Expression, the SourceID is always ""MEAS_UNKNOWN" if using Expression Plus protocol.
740 | NIBP Systolic | Possible units: mmHg or kPa.
741 |  | If PCMCIA protocol is used, the unit code of this variable is always 0 (unit unknown). A MeasurementTime attribute is also associated with this variable that is generated by the DDI which corresponds to the Smartlinx Server time when a new measurement is detected (i.e. either systolic, diastolic, or mean value has changed). When two consecutive NIBP measurements on the device yield identical values the DDI is not able to detect that a new measurement occurred and the MeasurementTime attribute value of this variable is not updated.
742 |  | If Expression Plus protocol is used, this variable is associated with a MeasurementTime attribute, which is transmitted by the device.
743 | NIBP Diastolic | Possible units: mmHg or kPa.
744 |  | If PCMCIA protocol is used, the unit code of this variable is always 0 (unit unknown). A MeasurementTime attribute is also associated with this variable that is generated by the DDI which corresponds to the Smartlinx Server time when a new measurement is detected (i.e. either systolic, diastolic, or mean value has changed). When two consecutive NIBP measurements on the device yield identical values the DDI is not able to detect that a new measurement occurred and the MeasurementTime attribute value of this variable is not updated.
745 |  | If Expression Plus protocol is used, this variable is associated with a MeasurementTime attribute, which is transmitted by the device.
746 | NIBP Mean | Possible units: mmHg or kPa.
747 |  | If PCMCIA protocol is used, the unit code of this variable is always 0 (unit unknown). A MeasurementTime attribute is also associated with this variable that is generated by the DDI which corresponds to the Smartlinx Server time when a new measurement is detected (i.e. either systolic, diastolic, or mean value has changed). When two consecutive NIBP measurements on the device yield identical values the DDI is not able to detect that a new measurement occurred and the MeasurementTime attribute value of this variable is not updated.
748 |  | If Expression Plus protocol is used, this variable is associated with a MeasurementTime attribute, which is transmitted by the device.
749 | INVP1 Systolic | Possible units: mmHg or kPa. If PCMCIA protocol is used, the unit code of this variable is always 0 (unit unknown).
750 | INVP1 Diastolic | Possible units: mmHg or kPa. If PCMCIA protocol is used, the unit code of this variable is always 0 (unit unknown).
751 | INVP1 Mean | Possible units: mmHg or kPa. If PCMCIA protocol is used, the unit code of this variable is always 0 (unit unknown).
752 | INVP2 Systolic | Possible units: mmHg or kPa. If PCMCIA protocol is used, the unit code of this variable is always 0 (unit unknown).
753 | INVP2 Diastolic | Possible units: mmHg or kPa. If PCMCIA protocol is used, the unit code of this variable is always 0 (unit unknown).
754 | INVP2 Mean | Possible units: mmHg or kPa. If PCMCIA protocol is used, the unit code of this variable is always 0 (unit unknown).
755 | SpO2
756 | End Tidal CO2 (EtCO2) | Possible units: mmHg or kPa. If PCMCIA protocol is used, the unit code of this variable is always 0 (unit unknown).
757 | Inspired CO2 (InCO2) | Possible units: mmHg or kPa. If PCMCIA protocol is used, the unit code of this variable is always 0 (unit unknown).
758 | Respiration Rate
759 | Device ID | 1–8 characters length alphanumeric string. For MR Patient Care Portal 5000, the length can be up to 31 characters.
760 | Inspired O2 (FiO2)
761 | Inspired N2O
762 | End Tidal N2O
763 | Patient ID | 1–16 characters length alphanumeric string.
764 | Device Date/Time | Format: YYYY-MM-DD HH:MM:SS
765 | Inspired Agent Concentration | This Variable is associated with a SourceID identifier.
766 | Expired Agent Concentration | This Variable is associated with a SourceID identifier.
767 | Patient Temperature | Possible units: °C or °F.
768 | Inspiratory 2nd Agent | This Variable is associated with a SourceID identifier.
769 | Expiratory 2nd Agent | This Variable is associated with a SourceID identifier.
770 | Device Configuration | For CritiCool: This variable contains information on the fixed step rate at which the patient gradually warms up. For Allon: This variable contains information on if the temperature alarms are on or off.Consult detailed help for more information on this variable
771 | Device Date/Time | Format: YYYYMMDDHHMMSS
772 | Device Status | This variable indicated if the water flow is through the patient wrap (garment) or is circulating internally (bypass). Possible values: 0 = Bypass 1 = Garment
773 | Device Software Revision | Format: XY.ZZ
774 | Patient Temperature | This variable is associated with SourceID identifier.
775 | Patient Temperature Setting
776 | System Pressure High Alarm | Possible values: 0 = Active 1 = Not Active
777 | Low Water Level Alarm | Low water level alarm Possible values: 0 = On 1 = Off
778 | Device Operation Mode | Consult detailed help for more information on this variable
779 | General Temperature Alarm | Consult detailed help for more information on this variable
781 | NIBP Systolic
782 | NIBP Diastolic
783 | NIBP Mean
784 | Patient Delta Temperature
785 | SpO2
786 | NIBP Pulse Rate
787 | ECG Heart Rate
788 | Respiration Rate
789 | Device Inactivation Status | Possible values: "enabled" and "audio-paused". See "Multiset Alarms" on help file.
790 | Active Alarms | See "Multiset Alarms" on help file.
792 |  | Consult detailed help for more information on this variable
793 | Latched Alarms | See "Multiset Alarms" on help file.
795 |  | Consult detailed help for more information on this variable
796 | Inactive Alarms | See "Multiset Alarms" on help file.
798 |  | Consult detailed help for more information on this variable
799 | General ST Level 1 | Physiological identifier: NOM_ECG_AMPL_ST_I
800 | General ST Level 2 | Physiological identifier: NOM_ECG_AMPL_ST_II
801 | General ST Level 3 | Physiological identifier: NOM_ECG_AMPL_ST_III
802 | Temperature 3
803 | Temperature 4
804 | O2 Saturation (SO2)
805 | Pleth Waveform
806 | ECG Measurement Status | ECG Rhythm Status
808 |  | Consult detailed help for more information on this variable
809 | Patient Age | Patient age. The Monitor device calculates the age of the patient from the birth data manually entered on the Monitor.
810 | Patient Sex | Possible values: "unknown", "male", "female", and "unspecified".
811 | Generic O2 Saturation (1)
812 | Generic O2 Saturation (2)
813 | Cardiac Output
814 | Fetal Heart Rate (FHR) | This variable is associated with ChannelID identifier.
815 | Fetal Beat-to-Beat Rate (FHR) | This variable is associated with ChannelID identifier.
816 | TOCO
817 | ECG Lead Name | Physiological identifier: NOM_ECG_AMPL_ST
818 | Arterial Pressure (ART) Systolic
819 | Arterial Pressure (ART) Diastolic
820 | Arterial Pressure (ART) Mean
821 | Left Atrial Mean Pressure (LAP)
822 | Central Venous Pressure (CVP) | Available on IntelliVue devices with firmware revision G or later.
823 | Device Language | Consult detailed help for more information on this variable
824 | ECG Lead II
825 | Airway N2O Concentration | Measurement from VueLink
826 | Inspired Desflurane
827 | Inspired Sevoflurane
828 | Fresh Gas Agent Value
829 | Fresh Gas Halothane
830 | Fresh Gas Enflurane
831 | Fresh Gas Isoflurane
832 | Inspired Halothane
833 | Inspired Enflurane
834 | Inspired Isoflurane
835 | End Tidal CO2
836 | Airway Pressure | Measured in the inspiratory path.
837 | Airway N2O Wave
838 | Halothane Concentration Wave
839 | Enflurane Concentration Wave
840 | Isoflurane Concentration Wave
841 | Desflurane Concentration Wave
842 | Sevoflurane Concentration Wave
843 | Apnea Duration | This is the device's measurement with label="NLS_NOM_TIME_PD_APNEA" and physio_id="NOM_TIME_PD_APNEA".
844 | End Tidal CO2
845 | Cardiac Beat-to-Beat Rate | Cardiac Beat-to-Beat Rate (NOM_ECG_CARD_BEAT_RATE_BTB).
846 | Airway Delta O2 Concentration
847 | Cardiac Output Index
848 | Stroke Volume
849 | Stroke Volume Index
850 | Systemic Vascular Resistance
851 | Normalized Ionized Calcium Concentration (Ca)
852 | Potassium Ion Concentration (K+)
853 | QT Interval (QT)
854 | Device Specification ID | Variables 764, 765 and 4201 must be interpreted conjointly.
856 |  | Consult detailed help for more information on this variable
857 | Device Specification Type | Variables 764, 765 and 4201 must be interpreted conjointly.
859 |  | Consult detailed help for more information on this variable
860 | Pleth Wave (2)
861 | Total Number of Used Disposable
862 | Cumulated Syringe Infused Volume
863 | Current Therapy Time
864 | Manually Entered Hematocrit
865 | Infusion Drug Name
866 | Infusion Time Remaining
867 | Infusion Dose Rate Setting
868 | Volume To Be Infused Original Setting
869 | KVO Rate Setting
870 | Inspired Minute Volume
871 | I:E Ratio | This variable contains expiratory component of I:E Ratio assuming that the inspiratory component is "1".
872 | SpO2
873 | Patient Birth Date | Format: 14 ascii digits string of "YYYYMMDDhhmmss"
874 | SpO2 (Pre-Ductal)
875 | SpO2 (Post-Ductal)
876 | ECG Lead III
877 | Hematocrit
878 | Pulmonary Vascular Resistance
879 | Body Surface Area
880 | ECG Lead MCL1 | Physiological identifier: NOM_ECG_AMPL_ST_MCL
881 | ECG Lead V1 | Physiological identifier: NOM_ECG_AMPL_ST_V1
882 | ECG Lead V6 | Physiological identifier: NOM_ECG_AMPL_ST_V6
883 | ECG Lead aVL | Physiological identifier: NOM_ECG_AMPL_ST_AVL
884 | ECG Lead aVF | Physiological identifier: NOM_ECG_AMPL_ST_AVF
885 | ECG Lead aVR | Physiological identifier: NOM_ECG_AMPL_ST_AVR
886 | ECG Lead V2 | Physiological identifier: NOM_ECG_AMPL_ST_V2
887 | ECG Lead V3 | Physiological identifier: NOM_ECG_AMPL_ST_V3
888 | ECG Lead V4 | Physiological identifier: NOM_ECG_AMPL_ST_V4
889 | ECG Lead V5 | Physiological identifier: NOM_ECG_AMPL_ST_V5
890 | Patient Temperature | This variable is associated with a SourceID Identifier.
891 | Micro Environment Air Temperature
892 | Tidal Volume
893 | Static Resistance
894 | Tidal Volume Setting
895 | Pressure Sensitivity Setting
896 | Sigh Interval Setting
897 | Sigh Volume Setting
898 | Pressure Support Level Above PEEP Setting
899 | Spontaneous Minute Volume
900 | Negative Inspiratory Pressure
901 | Peak Spontaneous Flow
902 | Flow-by Base Flow Setting
903 | Flow Sensitivity Setting
904 | Expiratory Time
905 | Inspiratory Flow Setting
906 | Expiratory Flow Setting
907 | Inspiratory Flow Trigger Setting | Delivered by the Evita 2 VueLink Driver.
908 | APRV Low Pressure Setting
909 | APRV High Pressure Setting
910 | Apnea Time Setting
911 | Minute Volume Setting
912 | Positive End Expiratory Pressure (PEEP)
913 | Ventilator Work of Breathing
914 | Rapid Shallow Breathing Index
915 | Patient Work of Breathing
916 | Systemic Vascular Resistance Index
917 | Pulmonary Vascular Resistance Index
918 | Continuous Cardiac Output
919 | Continuous Cardiac Output Index
920 | Pump Infusion Status
921 | Hemoglobin (Hb) | This variable is associated with SourceID identifier.
922 | Pleth Wave (Pre-Ductal)
923 | Pleth Wave (Post-Ductal)
924 | Arterial Venous O2 Content Difference
925 | O2 Delivery (DO2)
926 | O2 Extraction Ratio (O2ER)
927 | O2 Delivery Index (DO2I)
928 | O2 Consumption Index (VO2I)
929 | Oxygenation Ratio (PaO2/FiO2) | PaO2 to FIO2 ratio expressed in mmHg to % ratio.
930 | Physiologic Dead Space
931 | Alveolar Ventilation Volume (VA)
932 | Urea Nitrogen Concentration (BUN)
933 | Creatinine Concentration (Crea) | Measured by the i-Stat module.
934 | Plasma Osmolarity (Osm)
935 | Urine Creatinine
936 | Urine Potassium (K+)
937 | Urine Osmolarity
938 | Blood Urea Nitrogen Creatinine Ratio (BUN/cr)
939 | Osmolar Clearance (COSM)
940 | Creatinine Clearance (CreaCl)
941 | Fractional Excretion of Sodium (FeNa)
942 | Urine Serum Creatinine Ratio
943 | Urine Potassium Excretion
944 | Urine Sodium Excretion
945 | Free Water Clearance (CH2O)
946 | Inspired N2O
947 | Premature Ventricular Contractions
948 | LAP Systolic
949 | LAP Diastolic
950 | TcGas Heating Power
951 | Abdominal Fetal Heart Rate (FHR)
952 | Abdominal Fetal Beat-to-Beat Rate (FHR)
953 | Patient Activity Level (Movement)
954 | Patient Activity Level (Wearable)
955 | Minimum Airway Pressure
956 | P0.1 Occlusion Pressure
957 | Patient CO2 Production | Volume of CO2 produced by respiration.
958 | Relative Dead Space | Relative Dead Space Volume.
959 | Procalcitonin (PCT)
960 | Fetal Heart Rate (USound) | This variable is associated with ChannelID identifier.
961 | Fetal Beat-to-Beat Rate (USound) | This variable is associated with ChannelID identifier.
962 | HFV Mean Pressure Setting
963 | HFV Tidal Volume Setting
964 | PAV-based Resistance Setting
965 | PAV-based Elastance Setting
966 | Abdominal TOCO
967 | Chest Movement Respiration Rate
968 | Patient Movement Respiration Rate
969 | Patient Movement Pulse Rate
970 | Respiration Signal Quality Index (Movement)
971 | Cardiac Power Index (CCO)
972 | Cardiac Power Output (CCO)
973 | Central Venous Pressure (Thermodilution)
974 | Mean Arterial Pressure (Thermodilution)
975 | Average Blood Temperature (Thermodilution)
976 | O2 Delivery DO2 (Thermodilution)
977 | O2 Consumption (Thermodilution)
978 | O2 Consumption Index (Thermodilution)
979 | Global End-Diastolic Volume (Thermodilution)
980 | Global End-Diastolic Volum Index (Thermodilution)
981 | Pulmonary Vascular Permeability Index (Thermodilution)
982 | Stroke Volume (Thermodilution)
983 | Stroke Volume Index (Thermodilution)
984 | O2 Consumption Estimate
985 | EEG1 Wave
986 | EEG2 Wave
987 | EEG3 Wave
988 | EEG4 Wave
989 | Spectral Edge Frequency
990 | Burst Suppression Ratio | This variable is associated with ChannelID identifier.
991 | Patient Energy Expenditure
992 | Patient Respiratory Quotient
993 | Spectral Edge Frequency (CH3)
994 | Relative Power In Delta Band (CH3)
995 | Relative Power In Theta Band (CH3)
996 | Relative Power In Alpha Band (CH3)
997 | Relative Power In Beta Band (CH3)
998 | Spectral Edge Frequency (CH4)
999 | Relative Power In Delta Band (CH4)
1000 | Relative Power In Theta Band (CH4)
1001 | Relative Power In Alpha (CH4)
1002 | Relative Power In Beta Band (CH4)
1003 | Device Unique ID
1004 | Patient Bed ID | This variable contains a printable string identifying the system location.
1005 | Patient ID
1006 | Patient Name
1007 | Patient Comment Text
1008 | Patient Stress Index
1009 | Generic Patient Temperature
1010 | Breath Cycle Type
1011 | Airway Gas Flow Rate | Measured in the inspiratory path.
1012 | Inspired Minute Volume
1013 | Positive End Expiratory Pressure (PEEP) Setting
1014 | Spectral Edge Frequency - 1st EEG
1015 | Bispectral Index | This variable is associated with ChannelID identifier.
1016 | Total Power (Ch1)
1017 | Spike Rate | This variable is associated with ChannelID identifier.
1018 | Spectral Edge Frequency - 2nd EEG
1019 | Total Power (Ch2)
1020 | Narcotrend Index
1021 | Sharp Transient Intensity
1022 | Total Power Relative
1023 | Device Hardware Revision | This variable is obsolete because it may contains incorrect information. Please use variables 764, 765 and 4201 must be used instead.
1024 | Serial Number | This variable is obsolete because it may contains incorrect information. Please use variables 764, 765 and 4201 must be used instead.
1025 | O2 Flush Percentage Setting
1026 | CO2 Tidal Production
1027 | Inspiratory Resistance
1028 | Device Date/Time | Format: 16 ascii digits string of "YYYYMMDDhhmmssff", where the "ff" is second_fractions that is not used by IntelliVue monitors.
1029 | Arterial BP Systolic
1030 | Arterial BP Diastolic
1031 | Arterial BP Mean
1032 | Pulmonary Artery BP - Systolic
1033 | Pulmonary Artery BP - Diastolic
1034 | Pulmonary Artery BP Mean
1035 | Right Atrial Pressure (RAP) | Available on IntelliVue devices with firmware revision G or later.
1036 | Left Atrial Pressure (LAP) | Available on IntelliVue devices with firmware revision G or later.
1037 | Intra Cranial Pressure (ICP) | Available on IntelliVue devices with firmware revision G or later.
1038 | Pulmonary Artery Wedge Pressure
1039 | Generic NIBP Systolic Pressure | This variable is same as variable 2, except that the early estimated values are marked as NonSignificant in this variable.
1040 | Generic NIBP Diastolic Pressure | This variable is same as variable 3, except that the early estimated values are marked as NonSignificant in this variable.
1041 | Generic NIBP Mean Pressure | This variable is same as variable 4, except that the early estimated values are marked as NonSignificant in this variable.
1042 | Inspired Tidal Volume
1043 | Expired Tidal Volume
1044 | Expired Minute Volume
1045 | Patient Blood Temperature
1046 | Injectate Temperature
1047 | Inspired O2
1048 | Inspired O2 Pressure | This variable is marked as obsolete because the label is incorrect. Please see variable 52523 for concentration and 52527 for partial pressure.
1049 | End Tidal O2
1050 | CO2 (Gas Analyzer)
1051 | End Tidal N2O
1052 | End Tidal Agent Concentration
1053 | End Tidal Desflurane
1054 | End Tidal Sevoflurane
1055 | End Tidal Halothane
1056 | End Tidal Enflurane
1057 | End Tidal Isoflurane
1058 | Fresh Gas N2O Flow | Mixer parameter. Setting of fresh gas N2O flow.
1059 | Fresh Gas Air Flow
1060 | Transcutaneous O2 Pressure
1061 | Transcutaneous CO2 Pressure
1062 | Arterial O2 Saturation (SaO2)
1063 | EEG Total Power
1064 | Bispectral Index (BIS)
1065 | EEG Burst Suppression Ratio
1066 | Spontaneous Alveolar Minute Volume
1067 | Airway Dead Space
1068 | Expired Spontaneous Tidal Volume
1069 | Spontaneous Expired Minute Volume
1070 | Total Alveolar Tidal Volume
1071 | Respiration Waveform
1072 | EEG Signal Quality Index
1073 | Wearable Respiration Rate
1074 | Wearable Pulse Rate
1075 | Wearable Patient Temperature
1076 | Respiration Signal Quality Index (Wearable)
1077 | Pulse Rate Signal Quality Index (Wearable)
1078 | Airway Pressure Wave
1079 | Citrate Blood Concentration Setting
1080 | Neutrophil Count
1081 | Fresh Gas Total Flow Setting | Mixer parameter.
1082 | I:E Ratio Setting | This variable is associated with ChannelID identifier. This variable is calculated from variable 2773, which is sent as a string in the format "1:xx.xx".
1083 | Patient Admission Status | Possible values: "empty", "pre-admitted", "admitted", and "discharged".
1084 | ECG Lead V | Physiological identifier: NOM_ECG_AMPL_ST_V
1085 | Plateau Pressure | This variable contains parameter from Philips label "NLS_NOM_VENT_PRESS_RESP_PLAT" and observed value "NLS_NOM_VENT_PRESS_RESP_PLAT".
1086 | I:E Ratio Setting
1087 | Total PEEP
1088 | APRV Low Time (Tlow)
1089 | System General Status | This variable identifies the current operating mode of the device. Possible values: "not specified", "monitoring", "demo", and "service". Moreover, "- standby" or "- config" can be appended
1090 | Frequency | This variable describes the frequency of the main power supply. Possible values: "unspecified", "50", and "60".
1091 | Device Status | Possible values: "disconnected", "unassociated", "operating".
1092 | Device Software Revision | This variable is obsolete because it may contains incorrect information. Please use variables 764, 765 and 4201 must be used instead.
1093 | Maximum Apnea Time Setting
1094 | Inspired Agent Concentration
1095 | Patient Weight
1096 | Patient First Name
1097 | Patient Temperature
1098 | Ejection Fraction
1099 | End Diastolic Volume
1100 | End Diastolic Volume Index
1101 | End Systolic Volume
1102 | End Systolic Volume Index
1103 | EEG Electromyography (EMG) | This variable is associated with ChannelID identifier.
1104 | Generic EEG Wave
1105 | ECG Pad Wave
1106 | ECG Paddle Wave
1107 | Unspecific Pressure Pulse Rate
1108 | Patient Temperature 1
1109 | Patient Temperature 2
1110 | APRV High Pressure (Phigh)
1111 | APRV Low Pressure (Plow)
1112 | I:E Ratio | This variable is associated with ChannelID identifier This variable is calculated from variable 1190 and sent as a string in the format "1:xx.xx".
1113 | Plateau Airway Pressure
1114 | Plateau Pressure Setting
1115 | FiO2 Upper Alarm Limit
1116 | FiO2 Lower Alarm Limit
1117 | Device Altitude
1118 | Blood Flow Rate Setting
1119 | Dialysate Flow Rate Setting
1120 | Effluent Flow Rate Setting
1121 | Access Pod Pressure
1122 | Return Pod Pressure
1123 | Filter Pod Pressure
1124 | Effluent Pod Pressure
1125 | Transmembrane Pressure
1126 | Hb Calculation Offset
1127 | Patient Time Since Last Ward Admission
1128 | Patient Total Admission Time
1129 | Patient Drainage Volume
1130 | Fluid Removal Actual Rate
1131 | Sum of Minimum Alveolar Concentration (MAC)
1132 | HFV Frequency Setting
1133 | Ventilator Base Flow Setting
1134 | Blood Pressure Difference
1135 | Generic Blood Pressure Difference 1
1136 | Generic Blood Pressure Difference 2
1137 | Patient Height
1138 | Patient Temperature Setting
1139 | Arterial O2 Content (CaO2)
1140 | Alveolar Arterial O2 Difference (AaDO2)
1141 | Venous O2 Content (CvO2)
1142 | Ionized Calcium Concentration (Ca)
1143 | Chloride Concentration (Cl-)
1144 | Activated Clotting Time | Measured by the i-Stat module.
1145 | Generic Temperature
1146 | Generic Temperature 2
1147 | Generic Temperature 3
1148 | Generic Temperature 4
1149 | Expiratory Resistance
1150 | TcGas Site Timer
1151 | O2 Saturation (SO2)
1152 | TcGas Transducer Temperature
1153 | Peak Expiratory Flow Setting
1154 | Auto Flow Active Status
1155 | Inspiratory Minimum CO2
1156 | Expired Tidal Volume Indexed by Body Weight
1157 | ATC Activation Status
1158 | Skin Temperature
1159 | Rectal Temperature
1160 | Nasopharyngeal Temperature
1161 | Esophageal Temperature
1162 | Core Temperature
1163 | Arterial Temperature
1164 | Venous Temperature
1165 | Generic Core Temperature 1
1166 | Generic Core Temperature 2
1167 | Device Number | Device Model Number (example: "M8000")
1168 | Drug Delivery Rate
1169 | Drug Concentration
1170 | Manufacturers Name | "Philips"
1171 | Pressure Alarm Delay
1172 | Arterial Blood Pressure (ART) | Available on IntelliVue devices with firmware revision G or later. Arterial Blood Pressure (NOM_PRESS_BLD_ART).
1173 | Pulmonary Artery Blood Pressure (PA) | Available on IntelliVue devices with firmware revision G or later.
1174 | Arterial/Main Pump Flow
1175 | Slave Pump Flow
1176 | Suction Pump Flow
1177 | Auxiliary Pump Flow
1178 | Auxiliary Pump Time Since Start of Current Plegia Session
1179 | Auxiliary Pump Time Since Stop of Plegia Session
1180 | Auxiliary Pump Delivered Volume in Current Plegia Session
1181 | Auxiliary Pump Total Volume
1182 | Auxiliary Pump Total Plegia Time
1183 | Cardioplegia Main Pump Flow
1184 | Cardioplegia Main Pump Delivered Volume of Current Plegia Session
1185 | Cardioplegia Main Pump Total Volume
1186 | Cardioplegia Slave Pump Flow
1187 | Cardioplegia Slave Pump Delivered Volume in Current Plegia Session
1188 | Cardioplegia Slave Pump Total Volume
1189 | Cycle Flow Sensitivity Setting
1190 | Hydrocarbon Plasma Concentration (HCO3)
1191 | Corpuscular Mean Hemoglobin Concentration
1192 | Arterial O2 Partial Pressure
1193 | Venous O2 Partial Pressure
1194 | Arterial CO2 Partial Pressure
1195 | Arterial O2 Saturation (SaO2) | Obsolete variable that may be removed in a future release. The label of this variable is not correct. Please use variable 14 instead.
1196 | Hematocrit (HCT) | Post-Dialysis Hematocrit
1197 | Respiratory Rate (Flow)
1198 | Respiratory Rate (CO2)
1199 | Patient Type | Possible values: "unspecified", "adult", "pediatric", and "neonatal".
1200 | Patient Pace Maker Status | Possible values: "not paced", "paced gen".
1201 | Heart Rate Signal Quality
1202 | Tracheal Pressure
1203 | Ventilator Leakage
1204 | Plateau Time Setting | This variable is obsolete that may be removed in a future release. The label of this variable is incorrect, please use variable 9585 instead.
1205 | Warmer Mattress Temperature
1206 | Spontaneous Fraction of Minute Volume
1207 | Breathing System Leakage
1208 | Inspired MAC Concentration
1209 | Inspiratory 2nd Agent
1210 | Expiratory 2nd Agent | This is the device's measurement with label="NLS_NOM_EMFC_etAGTs" and physio_id="NOM_CONC_AWAY_AGENT_ET".
1211 | Volume Assist Level Setting
1212 | Backup Tidal Volume Setting
1213 | Flow Acceleration Setting
1214 | Cumulated Halothane Consumption
1215 | Cumulated Enflurane Consumption
1216 | Cumulated Isoflurane Consumption
1217 | Cumulated Desflurane Consumption
1218 | Cumulated Sevoflurane Consumption
1219 | Air Humidity Setting
1220 | Mattress Temperature Setting
1221 | Inspired Agent Setting
1222 | Expired Agent Setting | End-Tidal agent concentration
1223 | Inspired Desflurane Setting
1224 | Expired Desflurane Setting | End-Tidal desflurane concentration
1225 | Inspired Sevoflurane Setting
1226 | Expired Sevoflurane Setting | End-Tidal sevoflurane concentration
1227 | Fresh Gas Sevoflurane
1228 | Fresh Gas Desflurane
1229 | Ambient Temperature
1230 | Ambient Pressure
1231 | Pulse Rate ( Unspecified Source )
1232 | Central Venous Pressure Diastolic
1233 | Central Venous Mean Pressure
1234 | Central Venous Pressure Systolic
1235 | Expiratory Volume Wave
1236 | Apnea Ventilation I:E Ratio Setting (Number)
1237 | Dynamic Plateau Airway Pressure
1238 | Transcutaneous Relative Heater Power
1239 | Spontaneous Respiratory Rate Setting
1240 | Product Description | Variables 764, 765 and 4201 must be interpreted conjointly.
1242 |  | Consult detailed help for more information on this variable
1243 | Replacement Cumulated Volume
1244 | Dialysate Cumulated Volume | This variable is associated with ChannelID identifier.
1245 | Effluent Cumulated Volume | This variable is associated with ChannelID identifier.
1246 | Fluid Removal Cumulated Volume
1247 | Base Pressure
1248 | PEEP/CPAP Setting
1249 | Sigh Rate Setting
1250 | Sigh Tidal Volume Setting
1251 | Data Averaging Interval Setting
1252 | Chart Interval Setting
1253 | Static Lung Compliance
1254 | Static Lung Resistance
1255 | Dynamic Lung Compliance
1256 | Dynamic Lung Resistance | NLS_NOM_RES_AWAY_DYN
1257 | Apnea Inspiratory Pressure Setting
1258 | Patient Hose Compliance
1259 | Oximeter Averaging Interval Setting
1260 | System Alarm Status | Overall alert state of the system.
1262 |  | Consult detailed help for more information on this variable
1263 | Patient Alarm Priority | Overall Patient alarm priority. This is the current highest priority value of currently active Patient Alarms.
1265 |  | Consult detailed help for more information on this variable
1266 | Technical Alarm Priority | Overall Technical alarm priority. This is the current highest priority value of currently active Technical Alarms.
1268 |  | Consult detailed help for more information on this variable
1269 | Audio Alarm Priority | Overall audible alarm priority. This is the current highest priority value of currently active audible alarms.
1271 |  | Consult detailed help for more information on this variable
1272 | Technical Alarm Source ID
1273 | Technical Alarm Global State | Consult detailed help for more information on this variable
1274 | Technical Alarm Priority | Consult detailed help for more information on this variable
1275 | Technical Alarm Flag | Consult detailed help for more information on this variable
1276 | Technical Alarm Priority
1277 | Technical Alarm Description
1278 | Patient Alarm Source ID
1279 | Patient Alarm Global State | Consult detailed help for more information on this variable
1280 | Patient Alarm Priority | Consult detailed help for more information on this variable
1281 | Patient Alarm Flag | Consult detailed help for more information on this variable
1282 | Patient Alarm Priority
1283 | Patient Alarm Description
1284 | Device Nomenclature Revision | Format: major.minor (such as 1.0)
1285 | Device Text Catalog Revision | Format: "MM.NN mm.nn", where "MM.NN" is text catalog revision and "mm.nn" is language revision.
1286 | Device Data Format | Always 11 (Unicode).
1287 | Device Operating Area | Possible values: "unspecified", "operating room", "intensive care unit", "neonatal intensive care unit", and "cardiology care unit".
1288 | Patient O2 Consumption (VO2)
1289 | Patient CO2 Production (VCO2)
1290 | Arterial O2 Saturation (Left)
1291 | Arterial O2 Saturation (Right)
1292 | SpO2 Difference
1293 | Perfusion Indicator
1294 | Relative Perfusion (Left)
1295 | Relative Perfusion (Right)
1296 | Ao Systolic
1297 | Ao Diastolic
1298 | Ao Mean
1299 | RAP Systolic
1300 | RAP Diastolic
1301 | RAP Mean
1302 | Intracranial Pressure ICP Systolic
1303 | Intracranial Pressure ICP Diastolic
1304 | ICP Mean
1305 | UAP Systolic
1306 | UAP Diastolic
1307 | UAP Mean
1308 | UVP Systolic
1309 | UVP Diastolic
1310 | UVP Mean
1311 | Unspecific Pressure Systolic
1312 | Unspecific Pressure Diastolic
1313 | Unspecific Pressure Mean
1314 | Cerebral Perfusion Pressure (CPP)
1315 | Stroke Volume Variation
1316 | Index of Left Ventricular Contractility | This variable contains measurement from label "NLS_NOM_GRAD_PRESS_BLD_AORT_POS_MAX" and observed value "NOM_GRAD_PRESS_BLD_AORT_POS_MAX".
1317 | Aortic Blood Pressure (Ao) Wave
1318 | Auxiliary Waveform
1319 | Intermittent Mandatory Ventilation Status
1320 | Left Cardiac Work
1321 | Right Cardiac Work
1322 | Serum Creatinine
1323 | Urine Plasma Osmolarity Ratio
1324 | Urine Sodium Potassium Ratio
1325 | Urine Volume
1326 | Left Cardiac Work Index
1327 | Right Cardiac Work Index
1328 | Light Intensity (SvO2 Module)
1329 | SpO2 SvO2 Difference
1330 | Normalized Post Replacement Volume
1331 | Normalized Effluent Volume | This variable is associated with ChannelID identifier.
1332 | Normalized Total Replacement Volume (PBP) | This variable is associated with ChannelID identifier.
1333 | Normalized Replacement Volume (Pre)
1334 | Normalized Ultrafiltration Volume | This variable is associated with ChannelID identifier.
1335 | Patient Fluid Balance Limit Setting
1336 | Total Replacement Volume (PBP) | This variable is associated with ChannelID identifier.
1337 | Cumulated Liquid Volume (Pre) | This variable is associated with ChannelID identifier.
1338 | Filtration Fraction Setting
1339 | Diff.-Temperature between Tblood and Tskin
1340 | Diff.-Temperature between Tblood and Trect
1341 | Diff.-Temperature between Tblood and Tnaso
1342 | Diff.-Temperature between Tblood and Tesop
1343 | Diff.-Temperature between Tblood and Tcore
1344 | Diff.-Temperature between Tblood and Tart
1345 | Diff.-Temperature between Tblood and Tven
1346 | Diff.-Temperature between Tskin and Trect
1347 | Diff.-Temperature between Tskin and Tnaso
1348 | Diff.-Temperature between Tskin and Tesop
1349 | Diff.-Temperature between Tskin and Tcore
1350 | Diff.-Temperature between Tskin and Tart
1351 | Diff.-Temperature between Tskin and Tven
1352 | Diff.-Temperature between Trect and Tnaso
1353 | Diff.-Temperature between Trect and Tesop
1354 | Diff.-Temperature between Trect and Tcore
1355 | Diff.-Temperature between Trect and Tart
1356 | Diff.-Temperature between Trect and Tven
1357 | Diff.-Temperature between Tnaso and Tesop
1358 | Diff.-Temperature between Tnaso and Tcore
1359 | Diff.-Temperature between Tnaso and Tart
1360 | Diff.-Temperature between Tnaso and Tven
1361 | Diff.-Temperature between Tesop and Tcore
1362 | Diff.-Temperature between Tesop and Tart
1363 | Diff.-Temperature between Tesop and Tven
1364 | Diff.-Temperature between Tcore and Tart
1365 | Diff.-Temperature between Tcore and Tven
1366 | Diff.-Temperature between Tart and Tven
1367 | Newborn Body Length
1368 | Newborn Gestational Age
1369 | Calories
1370 | Input/Output Totals
1371 | Newborn Apgar Scores
1372 | BSA Formula Boyd
1373 | BSA Formula Dubois
1374 | Pupil Reaction (Left Eye) | Light reaction of left eye's pupil.
1375 | Pupil Reaction (Right Eye) | Light reaction of right eye's pupil.
1376 | Pleth Gain
1377 | Airway Temperature
1378 | Ventilation Perfusion Index
1379 | O2 Extraction Index
1380 | Duration Above Base Pressure
1381 | Inspiration Rise Time
1382 | Continuous Positive Airway Pressure (CPAP)
1383 | HFV Amplitude
1384 | Urine Urea
1385 | Urine pH Value
1386 | Total of CO2 (tCO2) | Result of Blood gas Analysis.
1387 | Total Bilirubin
1388 | Serum Glucose
1389 | Urine Glucose
1390 | Direct Bilirubin
1391 | Serum Calcium
1392 | Total Serum Calcium
1393 | Serum Magnesium
1394 | Serum Phosphat
1395 | Serum Kalium (Potassium)
1396 | Serum Clorid
1397 | Serum Albumine
1398 | Urine Clorid
1399 | Serum Globulin
1400 | Serum Protein
1401 | Serum Urea
1402 | Partial Thromboplastin Time
1403 | Prothrombin Time
1404 | Thrombin Time
1405 | Thromboplastine Time (Quick Time)
1406 | Alkalische Phosphatase
1407 | Alpha Amylase
1408 | Cholesterinesterase
1409 | Creatinin Kinase
1410 | Creatine Kinase of Type Muscle-Brain
1411 | Creatine Kinase of Type Muscle
1412 | Gamma Glutamyl Transpeptidase
1413 | Glutamic Oxaloacetic Transaminase
1414 | Glutamic Pyruvic Transaminase
1415 | Cholesterin
1416 | Triglyzeride
1417 | Urine Protein | (Total) Protein in Urine
1418 | Urine Calcium
1419 | Carboxy Hemoglobin
1420 | Total Protein
1421 | Lactate Dehydrogenase
1422 | Aspartin Aminotransferase
1423 | Alveolarproteinose (Alkaline Phosphatase)
1424 | Reticulocyte Count
1425 | Coagulation Time
1426 | Erithrocyte Sedimentation Rate
1427 | Packed Cell Volume
1428 | Kaolin Cephalin Time
1429 | Minimum Pressure During Breathing Cycle
1430 | Infusion Pump Accumulated Volume
1431 | Transthoracic Impedance
1432 | Inspired Expired N2O Difference
1433 | Inspired Expired Halothane Difference
1434 | Inspired Expired Enfluran Difference
1435 | Inspired Expired Isofluran Difference
1436 | Inspired Expired Sevofluran Difference
1437 | Inspired Expired Desfluran Difference
1438 | Inspired Expired Agent Difference
1439 | Inspired Expired O2 Difference
1440 | Patient Circuit O2 | O2 measured in the Patient Circuit
1441 | Barometric Pressure in Sample Chamber | This variable is obsolete that may be removed in a future release. It is duplicated with variable 4082.
1442 | Minute Volume
1443 | Mean Airway Pressure
1444 | Airway Respiration Rate
1445 | Inspiratory Airway Pressure Waveform
1446 | Inspiratory Flow Waveform
1447 | HFV Inspired Mandatory Minute Volume
1448 | HFV Gas Transport Coefficient Value
1449 | HFV Inspired Tidal Volume
1450 | HFV Tidal Volume
1451 | Generic Heart Rate | Heart Rate received from an external device
1452 | PERF from the 2nd SpO2/PLETH module
1453 | Percent Alpha (Left)
1454 | Percent Beta (Left)
1455 | Absolute Alpha (Left)
1456 | Absolute Beta (Left)
1457 | Absolute Delta (Left)
1458 | Absolute Theta (Left)
1459 | Mean Dominant Frequency (Left)
1460 | Median Power Frequency (Left)
1461 | Peak Power Frequency (Left)
1462 | Spectral Edge Frequency (Left)
1463 | Total Power (Left)
1464 | Scale of the Left Channel EEG Wave
1465 | Percent Alpha (Right)
1466 | Percent Beta (Right)
1467 | Absolute Alpha (Right)
1468 | Absolute Beta (Right)
1469 | Absolute Delta (Right)
1470 | Absolute Theta (Right)
1471 | Mean Dominant Frequency (Right)
1472 | Median Power Frequency (Right)
1473 | Peak Power Frequency (Right)
1474 | Spectral Edge Frequency (Right)
1475 | Total Power (Right)
1476 | Scale of the Right Channel EEG Wave
1477 | Duration of Positive Pressure
1478 | Fresh Gas N2O Concentration
1479 | Liquid Level (Anesthetic Agent Bottle)
1480 | Liquid Level (ISOflurane Bottle)
1481 | Liquid Level (ENFlurane Bottle)
1482 | Liquid Level (HALothane Bottle)
1483 | Liquid Level (DESflurane Bottle)
1484 | Liquid Level (SEVoflurane Bottle)
1485 | Invasive Blood Pressure Wave
1486 | Urine Shift Volume
1487 | Urine Flow Rate
1488 | Mean Dominant Frequency (CH1)
1489 | Mean Dominant Frequency (CH2)
1490 | Peak Power Frequency (CH1)
1491 | Peak Power Frequency (CH2)
1492 | % of Power in the Alpha Band (CH1)
1493 | % of Power in the Alpha Band (CH2)
1494 | % of Power in the Beta Band (1st EEG)
1495 | % of Power in the Beta Band (2nd EEG)
1496 | Hemoglobin
1497 | Patient Temperature
1498 | Capillaries pH Value
1499 | Adjusted pH (at Patient Temperature) | Adjusted pH at Patient Temperature
1500 | Adjusted Capillaries pH Value
1501 | Partial O2 in the Capillaries
1502 | Adjusted PO2 (at Patient Temperature) | Adjusted PO2 at Patient Temperature
1503 | Adjusted PcO2 (at Patient Temperature)
1504 | Venous CO2 Partial Pressure (PvCO2)
1505 | Partial CO2 in the Capillaries
1506 | Adjusted PcCO2 (at Patient Temperature)
1507 | Calculated Total CO2 (tCO2)
1508 | Calculated SO2 | Calculated SaO2
1509 | Calculated SvO2
1510 | Calculated ScO2
1511 | Calculated HCO3
1512 | Base Excess of Extra Cellular Fluid
1513 | Calculated Anion Gap
1514 | Urea | Measured by the i-Stat module.
1515 | Manually Entered FiO2 | In the Blood Review window.
1516 | Calculated Arterial Blood Base Excess
1517 | Ionized Magnesium
1518 | Desired Module Temperature | It's the selected/adjusted Temp in i-Stat Window.
1519 | Calculated BUN/Creatinine Ratio | Calculated by the i-Stat module.
1520 | Calculated Urea/Creatinine Ratio | Calculated by the i-Stat module.
1521 | Lactate | Measured value by the i-Stat module.
1522 | Device Simulator Status | Device data with the values: Passed and Failed.
1523 | Pressure of Sensor 1 | Pressure value of sensor 1. Device specific data.
1524 | Pressure of Sensor 2 | Pressure value of sensor 2. Device specific data.
1525 | Module Temperature
1526 | Air Temperature
1527 | Micro Environment Air Temperature Setting
1528 | Weight of the Urine Disposable Bag
1529 | Total Urine Volume of the Current Measurement Period
1530 | Urine Fluid Density
1531 | Urine Fluid Temperature
1532 | Ventilation Activated Status | Parameter to Informs whether the Patient is Ventilated
1533 | Simplified Acuity Physiologic Score
1534 | Value Normal | Missing parameters in SAPS parameter set which are considered in normal range.
1535 | Aspartate Aminotransferase
1536 | Spontaneous Respiration Rate
1537 | Current Fluid in Urine Bag
1538 | Cardiac Function Index (Thermodilution)
1539 | Intrathoracic Blood Volume (Thermodilution)
1540 | Intrathoracic Blood Volume Index (Thermodilution)
1541 | Extravascular Lung Water (Thermodilution)
1542 | Extravascular Lung Water Index (Thermodilution)
1543 | NMT TOF Count
1544 | NMT TOF Ratio
1545 | Twitch Height of the 1Hz/0.1Hz Stimulation Response
1546 | Post Tetatic Count
1547 | Remaining Time to Next Stimulation
1548 | TOF Response TOF1
1549 | TOF Response TOF2
1550 | TOF Response TOF3
1551 | TOF Response TOF4
1552 | Activated Partial Thromboplastin Time (aPTT Whole Blood Time)
1553 | Activated Partial Thromboplastin Time (aPTT Plasma Equivalent Time)
1554 | Prothrombin Time (Whole Blood Time)
1555 | Prothrombin Time (Plasma Equivalent Time)
1556 | Prothrombin Time Ratio
1557 | Spontaneous Inspiration Time
1558 | Overdistension Index
1559 | Lung Time Constant
1560 | Airway Correlation Coefficient
1561 | Rate Volume Ratio
1562 | Normalized Ionized Calcium (Ca)
1563 | Peak Inspiratory Flow Setting
1564 | Trigger Sensitivity Setting
1565 | Plateau Time Setting
1566 | Multiple Sigh Number Setting
1567 | Apnea Tidal Volume Setting
1568 | Apnea Respiration Rate Setting
1569 | Apnea Peak Flow Setting
1570 | Apnea Inspired O2 Setting
1571 | PSV Pressure Setting | Pressure Support Ventilation Setting
1572 | Sigh Activation Status
1573 | Suction O2 Concentration Setting
1574 | PCV Control Pressure Setting
1575 | PCV Inspiratory Time Setting
1576 | PCV Apnea Respiration Rate Setting
1577 | PCV Apnea Inspiratory Time Setting
1578 | PCV Apnea O2 Setting
1579 | PCV Apnea High Airway Pressure Setting
1580 | Inspiratory Component of I:E Ratio Setting | Pressure Ventilation I component of I:E Ratio setting
1581 | Expiratory Component of I:E Ratio Setting | Pressure Ventilation E component of I:E Ratio setting
1582 | Inspiratory Component of Apnea Ventilation I:E Ratio Setting
1583 | Expiratory Component of Apnea Ventilation I:E Ratio Setting
1584 | Setting: Pressure Ventilation Control
1585 | Cycle Time Setting
1586 | IPPV Ventilation Respiration Rate Setting
1587 | IMV Ventilation Respiration Rate Setting
1588 | Pressure Support PEEP Setting
1589 | Apnea Ventilation Delay
1590 | O2 Monitoring Mode
1591 | O2 Calibration Mode
1592 | Inspiratory Time Setting
1593 | Expiratory Time Setting
1594 | Alarm Percentage on Rise Time Setting
1595 | Continuous Positive Airway Pressure (CPAP) Setting | This is the device's measurement with label="NLS_NOM_SETT_CONC_AWAY_O2" and physio_id="NOM_CONC_AWAY_O2".
1596 | Ventilator Flow Setting
1597 | Low Inspiratory Pressure Setting
1598 | HFV Flow Setting
1599 | HFV Respiration Rate Setting
1600 | Mechanical Ventilation Activated Status | Controlled Mechanical Ventilation Setting
1601 | SIMV Mode | Synchronized Intermittent Mandatory Ventilation Setting
1602 | Mandatory Inspired Minute Volume Setting
1603 | Infusion Pump Delivery Rate Setting
1604 | Airway Respiration Rate Setting
1605 | Trigger Flow/Volume Setting | This is the device's measurement with label="NLS_NOM_SETT_VENT_VOL_LUNG_TRAPD" and physio_id="NOM_SETT_VENT_VOL_LUNG_TRAPD".
1606 | Apnea Alarm Delay
1607 | HFV Amplitude Setting
1608 | Volume Alarm Delay
1609 | Pressure Increase Rate Setting
1610 | Volume Warning Setting | Delivered by the Evita 2 VueLink Driver.
1611 | Airway Inspiratory Pressure Lower Alarm Limit
1612 | Apnea Alarm Delay Setting
1613 | Pressure Ramp Setting (ASB Ramp)
1614 | Applied Tidal Volume Setting
1615 | Assist Pressure Sensitivity Setting | Assist Sensitivity. Used by the Bear 1000 ventilator.
1616 | High Inspiratory Pressure Setting
1617 | Background Flow Setting
1618 | Vaporizer Agent Concentration Setting
1619 | Vaporizer Isoflurane Concentration Setting
1620 | Vaporizer Enflurane Concentration Setting
1621 | Vaporizer Halothane Concentration Setting
1622 | Vaporizer Desflurane Concentration Setting
1623 | Vaporizer Sevoflurane Concentration Setting
1624 | Gas Measurement Sensor Position
1625 | O2 Measurement Sensor Position
1626 | Inspired Tidal Volume Setting
1627 | Incubator/Warmer Temperature Setting
1628 | Preset Period of UrVol Setting
1629 | Low Time Setting | Setting variable for Evita 4 Airway Pressure Release Ventilation Mode.
1630 | High Time Setting | Setting variable for Evita 4 Airway Pressure Release Ventilation Mode.
1631 | Low Pressure Setting | Setting variable for Evita 4 Airway Pressure Release Ventilation Mode.
1632 | High Pressure Setting | Setting variable for Evita 4 Airway Pressure Release Ventilation Mode.
1633 | Volume Assist Level Setting
1634 | CPAP Flow Assist Level Setting
1635 | Preset Stimulation Current Setting
1636 | Preset Stimulation Charge Setting
1637 | Preset Stimulation Impulse Duration Setting
1638 | NMT TOF Repetition Time | Preset Train of Four (Slow TOF) Repetition Time Setting
1639 | Panting Limit Setting
1640 | Airway Inspiratory Pressure Upper Alarm Limit | Alarm Limit: High Pressure
1641 | Sigh Volume
1642 | Sustained Pressure Alarm Limit
1643 | Intrathoracic Blood Volume (ITBV)
1644 | Intrathoracic Blood Volume Index (ITBI)
1645 | Extravascular Lung Water
1646 | Extravascular Lung Water Index
1647 | Global End Diastolic Volume
1648 | Global End Diastolic Volume Index
1649 | Cardiac Function Index
1650 | Mixed Venous O2 Saturation
1651 | Generic Term for the Transcutaneous Gases
1652 | Intrinsic PEEP
1653 | Positive Peak Inspiratory Pressure
1654 | Dead Space to Tidal Volume Ratio
1655 | Lung Volume Trapped
1656 | Lung Vital Capacity
1657 | Generic Lung Compliance
1658 | EEG Electromyography
1659 | Suppression Ratio
1660 | EEG Peak Power Frequency
1661 | End Tidal N2
1662 | Inspired N2
1663 | Generic Agent Wave
1664 | End Tidal MAC Concentration
1665 | Left Ventricular Stroke Volume Index
1666 | Left Ventricular Stroke Volume
1667 | Right Ventricular Stroke Volume
1668 | Glasgow Coma Score
1669 | SubScore of the GCS - Eye Response
1670 | SubScore of the GCS - Motoric Response
1671 | SubScore of the GCS - Verbal Response
1672 | Head Circumference
1673 | pH in Arterial Blood
1674 | Arterial Blood Hemoglobin (HB)
1675 | pH in Venous Blood
1676 | Urine Natrium
1677 | Natrium in Serum
1678 | pH in the Blood Plasma
1679 | Blood Natrium (Sodium)
1680 | Kalium (Potassium)
1681 | Generic Glucose Parameter
1682 | Partial CO2
1683 | Partial O2
1684 | Base Excess of Blood
1685 | Percent Alveolarvenous Shunt Qs/Qt
1686 | Extravascular Thermal Volume Index
1687 | VPB Rate
1688 | ECG Lead V1 Wave
1689 | ECG Lead V2 Wave
1690 | ECG Lead V3 Wave
1691 | ECG Lead V4 Wave
1692 | ECG Lead V5 Wave
1693 | ECG Lead V6 Wave
1694 | ECG Lead I Wave
1695 | ECG Lead II Wave
1696 | ECG Lead III Wave
1697 | ECG Lead aVR Wave
1698 | ECG Lead aVL Wave
1699 | ECG Lead aVF Wave
1700 | Mean Dominant Frequency (CH3)
1701 | Mean Dominant Frequency (CH4)
1702 | NMT DBS Count
1703 | Peak Power Frequency (CH3)
1704 | Peak Power Frequency (CH4)
1705 | Pump Infusion Mode
1706 | Remaining Time Counter | Time to Elapse Counter used by the i-Stat module.
1707 | APRV Low Pressure Synchronisation Time Setting
1708 | APRV High Pressure Synchronisation Time Setting
1709 | Acoustic Respiration Rate
1710 | Regional O2 Saturation (rSO2 2)
1711 | Regional O2 Saturation (rSO2 3)
1712 | Regional O2 Saturation (rSO2 4)
1713 | Regional O2 Saturation (rSO2 5)
1714 | Regional O2 Saturation (rSO2 6)
1715 | O2 Delivery Index (Thermodilution)
1716 | Noninvasive Venous Hemoglobin (SpHbv)
1717 | Noninvasive Arterial Hemoglobin (SpHba)
1718 | Generic Pulse Rate
1719 | Corrected Total Hemoglobin
1720 | Corrected Arterial Total Hemoglobin
1721 | Corrected Venous Total Hemoglobin
1722 | Systemic Vascular Resistance (Thermodilution)
1723 | Systemic Vascular Resistance Index (Thermodilution)
1724 | Pulse Rate (Abdominal)
1725 | ECG Lead V3R
1726 | ECG Lead V4R
1727 | ECG Lead V5R
1728 | Inspired NO
1729 | ECG Lead V9
1730 | ECG Lead V7
1731 | ECG Lead V8
1732 | Apnea Ventilation I:E Ratio Setting | This variable is associated with ChannelID identifier. This variable is calculated from variable 4977 and 4978, which is sent as a string in the format "1:xx.xx".
1733 | Assisted Respiration Rate
1734 | Correlation Coefficient Squared
1735 | ECG Lead V6R
1736 | Citrate Solution Concentration
1737 | Calcium Solution Concentration
1738 | Citrate Acid Concentration
1739 | Citrate Load Setting
1740 | Backup Respiration Rate Setting
1741 | Inspired Expired Pressure Difference
1742 | Backup Inspired Expired Pressure Difference Setting
1743 | Femoral Arterial Pressure Mean
1744 | Femoral Arterial Pressure Systolic
1745 | Femoral Arterial Pressure Diastolic
1746 | Tidal Volume Lower Alarm Limit
1747 | Minute Volume Lower Alarm Limit
1748 | ECG Lead V7 Wave
1749 | ECG Lead V8 Wave
1750 | ECG Lead V9 Wave
1751 | Airway Aux Wave
1752 | Fetal DECG Wave
1753 | Patient BSA Formula | Consult detailed help for more information on this variable
1754 | Platelets (Thrombocyte Count)
1755 | Endotracheal Tube Type
1756 | Minute Volume Upper Alarm Limit
1757 | Pulse Rate from Plethysmogram
1758 | ECG Ectopic Status | Consult detailed help for more information on this variable
1759 | ST Lead ES
1760 | ST Lead AS
1761 | ST Lead AI
1762 | ST Index
1763 | Pulse Pressure
1764 | Mean Transit Time (MTt)
1765 | Alveolar Dead Space
1766 | Tympanic Temperature
1767 | Injectate Volume (CO)
1768 | SpO2 (Telemetry)
1769 | Pulse Rate (Telemetry)
1770 | Intracranial Pressure ICP1 (Systolic)
1771 | Intracranial Pressure ICP1 (Diastolic)
1772 | Intracranial Pressure ICP1 (Mean)
1773 | Intracranial Pressure ICP2 Systolic
1774 | Intracranial Pressure ICP2 Diastolic
1775 | Intracranial Pressure ICP2 Mean
1776 | Brachial Arterial Pressure (BA) Systolic
1777 | Brachial Arterial Pressure (BA) Diastolic
1778 | Brachial Arterial Pressure (BA) Mean
1779 | Vesical Temperature (Bladder)
1780 | Cerebral Temperature
1781 | ST Level for Lead I
1782 | ST Level for Lead II
1783 | ST Level for Lead III
1784 | ST Level for Lead aVR
1785 | ST Level for Lead aVL
1786 | ST Level for Lead aVF
1787 | ST Level for Lead V1
1788 | ST Level for Lead V2
1789 | ST Level for Lead V3
1790 | ST Level for Lead V4
1791 | ST Level for Lead V5
1792 | ST Level for Lead V6
1793 | Prothrombin Time (International Normalized Ratio)
1794 | PSV Tidal Volume | Tidal Volume in Pressure Support Ventilation Mode
1795 | Negative Inspiratory Force Index
1796 | A-Line ARX Index
1797 | Impedance Cardiography (ICG) Wave
1798 | Thoracic Fluid Content
1799 | Thoracic Fluid Content Index
1800 | Accelerated Cardiac Index
1801 | Heart Contractility Index
1802 | O2 Saturation (Left)
1803 | O2 Saturation (Right)
1804 | Pulse Pressure Variation
1805 | Downslope Time
1806 | Pulmonary Vascular Permeability Index | This is the device's measurement with label="NLS_NOM_PERM_VASC_PULM_INDEX" and physio_id="NOM_PERM_VASC_PULM_INDEX". See also variable 8791.
1807 | ART Pulse Rate | ART pulse rate
1808 | Umbilical Arterial Pressure Pulse Rate
1809 | Patient Visit Number | Patient encounter ID contains the ID of the current visit of the patient
1810 | Patient Height at Entered Time
1811 | Patient Weight at Entered Time
1812 | Patient Body Surface Area (BSA)
1813 | Inspired O2 Setting (FiO2)
1814 | Inspiratory Component of I:E Ratio Setting
1815 | Expiratory Component of I:E Ratio Setting
1816 | Lung Elastance
1817 | Backup Inspiratory Time Setting
1818 | Pressure Support Level Above Phigh
1819 | Total Respiration Duration Setting
1820 | Heater Power
1821 | Capillary Refill Time
1822 | Estimated Patient Circuit Leakage
1823 | Estimated Ventilation Total Leakage
1824 | Ventilation Mode Additional Information
1825 | Backup Plateau Time Setting
1826 | ECG Lead V3R Wave
1827 | ECG Lead V4R Wave
1828 | ECG Lead V5R Wave
1829 | ECG Lead V6R Wave
1830 | Signal Quality Index (BIS) | This variable is associated with ChannelID identifier.
1831 | Continuous NIBP Systolic
1832 | Continuous NIBP Diastolic
1833 | Continuous NIBP Mean
1834 | Next Intervention Time
1835 | Total Power (CH3)
1836 | Total Power (CH4)
1837 | Total Power (Left)
1838 | Total Power (Right)
1839 | PCV I:E Ratio Setting | This variable is associated with ChannelID identifier. This variable is calculated from variable 4975 and 4976, which is sent as a string in the format "1:xx.xx".
1840 | Inspiratory CO2 (FiCO2)
1841 | End Tidal NO2
1842 | Fresh Gas Total Flow
1843 | Bolus Volume Setting
1844 | Generic NIBP Pressure | Available on IntelliVue devices with firmware revision G or later.
1845 | Endotracheal Tube Diameter
1846 | Mandatory Respiration Rate
1847 | Mandatory Minute Volume
1848 | Mandatory Tidal Volume
1849 | N2O Consumption
1850 | Device Power Up Time | Number of seconds since the Monitor device's last power-on.
1851 | O2 Consumption
1852 | Air Consumption
1853 | Arterial Blood Pressure (ABP) | Available on IntelliVue devices with firmware revision G or later. Arterial Blood Pressure (NOM_PRESS_BLD_ART_ABP).
1854 | ABP Pulse Rate | ABP pulse rate
1855 | Signal Noise Ratio (CCO)
1856 | Spontaneous Airway Respiration Rate
1857 | Synchronized Respiration Rate
1858 | Mechanical Respiration Rate
1859 | Daily Urine Output
1860 | Activated Partial Thromboplastin Time Ratio
1861 | Global Ejection Fraction | This is the device's measurement with label="NLS_NOM_FRACT_EJECT" and physio_id="NOM_FRACT_EJECT". See also 8790.
1862 | Generic CO2 Parameter | This variable contains "PECO2, Partial expired CO2", which is measured from label="NLS_NOM_VENT_AWAY_CO2_EXP", physio_id="NOM_VENT_AWAY_CO2_EXP".
1863 | Central Venous O2 Saturation (ScvO2)
1864 | Suppression Ratio (BIS) | This variable is associated with ChannelID identifier.
1865 | Right Ventricular Stroke Work Index
1866 | Base Excess of Extracellular Fluid
1867 | Rise Time Setting
1868 | Time Duration at Low Pressure (Tlow)
1869 | Time Duration at High Pressure (Thigh)
1870 | Spontaneous Tidal Volume
1871 | Cardiac Troponin l
1872 | Cardiopulmonary Bypass Status
1873 | Cardiac Brain Natriuretic Peptide
1874 | Mean Dominant Frequency
1875 | Filtration Current Volume
1876 | Calcium Flow Rate Setting
1877 | Calcium Compensation Setting
1878 | Cerebral State Index (CSI)
1879 | EEG Amplitude | This variable is associated with ChannelID identifier.
1880 | Patient Middle Names | The Middle Name attribute contains the middle name of the patient.
1881 | Ao Pulse Rate
1882 | PAP Pulse Rate
1883 | CVP Pulse Rate
1884 | RAP Pulse Rate
1885 | LAP Pulse Rate
1886 | ICP Pulse Rate
1887 | Umbilical Venous Pressure Pulse Rate
1888 | FAP Pulse Rate
1889 | BAP Pulse Rate
1890 | IC1 Pulse Rate
1891 | Intracranial Pressure ICP2 Pulse Rate
1892 | Pulse Rate (P1)
1893 | Pulse Rate (P2)
1894 | Pulse Rate (P3)
1895 | Pulse Rate (P4)
1896 | Pulse Rate from Plethysmogram (Left)
1897 | Pleth Wave (Right)
1898 | Pleth Wave (Left)
1899 | Calculated SO2
1900 | Airway Respiration Rate (Ohmeda)
1901 | Airway Respiration Rate Setting (Ohmeda)
1902 | Fresh Gas O2 Concentration
1903 | Manually entered FIO2 or Flow | This variable contains measurement from label "NLS_NOM_CONC_AWAY_O2" and observed value "NOM_CONC_AWAY_O2".
1904 | Generic O2 Parameter | This is the device's measurement with label="NLS_NOM_SETT_CONC_AWAY_O2" and physio_id="NOM_CONC_AWAY_O2".
1905 | Fresh Gas O2 Flow Setting | Mixer parameter.
1906 | Airway Pressure Numeric Value (Ohmeda)
1907 | Inspiratory Peak Flow
1908 | Maximum Pressure During Breathing Cycle
1909 | Peak Pressure
1910 | Peak Inspiratory Flow Setting
1911 | Airway Maximum Pressure Setting
1912 | Peak Inspiratory Pressure Setting
1913 | Setting: Inspiratory Pressure (Ohmeda)
1914 | Mean Airway Pressure (Ohmeda)
1915 | EEG Wave (Left)
1916 | EEG Wave (Right)
1917 | % of Power in the Delta Band - 1st EEG
1918 | % of Power in the Delta Band - 2nd EEG
1919 | Percent Delta - Left Side
1920 | Percent Delta - Right Side
1921 | % of Power in the Theta Band (1st EEG)
1922 | % of Power in the Theta Band (2nd EEG)
1923 | Percent Theta - Left Side
1924 | Percent Theta - Right Side
1925 | Generic Calculation Constant | This variable contains measurement from Philips label "NLS_NOM_NUM_CALC_CONST" and observed value "NOM_METRIC_NOS". See also variable 8493.
1926 | Pressure Ventilation Control Pressure
1927 | Right Heart Ejection Fraction
1928 | White Blood Count (Leucocyte Count)
1929 | Red Blood Count (Erithrocyte Count)
1930 | Mean Corpuscular Volume
1931 | Corpuscular Mean Hemoglobin Content
1932 | Ferrum
1933 | Fetal Hemoglobin
1934 | MetHemoglobin
1935 | Airway Volume Wave
1936 | Expiratory Peak Flow
1937 | Expiratory Peak Flow
1938 | The V stands for Volume
1939 | The R stands for Rate
1940 | Adjusted PaCO2 (at Patient Temperature)
1941 | Incubator Heating Power | Power Required to set Air&Pat Temp in the Incubator
1942 | Actual Patient Age | Actual patient age, provided by a module, measured in years. Please note that this Patient Age will be available only when the corresponding module is connected to the Monitor. Variable 169 contains the patient age as provided by the Monitor itself (calculated from the birth date entered on the Monitor device.)
1943 | The P stands for Pressure
1944 | PCV Apnea Control Pressure Setting
1945 | Fresh Gas Air Flow Setting | Mixer parameter.
1946 | PEEP/CPAP Lower Alarm Limit
1947 | Tidal Volume Upper Alarm Limit
1948 | T-Tre Temperature difference between Temp and Trect
1949 | T-Tbl Temperature difference between Temp and Tblood
1950 | T-Tco Temperature difference between Temp and Tcore
1951 | T-Tsk Temperature difference between Temp and Tskin
1952 | T-Tes Temperature difference between Temp and Tesoph
1953 | T-Tna Temperature difference between Temp and Tnaso
1954 | T-Tar Temperature difference between Temp and Tart
1955 | T-Tve Temperature difference between Temp and Tven
1956 | T-Taw Temperature difference between Temp and Tairwy
1957 | Tre-T Temperature difference between Trect and Temp
1958 | Tre-bl Temperature difference between Trect and Tblood
1959 | Tre-sk Temperature difference between Trect and Tskin
1960 | Tre-aw Temperature difference between Trect and Tairwy
1961 | Temperature difference between Tblood and Temp
1962 | Tbl-aw Temperature difference between Tblood and Tairwy
1963 | Tco-T Temperature difference between Tcore and Temp
1964 | Tco-re Temperature difference between Tcore and Trect
1965 | Tco-bl Temperature difference between Tcore and Tblood
1966 | Tco-sk Temperature difference between Tcore and Tskin
1967 | Tco-es Temperature difference between Tcore and Tesoph
1968 | Tco-na Temperature difference between Tcore and Tnaso
1969 | Tco-aw Temperature difference between Tcore and Tairwy
1970 | Tsk-T Temperature difference between Tskin and Temp
1971 | Tsk-bl Temperature difference between Tskin and Tblood
1972 | Tsk-aw Temperature difference between Tskin and Tairwy
1973 | Tes-T Temperature difference between Tesoph and Temp
1974 | Tes-re Temperature difference between Tesoph and Trect
1975 | Tes-bl Temperature difference between Tesoph and Tblood
1976 | Tes-sk Temperature difference between Tesoph and Tskin
1977 | Tes-na Temperature difference between Tesoph and Tnaso
1978 | Tes-aw Temperature difference between Tesoph and Tairwy
1979 | Tna-T Temperature difference between Tnaso and Temp
1980 | Tna-re Temperature difference between Tnaso and Trect
1981 | Tna-bl Temperature difference between Tnaso and Tblood
1982 | Tna-sk Temperature difference between Tnaso and Tskin
1983 | Tna-aw Temperature difference between Tnaso and Tairwy
1984 | Tar-T Temperature difference between Tart and Temp
1985 | Tar-re Temperature difference between Tart and Trect
1986 | Tar-bl Temperature difference between Tart and Tblood
1987 | Tar-co Temperature difference between Tart and Tcore
1988 | Tar-sk Temperature difference between Tart and Tskin
1989 | Tar-es Temperature difference between Tart and Tesoph
1990 | Tar-na Temperature difference between Tart and Tnaso
1991 | Tar-aw Temperature difference between Tart and Tairwy
1992 | Tve-T Temperature difference between Tven and Temp
1993 | Tve-re Temperature difference between Tven and Trect
1994 | Tve-bl Temperature difference between Tven and Tblood
1995 | Tve-co Temperature difference between Tven and Tcore
1996 | Tve-sk Temperature difference between Tven and Tskin
1997 | Tve-es Temperature difference between Tven and Tesoph
1998 | Tve-na Temperature difference between Tven and Tnaso
1999 | MAC Concentration (Awake)
2000 | Tve-ar Temperature difference between Tven and Tart
2001 | Tve-aw Temperature difference between Tven and Tairwy
2002 | Taw-T Temperature difference between Tairwy and Temp
2003 | Taw-re Temperature difference between Tairwy and Trect
2004 | Taw-bl Temperature difference between Tairwy and Tblood
2005 | Taw-co Temperature difference between Tairwy and Tcore
2006 | Taw-sk Temperature difference between Tairwy and Tskin
2007 | Taw-es Temperature difference between Tairwy and Tesoph
2008 | Taw-na Temperature difference between Tairwy and Tnaso
2009 | Taw-ar Temperature difference between Tairwy and Tart
2010 | Taw-ve Temperature difference between Tairwy and Tven
2011 | Generic Agent 1 Wave
2012 | End Tidal Generic Agent 1
2013 | Inspired Generic Agent 1
2014 | Generic Agent 2 Wave
2015 | End Tidal Generic Agent 2
2016 | Inspired Generic Agent 2
2017 | Adjusted Arterial Blood pH Value
2018 | Adjusted Venous Blood pH Value
2019 | Adjusted PvCO2 (at Patient Temperature)
2020 | Adjusted PvO2 (at Patient Temperature)
2021 | Generic Calculation Constant | This variable contains parameter from Philips label "NLS_NOM_NUM_CATHETER_CONST" and observed value "NOM_NUM_CATHETER_CONST". See also variable 8395.
2022 | ECG Lead V Wave
2023 | ECG Lead MCL Wave
2024 | ECG Lead MCL1 Wave
2025 | MAC Concentration
2026 | Pleth Wave (Telemetry)
2027 | NIBP Systolic (Telemetry)
2028 | NIBP Diastolic (Telemetry)
2029 | NIBP Mean (Telemetry)
2030 | Generic Pressure 1 (P1) Systolic
2031 | Generic Pressure 1 (P1) Diastolic
2032 | Generic Pressure 1 (P1) Mean
2033 | Generic Pressure 2 (P2) Systolic
2034 | Generic Pressure 2 (P2) Diastolic
2035 | Generic Pressure 2 (P2) Mean
2036 | Generic Pressure 3 (P3) Systolic
2037 | Generic Pressure 3 (P3) Diastolic
2038 | Generic Pressure 3 (P3) Mean
2039 | Generic Pressure 4 (P4) Systolic
2040 | Generic Pressure 4 (P4) Diastolic
2041 | Generic Pressure 4 (P4) Mean
2042 | Incubator Humidity
2043 | Predictive Oral Temperature
2044 | Predictive Rectal Temperature
2045 | Predictive Axillary Temperature
2046 | Perfusion (Telemetry)
2047 | Right to Left Heart Shunt
2048 | Adjusted PcO2 (at Patient Temperature)
2049 | Adjusted Arterial PaO2 (at Patient Temperature)
2050 | Pulse Rate from Plethysmogram (Right)
2051 | Patient O2 Consumption per Body Surface
2052 | Patient O2 Consumption per Body Weight
2053 | Ideal Patient Weight
2054 | End Flow Setting
2055 | PSV Backup Delay Setting
2056 | Technical Alarm Internal Source ID | Consult detailed help for more information on this variable
2057 | Technical Alarm Source Type
2058 | Device Technical Alarm | Consult detailed help for more information on this variable
2059 | Patient Alarm Internal Source ID | Consult detailed help for more information on this variable
2060 | Patient Alarm Source Type
2061 | Patient Alarm | Consult detailed help for more information on this variable
2062 | Accoustic Respiration Rate Wave
2063 | Blood Temperature Wave
2064 | Non-specific Pressure 1 Wave
2065 | Generic Invasive Blood Pressure P1 | Available on IntelliVue devices.
2066 | Generic Invasive Blood Pressure P1 Systolic
2067 | Generic Invasive Blood Pressure P1 Diastolic
2068 | Generic Invasive Blood Pressure P1 Mean
2069 | Non-specific Pressure 2 Wave
2070 | Generic Invasive Blood Pressure P2 | Available on IntelliVue devices.
2071 | Generic Invasive Blood Pressure P2 Systolic
2072 | Generic Invasive Blood Pressure P2 Diastolic
2073 | Generic Invasive Blood Pressure P2 Mean
2074 | Non-specific Pressure 3 Wave
2075 | Generic Invasive Blood Pressure P3 | Available on IntelliVue devices.
2076 | Generic Invasive Blood Pressure P3 Systolic
2077 | Generic Invasive Blood Pressure P3 Diastolic
2078 | Generic Invasive Blood Pressure P3 Mean
2079 | Non-specific Pressure 4 Wave
2080 | Generic Invasive Blood Pressure P4 | Available on IntelliVue devices.
2081 | Generic Invasive Blood Pressure P4 Systolic
2082 | Generic Invasive Blood Pressure P4 Diastolic
2083 | Generic Invasive Blood Pressure P4 Mean
2084 | Non-specific Pressure 5 Wave
2085 | Generic Invasive Blood Pressure P5 | Available on IntelliVue devices.
2086 | Generic Invasive Blood Pressure P5 Systolic
2087 | Generic Invasive Blood Pressure P5 Diastolic
2088 | Generic Invasive Blood Pressure P5 Mean
2089 | Non-specific Pressure 6 Wave
2090 | Generic Invasive Blood Pressure P6 | Available on IntelliVue devices.
2091 | Generic Invasive Blood Pressure P6 Systolic
2092 | Generic Invasive Blood Pressure P6 Diastolic
2093 | Generic Invasive Blood Pressure P6 Mean
2094 | Non-specific Pressure 7 Wave
2095 | Generic Invasive Blood Pressure P7 | Available on IntelliVue devices.
2096 | Generic Invasive Blood Pressure P7 Systolic
2097 | Generic Invasive Blood Pressure P7 Diastolic
2098 | Generic Invasive Blood Pressure P7 Mean
2099 | Non-specific Pressure 8 Wave
2100 | Generic Invasive Blood Pressure P8 | Available on IntelliVue devices.
2101 | Generic Invasive Blood Pressure P8 Systolic
2102 | Generic Invasive Blood Pressure P8 Diastolic
2103 | Generic Invasive Blood Pressure P8 Mean
2104 | Intra-Uterine Pressure (IUP) Wave
2105 | Intra-Uterine Pressure
2106 | Intra-Uterine Pressure Systolic
2107 | Intra-Uterine Pressure Diastolic
2108 | Intra-Uterine Pressure Mean
2109 | Unspecified Blood Pressure
2110 | Unspecified Blood Pressure Systolic
2111 | Unspecified Blood Pressure Diastolic
2112 | Unspecified Blood Pressure Mean
2113 | Airway Flow Wave
2114 | Auxiliary Wave/Parameter
2115 | Global Ejection Fraction (2) | This is the device's measurement with label="NLS_NOM_EMFC_GEF" and physio_id="NOM_METRIC_NOS". See also 7810.
2116 | Pulmonary Vascular Permeability Index (2) | This is the device's measurement with label="NLS_NOM_EMFC_PVPI" and physio_id="NOM_METRIC_NOS". See also variable 6493.
2117 | Pressure Drop
2118 | Replacement Pre-delivered Rate Setting
2119 | Replacement Post-delivered Rate Setting
2120 | Pre-blood Flow Rate Setting
2121 | Patient Weight Balance Setting
2122 | Anticoagulant Bolus Interval Setting
2123 | Anticoagulant Infusion Rate Setting
2124 | Ultrafiltration Ratio Setting
2125 | Therapy Elapsed Time
2126 | Airway Volume
2127 | Processed Cumulated Blood Volume
2128 | Replacement Pre-delivered Cumulated Volume
2129 | Replacement Post-delivered Cumulated Volume | This variable is associated with ChannelID identifier.
2130 | Pre-blood Cumulated Volume | This variable is associated with ChannelID identifier.
2131 | Syringe Speed
2132 | Maximum Volume Gain Setting
2133 | Patient Fluid Balance
2134 | Arterial Blood Pressure (ABP) Wave
2135 | Arterial Blood Pressure (ART) Wave
2136 | Pulmonary Arterial Pressure (PAP) Wave
2137 | Central Venous Pressure (CVP) Wave
2138 | Right Atrial Pressure (RAP) Wave
2139 | Left Atrial Pressure (LAP) Wave
2140 | Intra-cranial Pressure (ICP) Wave
2141 | Umbilical Arterial Pressure (UAP) Wave
2142 | Umbilical Venous Pressure (UVP) Wave
2143 | Airway CO2 Level Waveform
2144 | Airway O2 Wave
2145 | Airway N2 Wave
2146 | Generic Agent Wave (2nd)
2147 | Impedance Respiration Wave
2148 | Brachial Arterial Pressure (BAP) Wave
2149 | Intracranial Pressure 1 (IC1) Wave
2150 | Intracranial Pressure 2 (IC2) Wave
2151 | Femoral Arterial Pressure (FAP) Wave
2152 | Vector ECG Wave
2153 | Transcutaneous Heater Power
2154 | BIS Standard Deviation | This variable is associated with ChannelID identifier.
2155 | EMG Standard Deviation (BIS) | This variable is associated with ChannelID identifier.
2156 | Asymmetric Index | This variable is associated with ChannelID identifier.
2157 | O2 Consumption Index Estimate
2158 | Aortic Blood Pressure (Ao) | Available on IntelliVue devices with firmware revision G or later. Arterial Blood Pressure in the Aorta (NOM_PRESS_BLD_AORT).
2159 | Umbilical Arterial Pressure (UAP) | Available on IntelliVue devices with firmware revision G or later.
2160 | Umbilical Venous Pressure (UVP) | Available on IntelliVue devices with firmware revision G or later.
2161 | Humidifier Activation Status
2162 | Nebulizer Time Setting
2163 | Edi Diaphragm Electrical Activity Waveform
2164 | Peak Diaphragm Electrical Activity (Edi)
2165 | Minimum Diaphragm Electrical Activity (Edi)
2166 | % Edi Trigger In NAVA
2167 | % Edi Cycle Off In NAVA
2168 | NAVA Level Setting
2169 | Edi Trigger Setting
2170 | Patient Body Mass Index (BMI)
2171 | Maximum Positive Auxiliary Pressure
2172 | Mean Positive Auxiliary Pressure
2173 | Minimum Positive Auxiliary Pressure
2174 | Patient CO2 Production Indexed By Body Surface
2175 | Patient CO2 Production Indexed By Body Weight
2176 | Tidal Volume Indexed By Body Weight
2177 | Functional Residual Capacity
2178 | Corrected QT Interval (QTc)
2179 | Airway N2 Concentration | Measurement from VueLink
2180 | Airway Agent Concentration | Measurement from VueLink. This variable is associated with a SourceID identifier.
2181 | Respiration Rate (Spirometry)
2182 | QT Heart Rate (QT-HR) | QT heart rate
2183 | Delta QTc
2184 | Humidification Type Setting
2185 | Normalized Rapid Shallow Breathing Index
2186 | Percent of Spontaneous Inspiratory Time
2187 | Energy Balance Calculated by Metabolics
2188 | Alveolar Volume
2189 | Apparatus Dead Space
2190 | Pleth (Pre-Ductal) | Pulse Rate from Plethysmogram (pre ductal)
2191 | Pleth (Post-Ductal) | Pulse Rate from Plethysmogram (post ductal)
2192 | Relative Perfusion (Pre-Ductal)
2193 | Relative Perfusion (Post-Ductal)
2194 | Cardioplegia Main Pump Time of Current Plegia Session
2195 | Cardioplegia Main Pump Time Since Stop of Plegia Session
2196 | Cardioplegia Main Pump Total Plegia Time
2197 | Cardioplegia Slave Pump Time Since Start of Current Plegia Session
2198 | Cardioplegia Slave Pump Time Since Stop of Plegia Session
2199 | Cardioplegia Slave Pump Total Plegia Time
2200 | PAV-based Lung Compliance
2201 | PAV-based Patient Resistance
2202 | PAV-based Total Resistance Estimate
2203 | PAV-based Lung Elastance
2204 | Backup Inspiratory Pressure Setting
2205 | Percent of Leakage
2206 | Maximum Inspiratory Time Setting
2207 | Plateau Time
2208 | QT Baseline
2209 | QT Heart Rate Baseline
2210 | High Density Lipoprotein
2211 | Low Density Lipoprotein
2212 | C-reactive Protein
2213 | Urine Hemoglobin
2214 | Noninvasive Total Hemoglobin (SpHb)
2215 | Noninvasive Oxygen Content (SpOC)
2216 | Noninvasive Carboxyhemoglobin (SpCO)
2217 | Noninvasive Methemoglobin (SpMet)
2218 | Pleth Variability Index (PVI)
2219 | Regional O2 Saturation (rSO2)
2220 | rSO2 Area Under Curve (AUC) | This variable is associated with ChannelID identifier.
2221 | Generic O2 Saturation (3)
2222 | Generic O2 Saturation (4)
2223 | Tissue O2 Saturation (StO2)
2224 | Nitrogen Balance (NsLoss)
2225 | Inspired O2 Concentration | See variable 52527 for partial pressure
2226 | Inspired O2 Partial Pressure | This variable is calculated from variable 52523 and will output inspired O2 partial pressure.
2228 | NIBP Systolic
2229 | NIBP Diastolic
2230 | NIBP Mean
2231 | Patient Delta Temperature
2232 | SpO2
2233 | ECG Heart Rate
2234 | Respiration Rate
2235 | Device Inactivation Status | Possible values: "enabled" and "audio-paused". See "Multiset Alarms" on help file.
2236 | Active Alarms | See "Multiset Alarms" on help file.
2238 |  | Consult detailed help for more information on this variable
2239 | Latched Alarms | See "Multiset Alarms" on help file.
2241 |  | Consult detailed help for more information on this variable
2242 | Inactive Alarms | See "Multiset Alarms" on help file.
2244 |  | Consult detailed help for more information on this variable
2245 | General ST Level 1 | Physiological identifier: NOM_ECG_AMPL_ST_I
2246 | General ST Level 2 | Physiological identifier: NOM_ECG_AMPL_ST_II
2247 | General ST Level 3 | Physiological identifier: NOM_ECG_AMPL_ST_III
2248 | ECG Measurement Status | ECG Rhythm Status
2250 |  | Consult detailed help for more information on this variable
2251 | Patient Age | Patient age. The Monitor device calculates the age of the patient from the birth data manually entered on the Monitor.
2252 | Patient Sex | Possible values: "unknown", "male", "female", and "unspecified".
2253 | ECG Lead Name | Physiological identifier: NOM_ECG_AMPL_ST
2254 | Arterial Pressure (ART) Systolic
2255 | Arterial Pressure (ART) Diastolic
2256 | Arterial Pressure (ART) Mean
2257 | Left Atrial Mean Pressure (LAP)
2258 | Central Venous Pressure (CVP) | Available on IntelliVue devices with firmware revision G or later.
2259 | Device Language | Consult detailed help for more information on this variable
2260 | End Tidal CO2
2261 | Cardiac Beat-to-Beat Rate | Cardiac Beat-to-Beat Rate (NOM_ECG_CARD_BEAT_RATE_BTB).
2262 | Device Specification ID | Variables 764, 765 and 4201 must be interpreted conjointly.
2264 |  | Consult detailed help for more information on this variable
2265 | Device Specification Type | Variables 764, 765 and 4201 must be interpreted conjointly.
2267 |  | Consult detailed help for more information on this variable
2268 | Patient Birth Date | Format: 14 ascii digits string of "YYYYMMDDhhmmss"
2269 | ECG Lead MCL1 | Physiological identifier: NOM_ECG_AMPL_ST_MCL
2270 | ECG Lead V1 | Physiological identifier: NOM_ECG_AMPL_ST_V1
2271 | ECG Lead V6 | Physiological identifier: NOM_ECG_AMPL_ST_V6
2272 | ECG Lead aVL | Physiological identifier: NOM_ECG_AMPL_ST_AVL
2273 | ECG Lead aVF | Physiological identifier: NOM_ECG_AMPL_ST_AVF
2274 | ECG Lead aVR | Physiological identifier: NOM_ECG_AMPL_ST_AVR
2275 | ECG Lead V2 | Physiological identifier: NOM_ECG_AMPL_ST_V2
2276 | ECG Lead V3 | Physiological identifier: NOM_ECG_AMPL_ST_V3
2277 | ECG Lead V4 | Physiological identifier: NOM_ECG_AMPL_ST_V4
2278 | ECG Lead V5 | Physiological identifier: NOM_ECG_AMPL_ST_V5
2279 | Premature Ventricular Contractions
2280 | LAP Systolic
2281 | LAP Diastolic
2282 | Device Unique ID
2283 | Patient Bed ID | This variable contains a printable string identifying the system location.
2284 | Patient ID
2285 | Patient Name
2286 | Patient Comment Text
2287 | Device Hardware Revision | This variable is obsolete because it may contains incorrect information. Please use variables 764, 765 and 4201 must be used instead.
2288 | Serial Number | This variable is obsolete because it may contains incorrect information. Please use variables 764, 765 and 4201 must be used instead.
2289 | Device Date/Time | Format: 16 ascii digits string of "YYYYMMDDhhmmssff", where the "ff" is second_fractions that is not used by IntelliVue monitors.
2290 | Arterial BP Systolic
2291 | Arterial BP Diastolic
2292 | Arterial BP Mean
2293 | Pulmonary Artery BP - Systolic
2294 | Pulmonary Artery BP - Diastolic
2295 | Pulmonary Artery BP Mean
2296 | Right Atrial Pressure (RAP) | Available on IntelliVue devices with firmware revision G or later.
2297 | Left Atrial Pressure (LAP) | Available on IntelliVue devices with firmware revision G or later.
2298 | Intra Cranial Pressure (ICP) | Available on IntelliVue devices with firmware revision G or later.
2299 | Generic NIBP Systolic Pressure | This variable is same as variable 2, except that the early estimated values are marked as NonSignificant in this variable.
2300 | Generic NIBP Diastolic Pressure | This variable is same as variable 3, except that the early estimated values are marked as NonSignificant in this variable.
2301 | Generic NIBP Mean Pressure | This variable is same as variable 4, except that the early estimated values are marked as NonSignificant in this variable.
2302 | Patient Admission Status | Possible values: "empty", "pre-admitted", "admitted", and "discharged".
2303 | ECG Lead V | Physiological identifier: NOM_ECG_AMPL_ST_V
2304 | System General Status | This variable identifies the current operating mode of the device. Possible values: "not specified", "monitoring", "demo", and "service". Moreover, "- standby" or "- config" can be appended
2305 | Frequency | This variable describes the frequency of the main power supply. Possible values: "unspecified", "50", and "60".
2306 | Device Status | Possible values: "disconnected", "unassociated", "operating".
2307 | Device Software Revision | This variable is obsolete because it may contains incorrect information. Please use variables 764, 765 and 4201 must be used instead.
2308 | Patient First Name
2309 | Patient Temperature
2310 | Device Altitude
2311 | Inspiratory Minimum CO2
2312 | Skin Temperature
2313 | Rectal Temperature
2314 | Nasopharyngeal Temperature
2315 | Esophageal Temperature
2316 | Core Temperature
2317 | Arterial Temperature
2318 | Venous Temperature
2319 | Device Number | Device Model Number (example: "M8000")
2320 | Manufacturers Name | "Philips"
2321 | Arterial Blood Pressure (ART) | Available on IntelliVue devices with firmware revision G or later. Arterial Blood Pressure (NOM_PRESS_BLD_ART).
2322 | Pulmonary Artery Blood Pressure (PA) | Available on IntelliVue devices with firmware revision G or later.
2323 | Arterial O2 Saturation (SaO2) | Obsolete variable that may be removed in a future release. The label of this variable is not correct. Please use variable 14 instead.
2324 | Patient Type | Possible values: "unspecified", "adult", "pediatric", and "neonatal".
2325 | Patient Pace Maker Status | Possible values: "not paced", "paced gen".
2326 | Pulse Rate ( Unspecified Source )
2327 | Central Venous Pressure Diastolic
2328 | Central Venous Mean Pressure
2329 | Central Venous Pressure Systolic
2330 | Product Description | Variables 764, 765 and 4201 must be interpreted conjointly.
2332 |  | Consult detailed help for more information on this variable
2333 | System Alarm Status | Overall alert state of the system.
2335 |  | Consult detailed help for more information on this variable
2336 | Patient Alarm Priority | Overall Patient alarm priority. This is the current highest priority value of currently active Patient Alarms.
2338 |  | Consult detailed help for more information on this variable
2339 | Technical Alarm Priority | Overall Technical alarm priority. This is the current highest priority value of currently active Technical Alarms.
2341 |  | Consult detailed help for more information on this variable
2342 | Audio Alarm Priority | Overall audible alarm priority. This is the current highest priority value of currently active audible alarms.
2344 |  | Consult detailed help for more information on this variable
2345 | Technical Alarm Source ID
2346 | Technical Alarm Global State | Consult detailed help for more information on this variable
2347 | Technical Alarm Priority | Consult detailed help for more information on this variable
2348 | Technical Alarm Flag | Consult detailed help for more information on this variable
2349 | Technical Alarm Priority
2350 | Technical Alarm Description
2351 | Patient Alarm Source ID
2352 | Patient Alarm Global State | Consult detailed help for more information on this variable
2353 | Patient Alarm Priority | Consult detailed help for more information on this variable
2354 | Patient Alarm Flag | Consult detailed help for more information on this variable
2355 | Patient Alarm Priority
2356 | Patient Alarm Description
2357 | Device Nomenclature Revision | Format: major.minor (such as 1.0)
2358 | Device Text Catalog Revision | Format: "MM.NN mm.nn", where "MM.NN" is text catalog revision and "mm.nn" is language revision.
2359 | Device Data Format | Always 11 (Unicode).
2360 | Device Operating Area | Possible values: "unspecified", "operating room", "intensive care unit", "neonatal intensive care unit", and "cardiology care unit".
2361 | Ao Systolic
2362 | Ao Diastolic
2363 | Ao Mean
2364 | RAP Systolic
2365 | RAP Diastolic
2366 | RAP Mean
2367 | Intracranial Pressure ICP Systolic
2368 | Intracranial Pressure ICP Diastolic
2369 | ICP Mean
2370 | UAP Systolic
2371 | UAP Diastolic
2372 | UAP Mean
2373 | UVP Systolic
2374 | UVP Diastolic
2375 | UVP Mean
2376 | Unspecific Pressure Systolic
2377 | Unspecific Pressure Diastolic
2378 | Unspecific Pressure Mean
2379 | Transthoracic Impedance
2380 | Airway Respiration Rate
2381 | Regional O2 Saturation (rSO2 2)
2382 | Regional O2 Saturation (rSO2 3)
2383 | Regional O2 Saturation (rSO2 4)
2384 | Patient BSA Formula | Consult detailed help for more information on this variable
2385 | Pulse Rate from Plethysmogram
2386 | ECG Ectopic Status | Consult detailed help for more information on this variable
2387 | Patient Height at Entered Time
2388 | Patient Weight at Entered Time
2389 | Patient Body Surface Area (BSA)
2390 | Generic NIBP Pressure | Available on IntelliVue devices with firmware revision G or later.
2391 | Device Power Up Time | Number of seconds since the Monitor device's last power-on.
2392 | Arterial Blood Pressure (ABP) | Available on IntelliVue devices with firmware revision G or later. Arterial Blood Pressure (NOM_PRESS_BLD_ART_ABP).
2393 | Technical Alarm Internal Source ID | Consult detailed help for more information on this variable
2394 | Technical Alarm Source Type
2395 | Device Technical Alarm | Consult detailed help for more information on this variable
2396 | Patient Alarm Internal Source ID | Consult detailed help for more information on this variable
2397 | Patient Alarm Source Type
2398 | Patient Alarm | Consult detailed help for more information on this variable
2399 | Aortic Blood Pressure (Ao) | Available on IntelliVue devices with firmware revision G or later. Arterial Blood Pressure in the Aorta (NOM_PRESS_BLD_AORT).
2400 | Umbilical Arterial Pressure (UAP) | Available on IntelliVue devices with firmware revision G or later.
2401 | Umbilical Venous Pressure (UVP) | Available on IntelliVue devices with firmware revision G or later.
2402 | Regional O2 Saturation (rSO2)
2404 | NIBP Systolic
2405 | NIBP Diastolic
2406 | NIBP Mean
2407 | Patient Delta Temperature
2408 | SpO2
2409 | NIBP Pulse Rate
2410 | ECG Heart Rate
2411 | Respiration Rate
2412 | Device Inactivation Status | Possible values: "enabled" and "audio-paused". See "Multiset Alarms" on help file.
2413 | Active Alarms | See "Multiset Alarms" on help file.
2415 |  | Consult detailed help for more information on this variable
2416 | Latched Alarms | See "Multiset Alarms" on help file.
2418 |  | Consult detailed help for more information on this variable
2419 | Inactive Alarms | See "Multiset Alarms" on help file.
2421 |  | Consult detailed help for more information on this variable
2422 | General ST Level 1 | Physiological identifier: NOM_ECG_AMPL_ST_I
2423 | General ST Level 2 | Physiological identifier: NOM_ECG_AMPL_ST_II
2424 | General ST Level 3 | Physiological identifier: NOM_ECG_AMPL_ST_III
2425 | Temperature 3
2426 | Temperature 4
2427 | O2 Saturation (SO2)
2428 | Pleth Waveform
2429 | ECG Measurement Status | ECG Rhythm Status
2431 |  | Consult detailed help for more information on this variable
2432 | Patient Age | Patient age. The Monitor device calculates the age of the patient from the birth data manually entered on the Monitor.
2433 | Patient Sex | Possible values: "unknown", "male", "female", and "unspecified".
2434 | Generic O2 Saturation (1)
2435 | Generic O2 Saturation (2)
2436 | Cardiac Output
2437 | Fetal Heart Rate (FHR) | This variable is associated with ChannelID identifier.
2438 | Fetal Beat-to-Beat Rate (FHR) | This variable is associated with ChannelID identifier.
2439 | TOCO
2440 | ECG Lead Name | Physiological identifier: NOM_ECG_AMPL_ST
2441 | Arterial Pressure (ART) Systolic
2442 | Arterial Pressure (ART) Diastolic
2443 | Arterial Pressure (ART) Mean
2444 | Left Atrial Mean Pressure (LAP)
2445 | Central Venous Pressure (CVP) | Available on IntelliVue devices with firmware revision G or later.
2446 | Device Language | Consult detailed help for more information on this variable
2447 | ECG Lead II
2448 | Airway N2O Concentration | Measurement from VueLink
2449 | Inspired Desflurane
2450 | Inspired Sevoflurane
2451 | Fresh Gas Agent Value
2452 | Fresh Gas Halothane
2453 | Fresh Gas Enflurane
2454 | Fresh Gas Isoflurane
2455 | Inspired Halothane
2456 | Inspired Enflurane
2457 | Inspired Isoflurane
2458 | End Tidal CO2
2459 | Airway Pressure | Measured in the inspiratory path.
2460 | Airway N2O Wave
2461 | Halothane Concentration Wave
2462 | Enflurane Concentration Wave
2463 | Isoflurane Concentration Wave
2464 | Desflurane Concentration Wave
2465 | Sevoflurane Concentration Wave
2466 | Apnea Duration | This is the device's measurement with label="NLS_NOM_TIME_PD_APNEA" and physio_id="NOM_TIME_PD_APNEA".
2467 | End Tidal CO2
2468 | Cardiac Beat-to-Beat Rate | Cardiac Beat-to-Beat Rate (NOM_ECG_CARD_BEAT_RATE_BTB).
2469 | Airway Delta O2 Concentration
2470 | Cardiac Output Index
2471 | Stroke Volume
2472 | Stroke Volume Index
2473 | Systemic Vascular Resistance
2474 | Normalized Ionized Calcium Concentration (Ca)
2475 | Potassium Ion Concentration (K+)
2476 | QT Interval (QT)
2477 | Device Specification ID | Variables 764, 765 and 4201 must be interpreted conjointly.
2479 |  | Consult detailed help for more information on this variable
2480 | Device Specification Type | Variables 764, 765 and 4201 must be interpreted conjointly.
2482 |  | Consult detailed help for more information on this variable
2483 | Pleth Wave (2)
2484 | Total Number of Used Disposable
2485 | Cumulated Syringe Infused Volume
2486 | Current Therapy Time
2487 | Manually Entered Hematocrit
2488 | Infusion Drug Name
2489 | Infusion Time Remaining
2490 | Infusion Dose Rate Setting
2491 | Volume To Be Infused Original Setting
2492 | KVO Rate Setting
2493 | Inspired Minute Volume
2494 | I:E Ratio | This variable contains expiratory component of I:E Ratio assuming that the inspiratory component is "1".
2495 | SpO2
2496 | Patient Birth Date | Format: 14 ascii digits string of "YYYYMMDDhhmmss"
2497 | SpO2 (Pre-Ductal)
2498 | SpO2 (Post-Ductal)
2499 | ECG Lead III
2500 | Hematocrit
2501 | Pulmonary Vascular Resistance
2502 | Body Surface Area
2503 | ECG Lead MCL1 | Physiological identifier: NOM_ECG_AMPL_ST_MCL
2504 | ECG Lead V1 | Physiological identifier: NOM_ECG_AMPL_ST_V1
2505 | ECG Lead V6 | Physiological identifier: NOM_ECG_AMPL_ST_V6
2506 | ECG Lead aVL | Physiological identifier: NOM_ECG_AMPL_ST_AVL
2507 | ECG Lead aVF | Physiological identifier: NOM_ECG_AMPL_ST_AVF
2508 | ECG Lead aVR | Physiological identifier: NOM_ECG_AMPL_ST_AVR
2509 | ECG Lead V2 | Physiological identifier: NOM_ECG_AMPL_ST_V2
2510 | ECG Lead V3 | Physiological identifier: NOM_ECG_AMPL_ST_V3
2511 | ECG Lead V4 | Physiological identifier: NOM_ECG_AMPL_ST_V4
2512 | ECG Lead V5 | Physiological identifier: NOM_ECG_AMPL_ST_V5
2513 | Patient Temperature | This variable is associated with a SourceID Identifier.
2514 | Micro Environment Air Temperature
2515 | Tidal Volume
2516 | Static Resistance
2517 | Tidal Volume Setting
2518 | Pressure Sensitivity Setting
2519 | Sigh Interval Setting
2520 | Sigh Volume Setting
2521 | Pressure Support Level Above PEEP Setting
2522 | Spontaneous Minute Volume
2523 | Negative Inspiratory Pressure
2524 | Peak Spontaneous Flow
2525 | Flow-by Base Flow Setting
2526 | Flow Sensitivity Setting
2527 | Expiratory Time
2528 | Inspiratory Flow Setting
2529 | Expiratory Flow Setting
2530 | Inspiratory Flow Trigger Setting | Delivered by the Evita 2 VueLink Driver.
2531 | APRV Low Pressure Setting
2532 | APRV High Pressure Setting
2533 | Apnea Time Setting
2534 | Minute Volume Setting
2535 | Positive End Expiratory Pressure (PEEP)
2536 | Ventilator Work of Breathing
2537 | Rapid Shallow Breathing Index
2538 | Patient Work of Breathing
2539 | Systemic Vascular Resistance Index
2540 | Pulmonary Vascular Resistance Index
2541 | Continuous Cardiac Output
2542 | Continuous Cardiac Output Index
2543 | Pump Infusion Status
2544 | Hemoglobin (Hb) | This variable is associated with SourceID identifier.
2545 | Pleth Wave (Pre-Ductal)
2546 | Pleth Wave (Post-Ductal)
2547 | Arterial Venous O2 Content Difference
2548 | O2 Delivery (DO2)
2549 | O2 Extraction Ratio (O2ER)
2550 | O2 Delivery Index (DO2I)
2551 | O2 Consumption Index (VO2I)
2552 | Oxygenation Ratio (PaO2/FiO2) | PaO2 to FIO2 ratio expressed in mmHg to % ratio.
2553 | Physiologic Dead Space
2554 | Alveolar Ventilation Volume (VA)
2555 | Urea Nitrogen Concentration (BUN)
2556 | Creatinine Concentration (Crea) | Measured by the i-Stat module.
2557 | Plasma Osmolarity (Osm)
2558 | Urine Creatinine
2559 | Urine Potassium (K+)
2560 | Urine Osmolarity
2561 | Blood Urea Nitrogen Creatinine Ratio (BUN/cr)
2562 | Osmolar Clearance (COSM)
2563 | Creatinine Clearance (CreaCl)
2564 | Fractional Excretion of Sodium (FeNa)
2565 | Urine Serum Creatinine Ratio
2566 | Urine Potassium Excretion
2567 | Urine Sodium Excretion
2568 | Free Water Clearance (CH2O)
2569 | Inspired N2O
2570 | Premature Ventricular Contractions
2571 | LAP Systolic
2572 | LAP Diastolic
2573 | TcGas Heating Power
2574 | Abdominal Fetal Heart Rate (FHR)
2575 | Abdominal Fetal Beat-to-Beat Rate (FHR)
2576 | Patient Activity Level (Movement)
2577 | Patient Activity Level (Wearable)
2578 | Minimum Airway Pressure
2579 | P0.1 Occlusion Pressure
2580 | Patient CO2 Production | Volume of CO2 produced by respiration.
2581 | Relative Dead Space | Relative Dead Space Volume.
2582 | Procalcitonin (PCT)
2583 | Fetal Heart Rate (USound) | This variable is associated with ChannelID identifier.
2584 | Fetal Beat-to-Beat Rate (USound) | This variable is associated with ChannelID identifier.
2585 | HFV Mean Pressure Setting
2586 | HFV Tidal Volume Setting
2587 | PAV-based Resistance Setting
2588 | PAV-based Elastance Setting
2589 | Abdominal TOCO
2590 | Chest Movement Respiration Rate
2591 | Patient Movement Respiration Rate
2592 | Patient Movement Pulse Rate
2593 | Respiration Signal Quality Index (Movement)
2594 | Cardiac Power Index (CCO)
2595 | Cardiac Power Output (CCO)
2596 | Central Venous Pressure (Thermodilution)
2597 | Mean Arterial Pressure (Thermodilution)
2598 | Average Blood Temperature (Thermodilution)
2599 | O2 Delivery DO2 (Thermodilution)
2600 | O2 Consumption (Thermodilution)
2601 | O2 Consumption Index (Thermodilution)
2602 | Global End-Diastolic Volume (Thermodilution)
2603 | Global End-Diastolic Volum Index (Thermodilution)
2604 | Pulmonary Vascular Permeability Index (Thermodilution)
2605 | Stroke Volume (Thermodilution)
2606 | Stroke Volume Index (Thermodilution)
2607 | O2 Consumption Estimate
2608 | EEG1 Wave
2609 | EEG2 Wave
2610 | EEG3 Wave
2611 | EEG4 Wave
2612 | Spectral Edge Frequency
2613 | Burst Suppression Ratio | This variable is associated with ChannelID identifier.
2614 | Patient Energy Expenditure
2615 | Patient Respiratory Quotient
2616 | Spectral Edge Frequency (CH3)
2617 | Relative Power In Delta Band (CH3)
2618 | Relative Power In Theta Band (CH3)
2619 | Relative Power In Alpha Band (CH3)
2620 | Relative Power In Beta Band (CH3)
2621 | Spectral Edge Frequency (CH4)
2622 | Relative Power In Delta Band (CH4)
2623 | Relative Power In Theta Band (CH4)
2624 | Relative Power In Alpha (CH4)
2625 | Relative Power In Beta Band (CH4)
2626 | Device Unique ID
2627 | Patient Bed ID | This variable contains a printable string identifying the system location.
2628 | Patient ID
2629 | Patient Name
2630 | Patient Comment Text
2631 | Patient Stress Index
2632 | Generic Patient Temperature
2633 | Breath Cycle Type
2634 | Airway Gas Flow Rate | Measured in the inspiratory path.
2635 | Inspired Minute Volume
2636 | Positive End Expiratory Pressure (PEEP) Setting
2637 | Spectral Edge Frequency - 1st EEG
2638 | Bispectral Index | This variable is associated with ChannelID identifier.
2639 | Total Power (Ch1)
2640 | Spike Rate | This variable is associated with ChannelID identifier.
2641 | Spectral Edge Frequency - 2nd EEG
2642 | Total Power (Ch2)
2643 | Narcotrend Index
2644 | Sharp Transient Intensity
2645 | Total Power Relative
2646 | Device Hardware Revision | This variable is obsolete because it may contains incorrect information. Please use variables 764, 765 and 4201 must be used instead.
2647 | Serial Number | This variable is obsolete because it may contains incorrect information. Please use variables 764, 765 and 4201 must be used instead.
2648 | O2 Flush Percentage Setting
2649 | CO2 Tidal Production
2650 | Inspiratory Resistance
2651 | Device Date/Time | Format: 16 ascii digits string of "YYYYMMDDhhmmssff", where the "ff" is second_fractions that is not used by IntelliVue monitors.
2652 | Arterial BP Systolic
2653 | Arterial BP Diastolic
2654 | Arterial BP Mean
2655 | Pulmonary Artery BP - Systolic
2656 | Pulmonary Artery BP - Diastolic
2657 | Pulmonary Artery BP Mean
2658 | Right Atrial Pressure (RAP) | Available on IntelliVue devices with firmware revision G or later.
2659 | Left Atrial Pressure (LAP) | Available on IntelliVue devices with firmware revision G or later.
2660 | Intra Cranial Pressure (ICP) | Available on IntelliVue devices with firmware revision G or later.
2661 | Pulmonary Artery Wedge Pressure
2662 | Generic NIBP Systolic Pressure | This variable is same as variable 2, except that the early estimated values are marked as NonSignificant in this variable.
2663 | Generic NIBP Diastolic Pressure | This variable is same as variable 3, except that the early estimated values are marked as NonSignificant in this variable.
2664 | Generic NIBP Mean Pressure | This variable is same as variable 4, except that the early estimated values are marked as NonSignificant in this variable.
2665 | Inspired Tidal Volume
2666 | Expired Tidal Volume
2667 | Expired Minute Volume
2668 | Patient Blood Temperature
2669 | Injectate Temperature
2670 | Inspired O2
2671 | Inspired O2 Pressure | This variable is marked as obsolete because the label is incorrect. Please see variable 52523 for concentration and 52527 for partial pressure.
2672 | End Tidal O2
2673 | CO2 (Gas Analyzer)
2674 | End Tidal N2O
2675 | End Tidal Agent Concentration
2676 | End Tidal Desflurane
2677 | End Tidal Sevoflurane
2678 | End Tidal Halothane
2679 | End Tidal Enflurane
2680 | End Tidal Isoflurane
2681 | Fresh Gas N2O Flow | Mixer parameter. Setting of fresh gas N2O flow.
2682 | Fresh Gas Air Flow
2683 | Transcutaneous O2 Pressure
2684 | Transcutaneous CO2 Pressure
2685 | Arterial O2 Saturation (SaO2)
2686 | EEG Total Power
2687 | Bispectral Index (BIS)
2688 | EEG Burst Suppression Ratio
2689 | Spontaneous Alveolar Minute Volume
2690 | Airway Dead Space
2691 | Expired Spontaneous Tidal Volume
2692 | Spontaneous Expired Minute Volume
2693 | Total Alveolar Tidal Volume
2694 | Respiration Waveform
2695 | EEG Signal Quality Index
2696 | Wearable Respiration Rate
2697 | Wearable Pulse Rate
2698 | Wearable Patient Temperature
2699 | Respiration Signal Quality Index (Wearable)
2700 | Pulse Rate Signal Quality Index (Wearable)
2701 | Airway Pressure Wave
2702 | Citrate Blood Concentration Setting
2703 | Neutrophil Count
2704 | Fresh Gas Total Flow Setting | Mixer parameter.
2705 | I:E Ratio Setting | This variable is associated with ChannelID identifier. This variable is calculated from variable 2773, which is sent as a string in the format "1:xx.xx".
2706 | Patient Admission Status | Possible values: "empty", "pre-admitted", "admitted", and "discharged".
2707 | ECG Lead V | Physiological identifier: NOM_ECG_AMPL_ST_V
2708 | Plateau Pressure | This variable contains parameter from Philips label "NLS_NOM_VENT_PRESS_RESP_PLAT" and observed value "NLS_NOM_VENT_PRESS_RESP_PLAT".
2709 | I:E Ratio Setting
2710 | Total PEEP
2711 | APRV Low Time (Tlow)
2712 | System General Status | This variable identifies the current operating mode of the device. Possible values: "not specified", "monitoring", "demo", and "service". Moreover, "- standby" or "- config" can be appended
2713 | Frequency | This variable describes the frequency of the main power supply. Possible values: "unspecified", "50", and "60".
2714 | Device Status | Possible values: "disconnected", "unassociated", "operating".
2715 | Device Software Revision | This variable is obsolete because it may contains incorrect information. Please use variables 764, 765 and 4201 must be used instead.
2716 | Maximum Apnea Time Setting
2717 | Inspired Agent Concentration
2718 | Patient Weight
2719 | Patient First Name
2720 | Patient Temperature
2721 | Ejection Fraction
2722 | End Diastolic Volume
2723 | End Diastolic Volume Index
2724 | End Systolic Volume
2725 | End Systolic Volume Index
2726 | EEG Electromyography (EMG) | This variable is associated with ChannelID identifier.
2727 | Generic EEG Wave
2728 | ECG Pad Wave
2729 | ECG Paddle Wave
2730 | Unspecific Pressure Pulse Rate
2731 | Patient Temperature 1
2732 | Patient Temperature 2
2733 | APRV High Pressure (Phigh)
2734 | APRV Low Pressure (Plow)
2735 | I:E Ratio | This variable is associated with ChannelID identifier This variable is calculated from variable 1190 and sent as a string in the format "1:xx.xx".
2736 | Plateau Airway Pressure
2737 | Plateau Pressure Setting
2738 | FiO2 Upper Alarm Limit
2739 | FiO2 Lower Alarm Limit
2740 | Device Altitude
2741 | Blood Flow Rate Setting
2742 | Dialysate Flow Rate Setting
2743 | Effluent Flow Rate Setting
2744 | Access Pod Pressure
2745 | Return Pod Pressure
2746 | Filter Pod Pressure
2747 | Effluent Pod Pressure
2748 | Transmembrane Pressure
2749 | Hb Calculation Offset
2750 | Patient Time Since Last Ward Admission
2751 | Patient Total Admission Time
2752 | Patient Drainage Volume
2753 | Fluid Removal Actual Rate
2754 | Sum of Minimum Alveolar Concentration (MAC)
2755 | HFV Frequency Setting
2756 | Ventilator Base Flow Setting
2757 | Blood Pressure Difference
2758 | Generic Blood Pressure Difference 1
2759 | Generic Blood Pressure Difference 2
2760 | Patient Height
2761 | Patient Temperature Setting
2762 | Arterial O2 Content (CaO2)
2763 | Alveolar Arterial O2 Difference (AaDO2)
2764 | Venous O2 Content (CvO2)
2765 | Ionized Calcium Concentration (Ca)
2766 | Chloride Concentration (Cl-)
2767 | Activated Clotting Time | Measured by the i-Stat module.
2768 | Generic Temperature
2769 | Generic Temperature 2
2770 | Generic Temperature 3
2771 | Generic Temperature 4
2772 | Expiratory Resistance
2773 | TcGas Site Timer
2774 | O2 Saturation (SO2)
2775 | TcGas Transducer Temperature
2776 | Peak Expiratory Flow Setting
2777 | Auto Flow Active Status
2778 | Inspiratory Minimum CO2
2779 | Expired Tidal Volume Indexed by Body Weight
2780 | ATC Activation Status
2781 | Skin Temperature
2782 | Rectal Temperature
2783 | Nasopharyngeal Temperature
2784 | Esophageal Temperature
2785 | Core Temperature
2786 | Arterial Temperature
2787 | Venous Temperature
2788 | Generic Core Temperature 1
2789 | Generic Core Temperature 2
2790 | Device Number | Device Model Number (example: "M8000")
2791 | Drug Delivery Rate
2792 | Drug Concentration
2793 | Manufacturers Name | "Philips"
2794 | Pressure Alarm Delay
2795 | Arterial Blood Pressure (ART) | Available on IntelliVue devices with firmware revision G or later. Arterial Blood Pressure (NOM_PRESS_BLD_ART).
2796 | Pulmonary Artery Blood Pressure (PA) | Available on IntelliVue devices with firmware revision G or later.
2797 | Arterial/Main Pump Flow
2798 | Slave Pump Flow
2799 | Suction Pump Flow
2800 | Auxiliary Pump Flow
2801 | Auxiliary Pump Time Since Start of Current Plegia Session
2802 | Auxiliary Pump Time Since Stop of Plegia Session
2803 | Auxiliary Pump Delivered Volume in Current Plegia Session
2804 | Auxiliary Pump Total Volume
2805 | Auxiliary Pump Total Plegia Time
2806 | Cardioplegia Main Pump Flow
2807 | Cardioplegia Main Pump Delivered Volume of Current Plegia Session
2808 | Cardioplegia Main Pump Total Volume
2809 | Cardioplegia Slave Pump Flow
2810 | Cardioplegia Slave Pump Delivered Volume in Current Plegia Session
2811 | Cardioplegia Slave Pump Total Volume
2812 | Cycle Flow Sensitivity Setting
2813 | Hydrocarbon Plasma Concentration (HCO3)
2814 | Corpuscular Mean Hemoglobin Concentration
2815 | Arterial O2 Partial Pressure
2816 | Venous O2 Partial Pressure
2817 | Arterial CO2 Partial Pressure
2818 | Arterial O2 Saturation (SaO2) | Obsolete variable that may be removed in a future release. The label of this variable is not correct. Please use variable 14 instead.
2819 | Hematocrit (HCT) | Post-Dialysis Hematocrit
2820 | Respiratory Rate (Flow)
2821 | Respiratory Rate (CO2)
2822 | Patient Type | Possible values: "unspecified", "adult", "pediatric", and "neonatal".
2823 | Patient Pace Maker Status | Possible values: "not paced", "paced gen".
2824 | Heart Rate Signal Quality
2825 | Tracheal Pressure
2826 | Ventilator Leakage
2827 | Plateau Time Setting | This variable is obsolete that may be removed in a future release. The label of this variable is incorrect, please use variable 9585 instead.
2828 | Warmer Mattress Temperature
2829 | Spontaneous Fraction of Minute Volume
2830 | Breathing System Leakage
2831 | Inspired MAC Concentration
2832 | Inspiratory 2nd Agent
2833 | Expiratory 2nd Agent | This is the device's measurement with label="NLS_NOM_EMFC_etAGTs" and physio_id="NOM_CONC_AWAY_AGENT_ET".
2834 | Volume Assist Level Setting
2835 | Backup Tidal Volume Setting
2836 | Flow Acceleration Setting
2837 | Cumulated Halothane Consumption
2838 | Cumulated Enflurane Consumption
2839 | Cumulated Isoflurane Consumption
2840 | Cumulated Desflurane Consumption
2841 | Cumulated Sevoflurane Consumption
2842 | Air Humidity Setting
2843 | Mattress Temperature Setting
2844 | Inspired Agent Setting
2845 | Expired Agent Setting | End-Tidal agent concentration
2846 | Inspired Desflurane Setting
2847 | Expired Desflurane Setting | End-Tidal desflurane concentration
2848 | Inspired Sevoflurane Setting
2849 | Expired Sevoflurane Setting | End-Tidal sevoflurane concentration
2850 | Fresh Gas Sevoflurane
2851 | Fresh Gas Desflurane
2852 | Ambient Temperature
2853 | Ambient Pressure
2854 | Pulse Rate ( Unspecified Source )
2855 | Central Venous Pressure Diastolic
2856 | Central Venous Mean Pressure
2857 | Central Venous Pressure Systolic
2858 | Expiratory Volume Wave
2859 | Apnea Ventilation I:E Ratio Setting (Number)
2860 | Dynamic Plateau Airway Pressure
2861 | Transcutaneous Relative Heater Power
2862 | Spontaneous Respiratory Rate Setting
2863 | Product Description | Variables 764, 765 and 4201 must be interpreted conjointly.
2865 |  | Consult detailed help for more information on this variable
2866 | Replacement Cumulated Volume
2867 | Dialysate Cumulated Volume | This variable is associated with ChannelID identifier.
2868 | Effluent Cumulated Volume | This variable is associated with ChannelID identifier.
2869 | Fluid Removal Cumulated Volume
2870 | Base Pressure
2871 | PEEP/CPAP Setting
2872 | Sigh Rate Setting
2873 | Sigh Tidal Volume Setting
2874 | Data Averaging Interval Setting
2875 | Chart Interval Setting
2876 | Static Lung Compliance
2877 | Static Lung Resistance
2878 | Dynamic Lung Compliance
2879 | Dynamic Lung Resistance | NLS_NOM_RES_AWAY_DYN
2880 | Apnea Inspiratory Pressure Setting
2881 | Patient Hose Compliance
2882 | Oximeter Averaging Interval Setting
2883 | System Alarm Status | Overall alert state of the system.
2885 |  | Consult detailed help for more information on this variable
2886 | Patient Alarm Priority | Overall Patient alarm priority. This is the current highest priority value of currently active Patient Alarms.
2888 |  | Consult detailed help for more information on this variable
2889 | Technical Alarm Priority | Overall Technical alarm priority. This is the current highest priority value of currently active Technical Alarms.
2891 |  | Consult detailed help for more information on this variable
2892 | Audio Alarm Priority | Overall audible alarm priority. This is the current highest priority value of currently active audible alarms.
2894 |  | Consult detailed help for more information on this variable
2895 | Technical Alarm Source ID
2896 | Technical Alarm Global State | Consult detailed help for more information on this variable
2897 | Technical Alarm Priority | Consult detailed help for more information on this variable
2898 | Technical Alarm Flag | Consult detailed help for more information on this variable
2899 | Technical Alarm Priority
2900 | Technical Alarm Description
2901 | Patient Alarm Source ID
2902 | Patient Alarm Global State | Consult detailed help for more information on this variable
2903 | Patient Alarm Priority | Consult detailed help for more information on this variable
2904 | Patient Alarm Flag | Consult detailed help for more information on this variable
2905 | Patient Alarm Priority
2906 | Patient Alarm Description
2907 | Device Nomenclature Revision | Format: major.minor (such as 1.0)
2908 | Device Text Catalog Revision | Format: "MM.NN mm.nn", where "MM.NN" is text catalog revision and "mm.nn" is language revision.
2909 | Device Data Format | Always 11 (Unicode).
2910 | Device Operating Area | Possible values: "unspecified", "operating room", "intensive care unit", "neonatal intensive care unit", and "cardiology care unit".
2911 | Patient O2 Consumption (VO2)
2912 | Patient CO2 Production (VCO2)
2913 | Arterial O2 Saturation (Left)
2914 | Arterial O2 Saturation (Right)
2915 | SpO2 Difference
2916 | Perfusion Indicator
2917 | Relative Perfusion (Left)
2918 | Relative Perfusion (Right)
2919 | Ao Systolic
2920 | Ao Diastolic
2921 | Ao Mean
2922 | RAP Systolic
2923 | RAP Diastolic
2924 | RAP Mean
2925 | Intracranial Pressure ICP Systolic
2926 | Intracranial Pressure ICP Diastolic
2927 | ICP Mean
2928 | UAP Systolic
2929 | UAP Diastolic
2930 | UAP Mean
2931 | UVP Systolic
2932 | UVP Diastolic
2933 | UVP Mean
2934 | Unspecific Pressure Systolic
2935 | Unspecific Pressure Diastolic
2936 | Unspecific Pressure Mean
2937 | Cerebral Perfusion Pressure (CPP)
2938 | Stroke Volume Variation
2939 | Index of Left Ventricular Contractility | This variable contains measurement from label "NLS_NOM_GRAD_PRESS_BLD_AORT_POS_MAX" and observed value "NOM_GRAD_PRESS_BLD_AORT_POS_MAX".
2940 | Aortic Blood Pressure (Ao) Wave
2941 | Auxiliary Waveform
2942 | Intermittent Mandatory Ventilation Status
2943 | Left Cardiac Work
2944 | Right Cardiac Work
2945 | Serum Creatinine
2946 | Urine Plasma Osmolarity Ratio
2947 | Urine Sodium Potassium Ratio
2948 | Urine Volume
2949 | Left Cardiac Work Index
2950 | Right Cardiac Work Index
2951 | Light Intensity (SvO2 Module)
2952 | SpO2 SvO2 Difference
2953 | Normalized Post Replacement Volume
2954 | Normalized Effluent Volume | This variable is associated with ChannelID identifier.
2955 | Normalized Total Replacement Volume (PBP) | This variable is associated with ChannelID identifier.
2956 | Normalized Replacement Volume (Pre)
2957 | Normalized Ultrafiltration Volume | This variable is associated with ChannelID identifier.
2958 | Patient Fluid Balance Limit Setting
2959 | Total Replacement Volume (PBP) | This variable is associated with ChannelID identifier.
2960 | Cumulated Liquid Volume (Pre) | This variable is associated with ChannelID identifier.
2961 | Filtration Fraction Setting
2962 | Diff.-Temperature between Tblood and Tskin
2963 | Diff.-Temperature between Tblood and Trect
2964 | Diff.-Temperature between Tblood and Tnaso
2965 | Diff.-Temperature between Tblood and Tesop
2966 | Diff.-Temperature between Tblood and Tcore
2967 | Diff.-Temperature between Tblood and Tart
2968 | Diff.-Temperature between Tblood and Tven
2969 | Diff.-Temperature between Tskin and Trect
2970 | Diff.-Temperature between Tskin and Tnaso
2971 | Diff.-Temperature between Tskin and Tesop
2972 | Diff.-Temperature between Tskin and Tcore
2973 | Diff.-Temperature between Tskin and Tart
2974 | Diff.-Temperature between Tskin and Tven
2975 | Diff.-Temperature between Trect and Tnaso
2976 | Diff.-Temperature between Trect and Tesop
2977 | Diff.-Temperature between Trect and Tcore
2978 | Diff.-Temperature between Trect and Tart
2979 | Diff.-Temperature between Trect and Tven
2980 | Diff.-Temperature between Tnaso and Tesop
2981 | Diff.-Temperature between Tnaso and Tcore
2982 | Diff.-Temperature between Tnaso and Tart
2983 | Diff.-Temperature between Tnaso and Tven
2984 | Diff.-Temperature between Tesop and Tcore
2985 | Diff.-Temperature between Tesop and Tart
2986 | Diff.-Temperature between Tesop and Tven
2987 | Diff.-Temperature between Tcore and Tart
2988 | Diff.-Temperature between Tcore and Tven
2989 | Diff.-Temperature between Tart and Tven
2990 | Newborn Body Length
2991 | Newborn Gestational Age
2992 | Calories
2993 | Input/Output Totals
2994 | Newborn Apgar Scores
2995 | BSA Formula Boyd
2996 | BSA Formula Dubois
2997 | Pupil Reaction (Left Eye) | Light reaction of left eye's pupil.
2998 | Pupil Reaction (Right Eye) | Light reaction of right eye's pupil.
2999 | Pleth Gain
3000 | Airway Temperature
3001 | Ventilation Perfusion Index
3002 | O2 Extraction Index
3003 | Duration Above Base Pressure
3004 | Inspiration Rise Time
3005 | Continuous Positive Airway Pressure (CPAP)
3006 | HFV Amplitude
3007 | Urine Urea
3008 | Urine pH Value
3009 | Total of CO2 (tCO2) | Result of Blood gas Analysis.
3010 | Total Bilirubin
3011 | Serum Glucose
3012 | Urine Glucose
3013 | Direct Bilirubin
3014 | Serum Calcium
3015 | Total Serum Calcium
3016 | Serum Magnesium
3017 | Serum Phosphat
3018 | Serum Kalium (Potassium)
3019 | Serum Clorid
3020 | Serum Albumine
3021 | Urine Clorid
3022 | Serum Globulin
3023 | Serum Protein
3024 | Serum Urea
3025 | Partial Thromboplastin Time
3026 | Prothrombin Time
3027 | Thrombin Time
3028 | Thromboplastine Time (Quick Time)
3029 | Alkalische Phosphatase
3030 | Alpha Amylase
3031 | Cholesterinesterase
3032 | Creatinin Kinase
3033 | Creatine Kinase of Type Muscle-Brain
3034 | Creatine Kinase of Type Muscle
3035 | Gamma Glutamyl Transpeptidase
3036 | Glutamic Oxaloacetic Transaminase
3037 | Glutamic Pyruvic Transaminase
3038 | Cholesterin
3039 | Triglyzeride
3040 | Urine Protein | (Total) Protein in Urine
3041 | Urine Calcium
3042 | Carboxy Hemoglobin
3043 | Total Protein
3044 | Lactate Dehydrogenase
3045 | Aspartin Aminotransferase
3046 | Alveolarproteinose (Alkaline Phosphatase)
3047 | Reticulocyte Count
3048 | Coagulation Time
3049 | Erithrocyte Sedimentation Rate
3050 | Packed Cell Volume
3051 | Kaolin Cephalin Time
3052 | Minimum Pressure During Breathing Cycle
3053 | Infusion Pump Accumulated Volume
3054 | Transthoracic Impedance
3055 | Inspired Expired N2O Difference
3056 | Inspired Expired Halothane Difference
3057 | Inspired Expired Enfluran Difference
3058 | Inspired Expired Isofluran Difference
3059 | Inspired Expired Sevofluran Difference
3060 | Inspired Expired Desfluran Difference
3061 | Inspired Expired Agent Difference
3062 | Inspired Expired O2 Difference
3063 | Patient Circuit O2 | O2 measured in the Patient Circuit
3064 | Barometric Pressure in Sample Chamber | This variable is obsolete that may be removed in a future release. It is duplicated with variable 4082.
3065 | Minute Volume
3066 | Mean Airway Pressure
3067 | Airway Respiration Rate
3068 | Inspiratory Airway Pressure Waveform
3069 | Inspiratory Flow Waveform
3070 | HFV Inspired Mandatory Minute Volume
3071 | HFV Gas Transport Coefficient Value
3072 | HFV Inspired Tidal Volume
3073 | HFV Tidal Volume
3074 | Generic Heart Rate | Heart Rate received from an external device
3075 | PERF from the 2nd SpO2/PLETH module
3076 | Percent Alpha (Left)
3077 | Percent Beta (Left)
3078 | Absolute Alpha (Left)
3079 | Absolute Beta (Left)
3080 | Absolute Delta (Left)
3081 | Absolute Theta (Left)
3082 | Mean Dominant Frequency (Left)
3083 | Median Power Frequency (Left)
3084 | Peak Power Frequency (Left)
3085 | Spectral Edge Frequency (Left)
3086 | Total Power (Left)
3087 | Scale of the Left Channel EEG Wave
3088 | Percent Alpha (Right)
3089 | Percent Beta (Right)
3090 | Absolute Alpha (Right)
3091 | Absolute Beta (Right)
3092 | Absolute Delta (Right)
3093 | Absolute Theta (Right)
3094 | Mean Dominant Frequency (Right)
3095 | Median Power Frequency (Right)
3096 | Peak Power Frequency (Right)
3097 | Spectral Edge Frequency (Right)
3098 | Total Power (Right)
3099 | Scale of the Right Channel EEG Wave
3100 | Duration of Positive Pressure
3101 | Fresh Gas N2O Concentration
3102 | Liquid Level (Anesthetic Agent Bottle)
3103 | Liquid Level (ISOflurane Bottle)
3104 | Liquid Level (ENFlurane Bottle)
3105 | Liquid Level (HALothane Bottle)
3106 | Liquid Level (DESflurane Bottle)
3107 | Liquid Level (SEVoflurane Bottle)
3108 | Invasive Blood Pressure Wave
3109 | Urine Shift Volume
3110 | Urine Flow Rate
3111 | Mean Dominant Frequency (CH1)
3112 | Mean Dominant Frequency (CH2)
3113 | Peak Power Frequency (CH1)
3114 | Peak Power Frequency (CH2)
3115 | % of Power in the Alpha Band (CH1)
3116 | % of Power in the Alpha Band (CH2)
3117 | % of Power in the Beta Band (1st EEG)
3118 | % of Power in the Beta Band (2nd EEG)
3119 | Hemoglobin
3120 | Patient Temperature
3121 | Capillaries pH Value
3122 | Adjusted pH (at Patient Temperature) | Adjusted pH at Patient Temperature
3123 | Adjusted Capillaries pH Value
3124 | Partial O2 in the Capillaries
3125 | Adjusted PO2 (at Patient Temperature) | Adjusted PO2 at Patient Temperature
3126 | Adjusted PcO2 (at Patient Temperature)
3127 | Venous CO2 Partial Pressure (PvCO2)
3128 | Partial CO2 in the Capillaries
3129 | Adjusted PcCO2 (at Patient Temperature)
3130 | Calculated Total CO2 (tCO2)
3131 | Calculated SO2 | Calculated SaO2
3132 | Calculated SvO2
3133 | Calculated ScO2
3134 | Calculated HCO3
3135 | Base Excess of Extra Cellular Fluid
3136 | Calculated Anion Gap
3137 | Urea | Measured by the i-Stat module.
3138 | Manually Entered FiO2 | In the Blood Review window.
3139 | Calculated Arterial Blood Base Excess
3140 | Ionized Magnesium
3141 | Desired Module Temperature | It's the selected/adjusted Temp in i-Stat Window.
3142 | Calculated BUN/Creatinine Ratio | Calculated by the i-Stat module.
3143 | Calculated Urea/Creatinine Ratio | Calculated by the i-Stat module.
3144 | Lactate | Measured value by the i-Stat module.
3145 | Device Simulator Status | Device data with the values: Passed and Failed.
3146 | Pressure of Sensor 1 | Pressure value of sensor 1. Device specific data.
3147 | Pressure of Sensor 2 | Pressure value of sensor 2. Device specific data.
3148 | Module Temperature
3149 | Air Temperature
3150 | Micro Environment Air Temperature Setting
3151 | Weight of the Urine Disposable Bag
3152 | Total Urine Volume of the Current Measurement Period
3153 | Urine Fluid Density
3154 | Urine Fluid Temperature
3155 | Ventilation Activated Status | Parameter to Informs whether the Patient is Ventilated
3156 | Simplified Acuity Physiologic Score
3157 | Value Normal | Missing parameters in SAPS parameter set which are considered in normal range.
3158 | Aspartate Aminotransferase
3159 | Spontaneous Respiration Rate
3160 | Current Fluid in Urine Bag
3161 | Cardiac Function Index (Thermodilution)
3162 | Intrathoracic Blood Volume (Thermodilution)
3163 | Intrathoracic Blood Volume Index (Thermodilution)
3164 | Extravascular Lung Water (Thermodilution)
3165 | Extravascular Lung Water Index (Thermodilution)
3166 | NMT TOF Count
3167 | NMT TOF Ratio
3168 | Twitch Height of the 1Hz/0.1Hz Stimulation Response
3169 | Post Tetatic Count
3170 | Remaining Time to Next Stimulation
3171 | TOF Response TOF1
3172 | TOF Response TOF2
3173 | TOF Response TOF3
3174 | TOF Response TOF4
3175 | Activated Partial Thromboplastin Time (aPTT Whole Blood Time)
3176 | Activated Partial Thromboplastin Time (aPTT Plasma Equivalent Time)
3177 | Prothrombin Time (Whole Blood Time)
3178 | Prothrombin Time (Plasma Equivalent Time)
3179 | Prothrombin Time Ratio
3180 | Spontaneous Inspiration Time
3181 | Overdistension Index
3182 | Lung Time Constant
3183 | Airway Correlation Coefficient
3184 | Rate Volume Ratio
3185 | Normalized Ionized Calcium (Ca)
3186 | Peak Inspiratory Flow Setting
3187 | Trigger Sensitivity Setting
3188 | Plateau Time Setting
3189 | Multiple Sigh Number Setting
3190 | Apnea Tidal Volume Setting
3191 | Apnea Respiration Rate Setting
3192 | Apnea Peak Flow Setting
3193 | Apnea Inspired O2 Setting
3194 | PSV Pressure Setting | Pressure Support Ventilation Setting
3195 | Sigh Activation Status
3196 | Suction O2 Concentration Setting
3197 | PCV Control Pressure Setting
3198 | PCV Inspiratory Time Setting
3199 | PCV Apnea Respiration Rate Setting
3200 | PCV Apnea Inspiratory Time Setting
3201 | PCV Apnea O2 Setting
3202 | PCV Apnea High Airway Pressure Setting
3203 | Inspiratory Component of I:E Ratio Setting | Pressure Ventilation I component of I:E Ratio setting
3204 | Expiratory Component of I:E Ratio Setting | Pressure Ventilation E component of I:E Ratio setting
3205 | Inspiratory Component of Apnea Ventilation I:E Ratio Setting
3206 | Expiratory Component of Apnea Ventilation I:E Ratio Setting
3207 | Setting: Pressure Ventilation Control
3208 | Cycle Time Setting
3209 | IPPV Ventilation Respiration Rate Setting
3210 | IMV Ventilation Respiration Rate Setting
3211 | Pressure Support PEEP Setting
3212 | Apnea Ventilation Delay
3213 | O2 Monitoring Mode
3214 | O2 Calibration Mode
3215 | Inspiratory Time Setting
3216 | Expiratory Time Setting
3217 | Alarm Percentage on Rise Time Setting
3218 | Continuous Positive Airway Pressure (CPAP) Setting | This is the device's measurement with label="NLS_NOM_SETT_CONC_AWAY_O2" and physio_id="NOM_CONC_AWAY_O2".
3219 | Ventilator Flow Setting
3220 | Low Inspiratory Pressure Setting
3221 | HFV Flow Setting
3222 | HFV Respiration Rate Setting
3223 | Mechanical Ventilation Activated Status | Controlled Mechanical Ventilation Setting
3224 | SIMV Mode | Synchronized Intermittent Mandatory Ventilation Setting
3225 | Mandatory Inspired Minute Volume Setting
3226 | Infusion Pump Delivery Rate Setting
3227 | Airway Respiration Rate Setting
3228 | Trigger Flow/Volume Setting | This is the device's measurement with label="NLS_NOM_SETT_VENT_VOL_LUNG_TRAPD" and physio_id="NOM_SETT_VENT_VOL_LUNG_TRAPD".
3229 | Apnea Alarm Delay
3230 | HFV Amplitude Setting
3231 | Volume Alarm Delay
3232 | Pressure Increase Rate Setting
3233 | Volume Warning Setting | Delivered by the Evita 2 VueLink Driver.
3234 | Airway Inspiratory Pressure Lower Alarm Limit
3235 | Apnea Alarm Delay Setting
3236 | Pressure Ramp Setting (ASB Ramp)
3237 | Applied Tidal Volume Setting
3238 | Assist Pressure Sensitivity Setting | Assist Sensitivity. Used by the Bear 1000 ventilator.
3239 | High Inspiratory Pressure Setting
3240 | Background Flow Setting
3241 | Vaporizer Agent Concentration Setting
3242 | Vaporizer Isoflurane Concentration Setting
3243 | Vaporizer Enflurane Concentration Setting
3244 | Vaporizer Halothane Concentration Setting
3245 | Vaporizer Desflurane Concentration Setting
3246 | Vaporizer Sevoflurane Concentration Setting
3247 | Gas Measurement Sensor Position
3248 | O2 Measurement Sensor Position
3249 | Inspired Tidal Volume Setting
3250 | Incubator/Warmer Temperature Setting
3251 | Preset Period of UrVol Setting
3252 | Low Time Setting | Setting variable for Evita 4 Airway Pressure Release Ventilation Mode.
3253 | High Time Setting | Setting variable for Evita 4 Airway Pressure Release Ventilation Mode.
3254 | Low Pressure Setting | Setting variable for Evita 4 Airway Pressure Release Ventilation Mode.
3255 | High Pressure Setting | Setting variable for Evita 4 Airway Pressure Release Ventilation Mode.
3256 | Volume Assist Level Setting
3257 | CPAP Flow Assist Level Setting
3258 | Preset Stimulation Current Setting
3259 | Preset Stimulation Charge Setting
3260 | Preset Stimulation Impulse Duration Setting
3261 | NMT TOF Repetition Time | Preset Train of Four (Slow TOF) Repetition Time Setting
3262 | Panting Limit Setting
3263 | Airway Inspiratory Pressure Upper Alarm Limit | Alarm Limit: High Pressure
3264 | Sigh Volume
3265 | Sustained Pressure Alarm Limit
3266 | Intrathoracic Blood Volume (ITBV)
3267 | Intrathoracic Blood Volume Index (ITBI)
3268 | Extravascular Lung Water
3269 | Extravascular Lung Water Index
3270 | Global End Diastolic Volume
3271 | Global End Diastolic Volume Index
3272 | Cardiac Function Index
3273 | Mixed Venous O2 Saturation
3274 | Generic Term for the Transcutaneous Gases
3275 | Intrinsic PEEP
3276 | Positive Peak Inspiratory Pressure
3277 | Dead Space to Tidal Volume Ratio
3278 | Lung Volume Trapped
3279 | Lung Vital Capacity
3280 | Generic Lung Compliance
3281 | EEG Electromyography
3282 | Suppression Ratio
3283 | EEG Peak Power Frequency
3284 | End Tidal N2
3285 | Inspired N2
3286 | Generic Agent Wave
3287 | End Tidal MAC Concentration
3288 | Left Ventricular Stroke Volume Index
3289 | Left Ventricular Stroke Volume
3290 | Right Ventricular Stroke Volume
3291 | Glasgow Coma Score
3292 | SubScore of the GCS - Eye Response
3293 | SubScore of the GCS - Motoric Response
3294 | SubScore of the GCS - Verbal Response
3295 | Head Circumference
3296 | pH in Arterial Blood
3297 | Arterial Blood Hemoglobin (HB)
3298 | pH in Venous Blood
3299 | Urine Natrium
3300 | Natrium in Serum
3301 | pH in the Blood Plasma
3302 | Blood Natrium (Sodium)
3303 | Kalium (Potassium)
3304 | Generic Glucose Parameter
3305 | Partial CO2
3306 | Partial O2
3307 | Base Excess of Blood
3308 | Percent Alveolarvenous Shunt Qs/Qt
3309 | Extravascular Thermal Volume Index
3310 | VPB Rate
3311 | ECG Lead V1 Wave
3312 | ECG Lead V2 Wave
3313 | ECG Lead V3 Wave
3314 | ECG Lead V4 Wave
3315 | ECG Lead V5 Wave
3316 | ECG Lead V6 Wave
3317 | ECG Lead I Wave
3318 | ECG Lead II Wave
3319 | ECG Lead III Wave
3320 | ECG Lead aVR Wave
3321 | ECG Lead aVL Wave
3322 | ECG Lead aVF Wave
3323 | Mean Dominant Frequency (CH3)
3324 | Mean Dominant Frequency (CH4)
3325 | NMT DBS Count
3326 | Peak Power Frequency (CH3)
3327 | Peak Power Frequency (CH4)
3328 | Pump Infusion Mode
3329 | Remaining Time Counter | Time to Elapse Counter used by the i-Stat module.
3330 | APRV Low Pressure Synchronisation Time Setting
3331 | APRV High Pressure Synchronisation Time Setting
3332 | Acoustic Respiration Rate
3333 | Regional O2 Saturation (rSO2 2)
3334 | Regional O2 Saturation (rSO2 3)
3335 | Regional O2 Saturation (rSO2 4)
3336 | Regional O2 Saturation (rSO2 5)
3337 | Regional O2 Saturation (rSO2 6)
3338 | O2 Delivery Index (Thermodilution)
3339 | Noninvasive Venous Hemoglobin (SpHbv)
3340 | Noninvasive Arterial Hemoglobin (SpHba)
3341 | Generic Pulse Rate
3342 | Corrected Total Hemoglobin
3343 | Corrected Arterial Total Hemoglobin
3344 | Corrected Venous Total Hemoglobin
3345 | Systemic Vascular Resistance (Thermodilution)
3346 | Systemic Vascular Resistance Index (Thermodilution)
3347 | Pulse Rate (Abdominal)
3348 | ECG Lead V3R
3349 | ECG Lead V4R
3350 | ECG Lead V5R
3351 | Inspired NO
3352 | ECG Lead V9
3353 | ECG Lead V7
3354 | ECG Lead V8
3355 | Apnea Ventilation I:E Ratio Setting | This variable is associated with ChannelID identifier. This variable is calculated from variable 4977 and 4978, which is sent as a string in the format "1:xx.xx".
3356 | Assisted Respiration Rate
3357 | Correlation Coefficient Squared
3358 | ECG Lead V6R
3359 | Citrate Solution Concentration
3360 | Calcium Solution Concentration
3361 | Citrate Acid Concentration
3362 | Citrate Load Setting
3363 | Backup Respiration Rate Setting
3364 | Inspired Expired Pressure Difference
3365 | Backup Inspired Expired Pressure Difference Setting
3366 | Femoral Arterial Pressure Mean
3367 | Femoral Arterial Pressure Systolic
3368 | Femoral Arterial Pressure Diastolic
3369 | Tidal Volume Lower Alarm Limit
3370 | Minute Volume Lower Alarm Limit
3371 | ECG Lead V7 Wave
3372 | ECG Lead V8 Wave
3373 | ECG Lead V9 Wave
3374 | Airway Aux Wave
3375 | Fetal DECG Wave
3376 | Patient BSA Formula | Consult detailed help for more information on this variable
3377 | Platelets (Thrombocyte Count)
3378 | Endotracheal Tube Type
3379 | Minute Volume Upper Alarm Limit
3380 | Pulse Rate from Plethysmogram
3381 | ECG Ectopic Status | Consult detailed help for more information on this variable
3382 | ST Lead ES
3383 | ST Lead AS
3384 | ST Lead AI
3385 | ST Index
3386 | Pulse Pressure
3387 | Mean Transit Time (MTt)
3388 | Alveolar Dead Space
3389 | Tympanic Temperature
3390 | Injectate Volume (CO)
3391 | SpO2 (Telemetry)
3392 | Pulse Rate (Telemetry)
3393 | Intracranial Pressure ICP1 (Systolic)
3394 | Intracranial Pressure ICP1 (Diastolic)
3395 | Intracranial Pressure ICP1 (Mean)
3396 | Intracranial Pressure ICP2 Systolic
3397 | Intracranial Pressure ICP2 Diastolic
3398 | Intracranial Pressure ICP2 Mean
3399 | Brachial Arterial Pressure (BA) Systolic
3400 | Brachial Arterial Pressure (BA) Diastolic
3401 | Brachial Arterial Pressure (BA) Mean
3402 | Vesical Temperature (Bladder)
3403 | Cerebral Temperature
3404 | ST Level for Lead I
3405 | ST Level for Lead II
3406 | ST Level for Lead III
3407 | ST Level for Lead aVR
3408 | ST Level for Lead aVL
3409 | ST Level for Lead aVF
3410 | ST Level for Lead V1
3411 | ST Level for Lead V2
3412 | ST Level for Lead V3
3413 | ST Level for Lead V4
3414 | ST Level for Lead V5
3415 | ST Level for Lead V6
3416 | Prothrombin Time (International Normalized Ratio)
3417 | PSV Tidal Volume | Tidal Volume in Pressure Support Ventilation Mode
3418 | Negative Inspiratory Force Index
3419 | A-Line ARX Index
3420 | Impedance Cardiography (ICG) Wave
3421 | Thoracic Fluid Content
3422 | Thoracic Fluid Content Index
3423 | Accelerated Cardiac Index
3424 | Heart Contractility Index
3425 | O2 Saturation (Left)
3426 | O2 Saturation (Right)
3427 | Pulse Pressure Variation
3428 | Downslope Time
3429 | Pulmonary Vascular Permeability Index | This is the device's measurement with label="NLS_NOM_PERM_VASC_PULM_INDEX" and physio_id="NOM_PERM_VASC_PULM_INDEX". See also variable 8791.
3430 | ART Pulse Rate | ART pulse rate
3431 | Umbilical Arterial Pressure Pulse Rate
3432 | Patient Visit Number | Patient encounter ID contains the ID of the current visit of the patient
3433 | Patient Height at Entered Time
3434 | Patient Weight at Entered Time
3435 | Patient Body Surface Area (BSA)
3436 | Inspired O2 Setting (FiO2)
3437 | Inspiratory Component of I:E Ratio Setting
3438 | Expiratory Component of I:E Ratio Setting
3439 | Lung Elastance
3440 | Backup Inspiratory Time Setting
3441 | Pressure Support Level Above Phigh
3442 | Total Respiration Duration Setting
3443 | Heater Power
3444 | Capillary Refill Time
3445 | Estimated Patient Circuit Leakage
3446 | Estimated Ventilation Total Leakage
3447 | Ventilation Mode Additional Information
3448 | Backup Plateau Time Setting
3449 | ECG Lead V3R Wave
3450 | ECG Lead V4R Wave
3451 | ECG Lead V5R Wave
3452 | ECG Lead V6R Wave
3453 | Signal Quality Index (BIS) | This variable is associated with ChannelID identifier.
3454 | Continuous NIBP Systolic
3455 | Continuous NIBP Diastolic
3456 | Continuous NIBP Mean
3457 | Next Intervention Time
3458 | Total Power (CH3)
3459 | Total Power (CH4)
3460 | Total Power (Left)
3461 | Total Power (Right)
3462 | PCV I:E Ratio Setting | This variable is associated with ChannelID identifier. This variable is calculated from variable 4975 and 4976, which is sent as a string in the format "1:xx.xx".
3463 | Inspiratory CO2 (FiCO2)
3464 | End Tidal NO2
3465 | Fresh Gas Total Flow
3466 | Bolus Volume Setting
3467 | Generic NIBP Pressure | Available on IntelliVue devices with firmware revision G or later.
3468 | Endotracheal Tube Diameter
3469 | Mandatory Respiration Rate
3470 | Mandatory Minute Volume
3471 | Mandatory Tidal Volume
3472 | N2O Consumption
3473 | Device Power Up Time | Number of seconds since the Monitor device's last power-on.
3474 | O2 Consumption
3475 | Air Consumption
3476 | Arterial Blood Pressure (ABP) | Available on IntelliVue devices with firmware revision G or later. Arterial Blood Pressure (NOM_PRESS_BLD_ART_ABP).
3477 | ABP Pulse Rate | ABP pulse rate
3478 | Signal Noise Ratio (CCO)
3479 | Spontaneous Airway Respiration Rate
3480 | Synchronized Respiration Rate
3481 | Mechanical Respiration Rate
3482 | Daily Urine Output
3483 | Activated Partial Thromboplastin Time Ratio
3484 | Global Ejection Fraction | This is the device's measurement with label="NLS_NOM_FRACT_EJECT" and physio_id="NOM_FRACT_EJECT". See also 8790.
3485 | Generic CO2 Parameter | This variable contains "PECO2, Partial expired CO2", which is measured from label="NLS_NOM_VENT_AWAY_CO2_EXP", physio_id="NOM_VENT_AWAY_CO2_EXP".
3486 | Central Venous O2 Saturation (ScvO2)
3487 | Suppression Ratio (BIS) | This variable is associated with ChannelID identifier.
3488 | Right Ventricular Stroke Work Index
3489 | Base Excess of Extracellular Fluid
3490 | Rise Time Setting
3491 | Time Duration at Low Pressure (Tlow)
3492 | Time Duration at High Pressure (Thigh)
3493 | Spontaneous Tidal Volume
3494 | Cardiac Troponin l
3495 | Cardiopulmonary Bypass Status
3496 | Cardiac Brain Natriuretic Peptide
3497 | Mean Dominant Frequency
3498 | Filtration Current Volume
3499 | Calcium Flow Rate Setting
3500 | Calcium Compensation Setting
3501 | Cerebral State Index (CSI)
3502 | EEG Amplitude | This variable is associated with ChannelID identifier.
3503 | Patient Middle Names | The Middle Name attribute contains the middle name of the patient.
3504 | Ao Pulse Rate
3505 | PAP Pulse Rate
3506 | CVP Pulse Rate
3507 | RAP Pulse Rate
3508 | LAP Pulse Rate
3509 | ICP Pulse Rate
3510 | Umbilical Venous Pressure Pulse Rate
3511 | FAP Pulse Rate
3512 | BAP Pulse Rate
3513 | IC1 Pulse Rate
3514 | Intracranial Pressure ICP2 Pulse Rate
3515 | Pulse Rate (P1)
3516 | Pulse Rate (P2)
3517 | Pulse Rate (P3)
3518 | Pulse Rate (P4)
3519 | Pulse Rate from Plethysmogram (Left)
3520 | Pleth Wave (Right)
3521 | Pleth Wave (Left)
3522 | Calculated SO2
3523 | Airway Respiration Rate (Ohmeda)
3524 | Airway Respiration Rate Setting (Ohmeda)
3525 | Fresh Gas O2 Concentration
3526 | Manually entered FIO2 or Flow | This variable contains measurement from label "NLS_NOM_CONC_AWAY_O2" and observed value "NOM_CONC_AWAY_O2".
3527 | Generic O2 Parameter | This is the device's measurement with label="NLS_NOM_SETT_CONC_AWAY_O2" and physio_id="NOM_CONC_AWAY_O2".
3528 | Fresh Gas O2 Flow Setting | Mixer parameter.
3529 | Airway Pressure Numeric Value (Ohmeda)
3530 | Inspiratory Peak Flow
3531 | Maximum Pressure During Breathing Cycle
3532 | Peak Pressure
3533 | Peak Inspiratory Flow Setting
3534 | Airway Maximum Pressure Setting
3535 | Peak Inspiratory Pressure Setting
3536 | Setting: Inspiratory Pressure (Ohmeda)
3537 | Mean Airway Pressure (Ohmeda)
3538 | EEG Wave (Left)
3539 | EEG Wave (Right)
3540 | % of Power in the Delta Band - 1st EEG
3541 | % of Power in the Delta Band - 2nd EEG
3542 | Percent Delta - Left Side
3543 | Percent Delta - Right Side
3544 | % of Power in the Theta Band (1st EEG)
3545 | % of Power in the Theta Band (2nd EEG)
3546 | Percent Theta - Left Side
3547 | Percent Theta - Right Side
3548 | Generic Calculation Constant | This variable contains measurement from Philips label "NLS_NOM_NUM_CALC_CONST" and observed value "NOM_METRIC_NOS". See also variable 8493.
3549 | Pressure Ventilation Control Pressure
3550 | Right Heart Ejection Fraction
3551 | White Blood Count (Leucocyte Count)
3552 | Red Blood Count (Erithrocyte Count)
3553 | Mean Corpuscular Volume
3554 | Corpuscular Mean Hemoglobin Content
3555 | Ferrum
3556 | Fetal Hemoglobin
3557 | MetHemoglobin
3558 | Airway Volume Wave
3559 | Expiratory Peak Flow
3560 | Expiratory Peak Flow
3561 | The V stands for Volume
3562 | The R stands for Rate
3563 | Adjusted PaCO2 (at Patient Temperature)
3564 | Incubator Heating Power | Power Required to set Air&Pat Temp in the Incubator
3565 | Actual Patient Age | Actual patient age, provided by a module, measured in years. Please note that this Patient Age will be available only when the corresponding module is connected to the Monitor. Variable 169 contains the patient age as provided by the Monitor itself (calculated from the birth date entered on the Monitor device.)
3566 | The P stands for Pressure
3567 | PCV Apnea Control Pressure Setting
3568 | Fresh Gas Air Flow Setting | Mixer parameter.
3569 | PEEP/CPAP Lower Alarm Limit
3570 | Tidal Volume Upper Alarm Limit
3571 | T-Tre Temperature difference between Temp and Trect
3572 | T-Tbl Temperature difference between Temp and Tblood
3573 | T-Tco Temperature difference between Temp and Tcore
3574 | T-Tsk Temperature difference between Temp and Tskin
3575 | T-Tes Temperature difference between Temp and Tesoph
3576 | T-Tna Temperature difference between Temp and Tnaso
3577 | T-Tar Temperature difference between Temp and Tart
3578 | T-Tve Temperature difference between Temp and Tven
3579 | T-Taw Temperature difference between Temp and Tairwy
3580 | Tre-T Temperature difference between Trect and Temp
3581 | Tre-bl Temperature difference between Trect and Tblood
3582 | Tre-sk Temperature difference between Trect and Tskin
3583 | Tre-aw Temperature difference between Trect and Tairwy
3584 | Temperature difference between Tblood and Temp
3585 | Tbl-aw Temperature difference between Tblood and Tairwy
3586 | Tco-T Temperature difference between Tcore and Temp
3587 | Tco-re Temperature difference between Tcore and Trect
3588 | Tco-bl Temperature difference between Tcore and Tblood
3589 | Tco-sk Temperature difference between Tcore and Tskin
3590 | Tco-es Temperature difference between Tcore and Tesoph
3591 | Tco-na Temperature difference between Tcore and Tnaso
3592 | Tco-aw Temperature difference between Tcore and Tairwy
3593 | Tsk-T Temperature difference between Tskin and Temp
3594 | Tsk-bl Temperature difference between Tskin and Tblood
3595 | Tsk-aw Temperature difference between Tskin and Tairwy
3596 | Tes-T Temperature difference between Tesoph and Temp
3597 | Tes-re Temperature difference between Tesoph and Trect
3598 | Tes-bl Temperature difference between Tesoph and Tblood
3599 | Tes-sk Temperature difference between Tesoph and Tskin
3600 | Tes-na Temperature difference between Tesoph and Tnaso
3601 | Tes-aw Temperature difference between Tesoph and Tairwy
3602 | Tna-T Temperature difference between Tnaso and Temp
3603 | Tna-re Temperature difference between Tnaso and Trect
3604 | Tna-bl Temperature difference between Tnaso and Tblood
3605 | Tna-sk Temperature difference between Tnaso and Tskin
3606 | Tna-aw Temperature difference between Tnaso and Tairwy
3607 | Tar-T Temperature difference between Tart and Temp
3608 | Tar-re Temperature difference between Tart and Trect
3609 | Tar-bl Temperature difference between Tart and Tblood
3610 | Tar-co Temperature difference between Tart and Tcore
3611 | Tar-sk Temperature difference between Tart and Tskin
3612 | Tar-es Temperature difference between Tart and Tesoph
3613 | Tar-na Temperature difference between Tart and Tnaso
3614 | Tar-aw Temperature difference between Tart and Tairwy
3615 | Tve-T Temperature difference between Tven and Temp
3616 | Tve-re Temperature difference between Tven and Trect
3617 | Tve-bl Temperature difference between Tven and Tblood
3618 | Tve-co Temperature difference between Tven and Tcore
3619 | Tve-sk Temperature difference between Tven and Tskin
3620 | Tve-es Temperature difference between Tven and Tesoph
3621 | Tve-na Temperature difference between Tven and Tnaso
3622 | MAC Concentration (Awake)
3623 | Tve-ar Temperature difference between Tven and Tart
3624 | Tve-aw Temperature difference between Tven and Tairwy
3625 | Taw-T Temperature difference between Tairwy and Temp
3626 | Taw-re Temperature difference between Tairwy and Trect
3627 | Taw-bl Temperature difference between Tairwy and Tblood
3628 | Taw-co Temperature difference between Tairwy and Tcore
3629 | Taw-sk Temperature difference between Tairwy and Tskin
3630 | Taw-es Temperature difference between Tairwy and Tesoph
3631 | Taw-na Temperature difference between Tairwy and Tnaso
3632 | Taw-ar Temperature difference between Tairwy and Tart
3633 | Taw-ve Temperature difference between Tairwy and Tven
3634 | Generic Agent 1 Wave
3635 | End Tidal Generic Agent 1
3636 | Inspired Generic Agent 1
3637 | Generic Agent 2 Wave
3638 | End Tidal Generic Agent 2
3639 | Inspired Generic Agent 2
3640 | Adjusted Arterial Blood pH Value
3641 | Adjusted Venous Blood pH Value
3642 | Adjusted PvCO2 (at Patient Temperature)
3643 | Adjusted PvO2 (at Patient Temperature)
3644 | Generic Calculation Constant | This variable contains parameter from Philips label "NLS_NOM_NUM_CATHETER_CONST" and observed value "NOM_NUM_CATHETER_CONST". See also variable 8395.
3645 | ECG Lead V Wave
3646 | ECG Lead MCL Wave
3647 | ECG Lead MCL1 Wave
3648 | MAC Concentration
3649 | Pleth Wave (Telemetry)
3650 | NIBP Systolic (Telemetry)
3651 | NIBP Diastolic (Telemetry)
3652 | NIBP Mean (Telemetry)
3653 | Generic Pressure 1 (P1) Systolic
3654 | Generic Pressure 1 (P1) Diastolic
3655 | Generic Pressure 1 (P1) Mean
3656 | Generic Pressure 2 (P2) Systolic
3657 | Generic Pressure 2 (P2) Diastolic
3658 | Generic Pressure 2 (P2) Mean
3659 | Generic Pressure 3 (P3) Systolic
3660 | Generic Pressure 3 (P3) Diastolic
3661 | Generic Pressure 3 (P3) Mean
3662 | Generic Pressure 4 (P4) Systolic
3663 | Generic Pressure 4 (P4) Diastolic
3664 | Generic Pressure 4 (P4) Mean
3665 | Incubator Humidity
3666 | Predictive Oral Temperature
3667 | Predictive Rectal Temperature
3668 | Predictive Axillary Temperature
3669 | Perfusion (Telemetry)
3670 | Right to Left Heart Shunt
3671 | Adjusted PcO2 (at Patient Temperature)
3672 | Adjusted Arterial PaO2 (at Patient Temperature)
3673 | Pulse Rate from Plethysmogram (Right)
3674 | Patient O2 Consumption per Body Surface
3675 | Patient O2 Consumption per Body Weight
3676 | Ideal Patient Weight
3677 | End Flow Setting
3678 | PSV Backup Delay Setting
3679 | Technical Alarm Internal Source ID | Consult detailed help for more information on this variable
3680 | Technical Alarm Source Type
3681 | Device Technical Alarm | Consult detailed help for more information on this variable
3682 | Patient Alarm Internal Source ID | Consult detailed help for more information on this variable
3683 | Patient Alarm Source Type
3684 | Patient Alarm | Consult detailed help for more information on this variable
3685 | Accoustic Respiration Rate Wave
3686 | Blood Temperature Wave
3687 | Non-specific Pressure 1 Wave
3688 | Generic Invasive Blood Pressure P1 | Available on IntelliVue devices.
3689 | Generic Invasive Blood Pressure P1 Systolic
3690 | Generic Invasive Blood Pressure P1 Diastolic
3691 | Generic Invasive Blood Pressure P1 Mean
3692 | Non-specific Pressure 2 Wave
3693 | Generic Invasive Blood Pressure P2 | Available on IntelliVue devices.
3694 | Generic Invasive Blood Pressure P2 Systolic
3695 | Generic Invasive Blood Pressure P2 Diastolic
3696 | Generic Invasive Blood Pressure P2 Mean
3697 | Non-specific Pressure 3 Wave
3698 | Generic Invasive Blood Pressure P3 | Available on IntelliVue devices.
3699 | Generic Invasive Blood Pressure P3 Systolic
3700 | Generic Invasive Blood Pressure P3 Diastolic
3701 | Generic Invasive Blood Pressure P3 Mean
3702 | Non-specific Pressure 4 Wave
3703 | Generic Invasive Blood Pressure P4 | Available on IntelliVue devices.
3704 | Generic Invasive Blood Pressure P4 Systolic
3705 | Generic Invasive Blood Pressure P4 Diastolic
3706 | Generic Invasive Blood Pressure P4 Mean
3707 | Non-specific Pressure 5 Wave
3708 | Generic Invasive Blood Pressure P5 | Available on IntelliVue devices.
3709 | Generic Invasive Blood Pressure P5 Systolic
3710 | Generic Invasive Blood Pressure P5 Diastolic
3711 | Generic Invasive Blood Pressure P5 Mean
3712 | Non-specific Pressure 6 Wave
3713 | Generic Invasive Blood Pressure P6 | Available on IntelliVue devices.
3714 | Generic Invasive Blood Pressure P6 Systolic
3715 | Generic Invasive Blood Pressure P6 Diastolic
3716 | Generic Invasive Blood Pressure P6 Mean
3717 | Non-specific Pressure 7 Wave
3718 | Generic Invasive Blood Pressure P7 | Available on IntelliVue devices.
3719 | Generic Invasive Blood Pressure P7 Systolic
3720 | Generic Invasive Blood Pressure P7 Diastolic
3721 | Generic Invasive Blood Pressure P7 Mean
3722 | Non-specific Pressure 8 Wave
3723 | Generic Invasive Blood Pressure P8 | Available on IntelliVue devices.
3724 | Generic Invasive Blood Pressure P8 Systolic
3725 | Generic Invasive Blood Pressure P8 Diastolic
3726 | Generic Invasive Blood Pressure P8 Mean
3727 | Intra-Uterine Pressure (IUP) Wave
3728 | Intra-Uterine Pressure
3729 | Intra-Uterine Pressure Systolic
3730 | Intra-Uterine Pressure Diastolic
3731 | Intra-Uterine Pressure Mean
3732 | Unspecified Blood Pressure
3733 | Unspecified Blood Pressure Systolic
3734 | Unspecified Blood Pressure Diastolic
3735 | Unspecified Blood Pressure Mean
3736 | Airway Flow Wave
3737 | Auxiliary Wave/Parameter
3738 | Global Ejection Fraction (2) | This is the device's measurement with label="NLS_NOM_EMFC_GEF" and physio_id="NOM_METRIC_NOS". See also 7810.
3739 | Pulmonary Vascular Permeability Index (2) | This is the device's measurement with label="NLS_NOM_EMFC_PVPI" and physio_id="NOM_METRIC_NOS". See also variable 6493.
3740 | Pressure Drop
3741 | Replacement Pre-delivered Rate Setting
3742 | Replacement Post-delivered Rate Setting
3743 | Pre-blood Flow Rate Setting
3744 | Patient Weight Balance Setting
3745 | Anticoagulant Bolus Interval Setting
3746 | Anticoagulant Infusion Rate Setting
3747 | Ultrafiltration Ratio Setting
3748 | Therapy Elapsed Time
3749 | Airway Volume
3750 | Processed Cumulated Blood Volume
3751 | Replacement Pre-delivered Cumulated Volume
3752 | Replacement Post-delivered Cumulated Volume | This variable is associated with ChannelID identifier.
3753 | Pre-blood Cumulated Volume | This variable is associated with ChannelID identifier.
3754 | Syringe Speed
3755 | Maximum Volume Gain Setting
3756 | Patient Fluid Balance
3757 | Arterial Blood Pressure (ABP) Wave
3758 | Arterial Blood Pressure (ART) Wave
3759 | Pulmonary Arterial Pressure (PAP) Wave
3760 | Central Venous Pressure (CVP) Wave
3761 | Right Atrial Pressure (RAP) Wave
3762 | Left Atrial Pressure (LAP) Wave
3763 | Intra-cranial Pressure (ICP) Wave
3764 | Umbilical Arterial Pressure (UAP) Wave
3765 | Umbilical Venous Pressure (UVP) Wave
3766 | Airway CO2 Level Waveform
3767 | Airway O2 Wave
3768 | Airway N2 Wave
3769 | Generic Agent Wave (2nd)
3770 | Impedance Respiration Wave
3771 | Brachial Arterial Pressure (BAP) Wave
3772 | Intracranial Pressure 1 (IC1) Wave
3773 | Intracranial Pressure 2 (IC2) Wave
3774 | Femoral Arterial Pressure (FAP) Wave
3775 | Vector ECG Wave
3776 | Transcutaneous Heater Power
3777 | BIS Standard Deviation | This variable is associated with ChannelID identifier.
3778 | EMG Standard Deviation (BIS) | This variable is associated with ChannelID identifier.
3779 | Asymmetric Index | This variable is associated with ChannelID identifier.
3780 | O2 Consumption Index Estimate
3781 | Aortic Blood Pressure (Ao) | Available on IntelliVue devices with firmware revision G or later. Arterial Blood Pressure in the Aorta (NOM_PRESS_BLD_AORT).
3782 | Umbilical Arterial Pressure (UAP) | Available on IntelliVue devices with firmware revision G or later.
3783 | Umbilical Venous Pressure (UVP) | Available on IntelliVue devices with firmware revision G or later.
3784 | Humidifier Activation Status
3785 | Nebulizer Time Setting
3786 | Edi Diaphragm Electrical Activity Waveform
3787 | Peak Diaphragm Electrical Activity (Edi)
3788 | Minimum Diaphragm Electrical Activity (Edi)
3789 | % Edi Trigger In NAVA
3790 | % Edi Cycle Off In NAVA
3791 | NAVA Level Setting
3792 | Edi Trigger Setting
3793 | Patient Body Mass Index (BMI)
3794 | Maximum Positive Auxiliary Pressure
3795 | Mean Positive Auxiliary Pressure
3796 | Minimum Positive Auxiliary Pressure
3797 | Patient CO2 Production Indexed By Body Surface
3798 | Patient CO2 Production Indexed By Body Weight
3799 | Tidal Volume Indexed By Body Weight
3800 | Functional Residual Capacity
3801 | Corrected QT Interval (QTc)
3802 | Airway N2 Concentration | Measurement from VueLink
3803 | Airway Agent Concentration | Measurement from VueLink. This variable is associated with a SourceID identifier.
3804 | Respiration Rate (Spirometry)
3805 | QT Heart Rate (QT-HR) | QT heart rate
3806 | Delta QTc
3807 | Humidification Type Setting
3808 | Normalized Rapid Shallow Breathing Index
3809 | Percent of Spontaneous Inspiratory Time
3810 | Energy Balance Calculated by Metabolics
3811 | Alveolar Volume
3812 | Apparatus Dead Space
3813 | Pleth (Pre-Ductal) | Pulse Rate from Plethysmogram (pre ductal)
3814 | Pleth (Post-Ductal) | Pulse Rate from Plethysmogram (post ductal)
3815 | Relative Perfusion (Pre-Ductal)
3816 | Relative Perfusion (Post-Ductal)
3817 | Cardioplegia Main Pump Time of Current Plegia Session
3818 | Cardioplegia Main Pump Time Since Stop of Plegia Session
3819 | Cardioplegia Main Pump Total Plegia Time
3820 | Cardioplegia Slave Pump Time Since Start of Current Plegia Session
3821 | Cardioplegia Slave Pump Time Since Stop of Plegia Session
3822 | Cardioplegia Slave Pump Total Plegia Time
3823 | PAV-based Lung Compliance
3824 | PAV-based Patient Resistance
3825 | PAV-based Total Resistance Estimate
3826 | PAV-based Lung Elastance
3827 | Backup Inspiratory Pressure Setting
3828 | Percent of Leakage
3829 | Maximum Inspiratory Time Setting
3830 | Plateau Time
3831 | QT Baseline
3832 | QT Heart Rate Baseline
3833 | High Density Lipoprotein
3834 | Low Density Lipoprotein
3835 | C-reactive Protein
3836 | Urine Hemoglobin
3837 | Noninvasive Total Hemoglobin (SpHb)
3838 | Noninvasive Oxygen Content (SpOC)
3839 | Noninvasive Carboxyhemoglobin (SpCO)
3840 | Noninvasive Methemoglobin (SpMet)
3841 | Pleth Variability Index (PVI)
3842 | Regional O2 Saturation (rSO2)
3843 | rSO2 Area Under Curve (AUC) | This variable is associated with ChannelID identifier.
3844 | Generic O2 Saturation (3)
3845 | Generic O2 Saturation (4)
3846 | Tissue O2 Saturation (StO2)
3847 | Nitrogen Balance (NsLoss)
3848 | Inspired O2 Concentration | See variable 52527 for partial pressure
3849 | Inspired O2 Partial Pressure | This variable is calculated from variable 52523 and will output inspired O2 partial pressure.
3851 | NIBP Systolic
3852 | NIBP Diastolic
3853 | NIBP Mean
3854 | Patient Delta Temperature
3855 | SpO2
3856 | NIBP Pulse Rate
3857 | ECG Heart Rate
3858 | Respiration Rate
3859 | Device Inactivation Status | Possible values: "enabled" and "audio-paused". See "Multiset Alarms" on help file.
3860 | Active Alarms | See "Multiset Alarms" on help file.
3862 |  | Consult detailed help for more information on this variable
3863 | Latched Alarms | See "Multiset Alarms" on help file.
3865 |  | Consult detailed help for more information on this variable
3866 | Inactive Alarms | See "Multiset Alarms" on help file.
3868 |  | Consult detailed help for more information on this variable
3869 | General ST Level 1 | Physiological identifier: NOM_ECG_AMPL_ST_I
3870 | General ST Level 2 | Physiological identifier: NOM_ECG_AMPL_ST_II
3871 | General ST Level 3 | Physiological identifier: NOM_ECG_AMPL_ST_III
3872 | Temperature 3
3873 | Temperature 4
3874 | O2 Saturation (SO2)
3875 | Pleth Waveform
3876 | ECG Measurement Status | ECG Rhythm Status
3878 |  | Consult detailed help for more information on this variable
3879 | Patient Age | Patient age. The Monitor device calculates the age of the patient from the birth data manually entered on the Monitor.
3880 | Patient Sex | Possible values: "unknown", "male", "female", and "unspecified".
3881 | Generic O2 Saturation (1)
3882 | Generic O2 Saturation (2)
3883 | Cardiac Output
3884 | Fetal Heart Rate (FHR) | This variable is associated with ChannelID identifier.
3885 | Fetal Beat-to-Beat Rate (FHR) | This variable is associated with ChannelID identifier.
3886 | TOCO
3887 | ECG Lead Name | Physiological identifier: NOM_ECG_AMPL_ST
3888 | Arterial Pressure (ART) Systolic
3889 | Arterial Pressure (ART) Diastolic
3890 | Arterial Pressure (ART) Mean
3891 | Left Atrial Mean Pressure (LAP)
3892 | Central Venous Pressure (CVP) | Available on IntelliVue devices with firmware revision G or later.
3893 | Device Language | Consult detailed help for more information on this variable
3894 | ECG Lead II
3895 | Airway N2O Concentration | Measurement from VueLink
3896 | Inspired Desflurane
3897 | Inspired Sevoflurane
3898 | Fresh Gas Agent Value
3899 | Fresh Gas Halothane
3900 | Fresh Gas Enflurane
3901 | Fresh Gas Isoflurane
3902 | Inspired Halothane
3903 | Inspired Enflurane
3904 | Inspired Isoflurane
3905 | End Tidal CO2
3906 | Airway Pressure | Measured in the inspiratory path.
3907 | Airway N2O Wave
3908 | Halothane Concentration Wave
3909 | Enflurane Concentration Wave
3910 | Isoflurane Concentration Wave
3911 | Desflurane Concentration Wave
3912 | Sevoflurane Concentration Wave
3913 | Apnea Duration | This is the device's measurement with label="NLS_NOM_TIME_PD_APNEA" and physio_id="NOM_TIME_PD_APNEA".
3914 | End Tidal CO2
3915 | Cardiac Beat-to-Beat Rate | Cardiac Beat-to-Beat Rate (NOM_ECG_CARD_BEAT_RATE_BTB).
3916 | Airway Delta O2 Concentration
3917 | Cardiac Output Index
3918 | Stroke Volume
3919 | Stroke Volume Index
3920 | Systemic Vascular Resistance
3921 | Normalized Ionized Calcium Concentration (Ca)
3922 | Potassium Ion Concentration (K+)
3923 | QT Interval (QT)
3924 | Device Specification ID | Variables 764, 765 and 4201 must be interpreted conjointly.
3926 |  | Consult detailed help for more information on this variable
3927 | Device Specification Type | Variables 764, 765 and 4201 must be interpreted conjointly.
3929 |  | Consult detailed help for more information on this variable
3930 | Pleth Wave (2)
3931 | Total Number of Used Disposable
3932 | Cumulated Syringe Infused Volume
3933 | Current Therapy Time
3934 | Manually Entered Hematocrit
3935 | Infusion Drug Name
3936 | Infusion Time Remaining
3937 | Infusion Dose Rate Setting
3938 | Volume To Be Infused Original Setting
3939 | KVO Rate Setting
3940 | Inspired Minute Volume
3941 | I:E Ratio | This variable contains expiratory component of I:E Ratio assuming that the inspiratory component is "1".
3942 | SpO2
3943 | Patient Birth Date | Format: 14 ascii digits string of "YYYYMMDDhhmmss"
3944 | SpO2 (Pre-Ductal)
3945 | SpO2 (Post-Ductal)
3946 | ECG Lead III
3947 | Hematocrit
3948 | Pulmonary Vascular Resistance
3949 | Body Surface Area
3950 | ECG Lead MCL1 | Physiological identifier: NOM_ECG_AMPL_ST_MCL
3951 | ECG Lead V1 | Physiological identifier: NOM_ECG_AMPL_ST_V1
3952 | ECG Lead V6 | Physiological identifier: NOM_ECG_AMPL_ST_V6
3953 | ECG Lead aVL | Physiological identifier: NOM_ECG_AMPL_ST_AVL
3954 | ECG Lead aVF | Physiological identifier: NOM_ECG_AMPL_ST_AVF
3955 | ECG Lead aVR | Physiological identifier: NOM_ECG_AMPL_ST_AVR
3956 | ECG Lead V2 | Physiological identifier: NOM_ECG_AMPL_ST_V2
3957 | ECG Lead V3 | Physiological identifier: NOM_ECG_AMPL_ST_V3
3958 | ECG Lead V4 | Physiological identifier: NOM_ECG_AMPL_ST_V4
3959 | ECG Lead V5 | Physiological identifier: NOM_ECG_AMPL_ST_V5
3960 | Patient Temperature | This variable is associated with a SourceID Identifier.
3961 | Micro Environment Air Temperature
3962 | Tidal Volume
3963 | Static Resistance
3964 | Tidal Volume Setting
3965 | Pressure Sensitivity Setting
3966 | Sigh Interval Setting
3967 | Sigh Volume Setting
3968 | Pressure Support Level Above PEEP Setting
3969 | Spontaneous Minute Volume
3970 | Negative Inspiratory Pressure
3971 | Peak Spontaneous Flow
3972 | Flow-by Base Flow Setting
3973 | Flow Sensitivity Setting
3974 | Expiratory Time
3975 | Inspiratory Flow Setting
3976 | Expiratory Flow Setting
3977 | Inspiratory Flow Trigger Setting | Delivered by the Evita 2 VueLink Driver.
3978 | APRV Low Pressure Setting
3979 | APRV High Pressure Setting
3980 | Apnea Time Setting
3981 | Minute Volume Setting
3982 | Positive End Expiratory Pressure (PEEP)
3983 | Ventilator Work of Breathing
3984 | Rapid Shallow Breathing Index
3985 | Patient Work of Breathing
3986 | Systemic Vascular Resistance Index
3987 | Pulmonary Vascular Resistance Index
3988 | Continuous Cardiac Output
3989 | Continuous Cardiac Output Index
3990 | Pump Infusion Status
3991 | Hemoglobin (Hb) | This variable is associated with SourceID identifier.
3992 | Pleth Wave (Pre-Ductal)
3993 | Pleth Wave (Post-Ductal)
3994 | Arterial Venous O2 Content Difference
3995 | O2 Delivery (DO2)
3996 | O2 Extraction Ratio (O2ER)
3997 | O2 Delivery Index (DO2I)
3998 | O2 Consumption Index (VO2I)
3999 | Oxygenation Ratio (PaO2/FiO2) | PaO2 to FIO2 ratio expressed in mmHg to % ratio.
4000 | Physiologic Dead Space
4001 | Alveolar Ventilation Volume (VA)
4002 | Urea Nitrogen Concentration (BUN)
4003 | Creatinine Concentration (Crea) | Measured by the i-Stat module.
4004 | Plasma Osmolarity (Osm)
4005 | Urine Creatinine
4006 | Urine Potassium (K+)
4007 | Urine Osmolarity
4008 | Blood Urea Nitrogen Creatinine Ratio (BUN/cr)
4009 | Osmolar Clearance (COSM)
4010 | Creatinine Clearance (CreaCl)
4011 | Fractional Excretion of Sodium (FeNa)
4012 | Urine Serum Creatinine Ratio
4013 | Urine Potassium Excretion
4014 | Urine Sodium Excretion
4015 | Free Water Clearance (CH2O)
4016 | Inspired N2O
4017 | Premature Ventricular Contractions
4018 | LAP Systolic
4019 | LAP Diastolic
4020 | TcGas Heating Power
4021 | Abdominal Fetal Heart Rate (FHR)
4022 | Abdominal Fetal Beat-to-Beat Rate (FHR)
4023 | Patient Activity Level (Movement)
4024 | Patient Activity Level (Wearable)
4025 | Minimum Airway Pressure
4026 | P0.1 Occlusion Pressure
4027 | Patient CO2 Production | Volume of CO2 produced by respiration.
4028 | Relative Dead Space | Relative Dead Space Volume.
4029 | Procalcitonin (PCT)
4030 | Fetal Heart Rate (USound) | This variable is associated with ChannelID identifier.
4031 | Fetal Beat-to-Beat Rate (USound) | This variable is associated with ChannelID identifier.
4032 | HFV Mean Pressure Setting
4033 | HFV Tidal Volume Setting
4034 | PAV-based Resistance Setting
4035 | PAV-based Elastance Setting
4036 | Abdominal TOCO
4037 | Chest Movement Respiration Rate
4038 | Patient Movement Respiration Rate
4039 | Patient Movement Pulse Rate
4040 | Respiration Signal Quality Index (Movement)
4041 | Cardiac Power Index (CCO)
4042 | Cardiac Power Output (CCO)
4043 | Central Venous Pressure (Thermodilution)
4044 | Mean Arterial Pressure (Thermodilution)
4045 | Average Blood Temperature (Thermodilution)
4046 | O2 Delivery DO2 (Thermodilution)
4047 | O2 Consumption (Thermodilution)
4048 | O2 Consumption Index (Thermodilution)
4049 | Global End-Diastolic Volume (Thermodilution)
4050 | Global End-Diastolic Volum Index (Thermodilution)
4051 | Pulmonary Vascular Permeability Index (Thermodilution)
4052 | Stroke Volume (Thermodilution)
4053 | Stroke Volume Index (Thermodilution)
4054 | O2 Consumption Estimate
4055 | EEG1 Wave
4056 | EEG2 Wave
4057 | EEG3 Wave
4058 | EEG4 Wave
4059 | Spectral Edge Frequency
4060 | Burst Suppression Ratio | This variable is associated with ChannelID identifier.
4061 | Patient Energy Expenditure
4062 | Patient Respiratory Quotient
4063 | Spectral Edge Frequency (CH3)
4064 | Relative Power In Delta Band (CH3)
4065 | Relative Power In Theta Band (CH3)
4066 | Relative Power In Alpha Band (CH3)
4067 | Relative Power In Beta Band (CH3)
4068 | Spectral Edge Frequency (CH4)
4069 | Relative Power In Delta Band (CH4)
4070 | Relative Power In Theta Band (CH4)
4071 | Relative Power In Alpha (CH4)
4072 | Relative Power In Beta Band (CH4)
4073 | Device Unique ID
4074 | Patient Bed ID | This variable contains a printable string identifying the system location.
4075 | Patient ID
4076 | Patient Name
4077 | Patient Comment Text
4078 | Patient Stress Index
4079 | Generic Patient Temperature
4080 | Breath Cycle Type
4081 | Airway Gas Flow Rate | Measured in the inspiratory path.
4082 | Inspired Minute Volume
4083 | Positive End Expiratory Pressure (PEEP) Setting
4084 | Spectral Edge Frequency - 1st EEG
4085 | Bispectral Index | This variable is associated with ChannelID identifier.
4086 | Total Power (Ch1)
4087 | Spike Rate | This variable is associated with ChannelID identifier.
4088 | Spectral Edge Frequency - 2nd EEG
4089 | Total Power (Ch2)
4090 | Narcotrend Index
4091 | Sharp Transient Intensity
4092 | Total Power Relative
4093 | Device Hardware Revision | This variable is obsolete because it may contains incorrect information. Please use variables 764, 765 and 4201 must be used instead.
4094 | Serial Number | This variable is obsolete because it may contains incorrect information. Please use variables 764, 765 and 4201 must be used instead.
4095 | O2 Flush Percentage Setting
4096 | CO2 Tidal Production
4097 | Inspiratory Resistance
4098 | Device Date/Time | Format: 16 ascii digits string of "YYYYMMDDhhmmssff", where the "ff" is second_fractions that is not used by IntelliVue monitors.
4099 | Arterial BP Systolic
4100 | Arterial BP Diastolic
4101 | Arterial BP Mean
4102 | Pulmonary Artery BP - Systolic
4103 | Pulmonary Artery BP - Diastolic
4104 | Pulmonary Artery BP Mean
4105 | Right Atrial Pressure (RAP) | Available on IntelliVue devices with firmware revision G or later.
4106 | Left Atrial Pressure (LAP) | Available on IntelliVue devices with firmware revision G or later.
4107 | Intra Cranial Pressure (ICP) | Available on IntelliVue devices with firmware revision G or later.
4108 | Pulmonary Artery Wedge Pressure
4109 | Generic NIBP Systolic Pressure | This variable is same as variable 2, except that the early estimated values are marked as NonSignificant in this variable.
4110 | Generic NIBP Diastolic Pressure | This variable is same as variable 3, except that the early estimated values are marked as NonSignificant in this variable.
4111 | Generic NIBP Mean Pressure | This variable is same as variable 4, except that the early estimated values are marked as NonSignificant in this variable.
4112 | Inspired Tidal Volume
4113 | Expired Tidal Volume
4114 | Expired Minute Volume
4115 | Patient Blood Temperature
4116 | Injectate Temperature
4117 | Inspired O2
4118 | Inspired O2 Pressure | This variable is marked as obsolete because the label is incorrect. Please see variable 52523 for concentration and 52527 for partial pressure.
4119 | End Tidal O2
4120 | CO2 (Gas Analyzer)
4121 | End Tidal N2O
4122 | End Tidal Agent Concentration
4123 | End Tidal Desflurane
4124 | End Tidal Sevoflurane
4125 | End Tidal Halothane
4126 | End Tidal Enflurane
4127 | End Tidal Isoflurane
4128 | Fresh Gas N2O Flow | Mixer parameter. Setting of fresh gas N2O flow.
4129 | Fresh Gas Air Flow
4130 | Transcutaneous O2 Pressure
4131 | Transcutaneous CO2 Pressure
4132 | Arterial O2 Saturation (SaO2)
4133 | EEG Total Power
4134 | Bispectral Index (BIS)
4135 | EEG Burst Suppression Ratio
4136 | Spontaneous Alveolar Minute Volume
4137 | Airway Dead Space
4138 | Expired Spontaneous Tidal Volume
4139 | Spontaneous Expired Minute Volume
4140 | Total Alveolar Tidal Volume
4141 | Respiration Waveform
4142 | EEG Signal Quality Index
4143 | Wearable Respiration Rate
4144 | Wearable Pulse Rate
4145 | Wearable Patient Temperature
4146 | Respiration Signal Quality Index (Wearable)
4147 | Pulse Rate Signal Quality Index (Wearable)
4148 | Airway Pressure Wave
4149 | Citrate Blood Concentration Setting
4150 | Neutrophil Count
4151 | Fresh Gas Total Flow Setting | Mixer parameter.
4152 | I:E Ratio Setting | This variable is associated with ChannelID identifier. This variable is calculated from variable 2773, which is sent as a string in the format "1:xx.xx".
4153 | Patient Admission Status | Possible values: "empty", "pre-admitted", "admitted", and "discharged".
4154 | ECG Lead V | Physiological identifier: NOM_ECG_AMPL_ST_V
4155 | Plateau Pressure | This variable contains parameter from Philips label "NLS_NOM_VENT_PRESS_RESP_PLAT" and observed value "NLS_NOM_VENT_PRESS_RESP_PLAT".
4156 | I:E Ratio Setting
4157 | Total PEEP
4158 | APRV Low Time (Tlow)
4159 | System General Status | This variable identifies the current operating mode of the device. Possible values: "not specified", "monitoring", "demo", and "service". Moreover, "- standby" or "- config" can be appended
4160 | Frequency | This variable describes the frequency of the main power supply. Possible values: "unspecified", "50", and "60".
4161 | Device Status | Possible values: "disconnected", "unassociated", "operating".
4162 | Device Software Revision | This variable is obsolete because it may contains incorrect information. Please use variables 764, 765 and 4201 must be used instead.
4163 | Maximum Apnea Time Setting
4164 | Inspired Agent Concentration
4165 | Patient Weight
4166 | Patient First Name
4167 | Patient Temperature
4168 | Ejection Fraction
4169 | End Diastolic Volume
4170 | End Diastolic Volume Index
4171 | End Systolic Volume
4172 | End Systolic Volume Index
4173 | EEG Electromyography (EMG) | This variable is associated with ChannelID identifier.
4174 | Generic EEG Wave
4175 | ECG Pad Wave
4176 | ECG Paddle Wave
4177 | Unspecific Pressure Pulse Rate
4178 | Patient Temperature 1
4179 | Patient Temperature 2
4180 | APRV High Pressure (Phigh)
4181 | APRV Low Pressure (Plow)
4182 | I:E Ratio | This variable is associated with ChannelID identifier This variable is calculated from variable 1190 and sent as a string in the format "1:xx.xx".
4183 | Plateau Airway Pressure
4184 | Plateau Pressure Setting
4185 | FiO2 Upper Alarm Limit
4186 | FiO2 Lower Alarm Limit
4187 | Device Altitude
4188 | Blood Flow Rate Setting
4189 | Dialysate Flow Rate Setting
4190 | Effluent Flow Rate Setting
4191 | Access Pod Pressure
4192 | Return Pod Pressure
4193 | Filter Pod Pressure
4194 | Effluent Pod Pressure
4195 | Transmembrane Pressure
4196 | Hb Calculation Offset
4197 | Patient Time Since Last Ward Admission
4198 | Patient Total Admission Time
4199 | Patient Drainage Volume
4200 | Fluid Removal Actual Rate
4201 | Sum of Minimum Alveolar Concentration (MAC)
4202 | HFV Frequency Setting
4203 | Ventilator Base Flow Setting
4204 | Blood Pressure Difference
4205 | Generic Blood Pressure Difference 1
4206 | Generic Blood Pressure Difference 2
4207 | Patient Height
4208 | Patient Temperature Setting
4209 | Arterial O2 Content (CaO2)
4210 | Alveolar Arterial O2 Difference (AaDO2)
4211 | Venous O2 Content (CvO2)
4212 | Ionized Calcium Concentration (Ca)
4213 | Chloride Concentration (Cl-)
4214 | Activated Clotting Time | Measured by the i-Stat module.
4215 | Generic Temperature
4216 | Generic Temperature 2
4217 | Generic Temperature 3
4218 | Generic Temperature 4
4219 | Expiratory Resistance
4220 | TcGas Site Timer
4221 | O2 Saturation (SO2)
4222 | TcGas Transducer Temperature
4223 | Peak Expiratory Flow Setting
4224 | Auto Flow Active Status
4225 | Inspiratory Minimum CO2
4226 | Expired Tidal Volume Indexed by Body Weight
4227 | ATC Activation Status
4228 | Skin Temperature
4229 | Rectal Temperature
4230 | Nasopharyngeal Temperature
4231 | Esophageal Temperature
4232 | Core Temperature
4233 | Arterial Temperature
4234 | Venous Temperature
4235 | Generic Core Temperature 1
4236 | Generic Core Temperature 2
4237 | Device Number | Device Model Number (example: "M8000")
4238 | Drug Delivery Rate
4239 | Drug Concentration
4240 | Manufacturers Name | "Philips"
4241 | Pressure Alarm Delay
4242 | Arterial Blood Pressure (ART) | Available on IntelliVue devices with firmware revision G or later. Arterial Blood Pressure (NOM_PRESS_BLD_ART).
4243 | Pulmonary Artery Blood Pressure (PA) | Available on IntelliVue devices with firmware revision G or later.
4244 | Arterial/Main Pump Flow
4245 | Slave Pump Flow
4246 | Suction Pump Flow
4247 | Auxiliary Pump Flow
4248 | Auxiliary Pump Time Since Start of Current Plegia Session
4249 | Auxiliary Pump Time Since Stop of Plegia Session
4250 | Auxiliary Pump Delivered Volume in Current Plegia Session
4251 | Auxiliary Pump Total Volume
4252 | Auxiliary Pump Total Plegia Time
4253 | Cardioplegia Main Pump Flow
4254 | Cardioplegia Main Pump Delivered Volume of Current Plegia Session
4255 | Cardioplegia Main Pump Total Volume
4256 | Cardioplegia Slave Pump Flow
4257 | Cardioplegia Slave Pump Delivered Volume in Current Plegia Session
4258 | Cardioplegia Slave Pump Total Volume
4259 | Cycle Flow Sensitivity Setting
4260 | Hydrocarbon Plasma Concentration (HCO3)
4261 | Corpuscular Mean Hemoglobin Concentration
4262 | Arterial O2 Partial Pressure
4263 | Venous O2 Partial Pressure
4264 | Arterial CO2 Partial Pressure
4265 | Arterial O2 Saturation (SaO2) | Obsolete variable that may be removed in a future release. The label of this variable is not correct. Please use variable 14 instead.
4266 | Hematocrit (HCT) | Post-Dialysis Hematocrit
4267 | Respiratory Rate (Flow)
4268 | Respiratory Rate (CO2)
4269 | Patient Type | Possible values: "unspecified", "adult", "pediatric", and "neonatal".
4270 | Patient Pace Maker Status | Possible values: "not paced", "paced gen".
4271 | Heart Rate Signal Quality
4272 | Tracheal Pressure
4273 | Ventilator Leakage
4274 | Plateau Time Setting | This variable is obsolete that may be removed in a future release. The label of this variable is incorrect, please use variable 9585 instead.
4275 | Warmer Mattress Temperature
4276 | Spontaneous Fraction of Minute Volume
4277 | Breathing System Leakage
4278 | Inspired MAC Concentration
4279 | Inspiratory 2nd Agent
4280 | Expiratory 2nd Agent | This is the device's measurement with label="NLS_NOM_EMFC_etAGTs" and physio_id="NOM_CONC_AWAY_AGENT_ET".
4281 | Volume Assist Level Setting
4282 | Backup Tidal Volume Setting
4283 | Flow Acceleration Setting
4284 | Cumulated Halothane Consumption
4285 | Cumulated Enflurane Consumption
4286 | Cumulated Isoflurane Consumption
4287 | Cumulated Desflurane Consumption
4288 | Cumulated Sevoflurane Consumption
4289 | Air Humidity Setting
4290 | Mattress Temperature Setting
4291 | Inspired Agent Setting
4292 | Expired Agent Setting | End-Tidal agent concentration
4293 | Inspired Desflurane Setting
4294 | Expired Desflurane Setting | End-Tidal desflurane concentration
4295 | Inspired Sevoflurane Setting
4296 | Expired Sevoflurane Setting | End-Tidal sevoflurane concentration
4297 | Fresh Gas Sevoflurane
4298 | Fresh Gas Desflurane
4299 | Ambient Temperature
4300 | Ambient Pressure
4301 | Pulse Rate ( Unspecified Source )
4302 | Central Venous Pressure Diastolic
4303 | Central Venous Mean Pressure
4304 | Central Venous Pressure Systolic
4305 | Expiratory Volume Wave
4306 | Apnea Ventilation I:E Ratio Setting (Number)
4307 | Dynamic Plateau Airway Pressure
4308 | Transcutaneous Relative Heater Power
4309 | Spontaneous Respiratory Rate Setting
4310 | Product Description | Variables 764, 765 and 4201 must be interpreted conjointly.
4312 |  | Consult detailed help for more information on this variable
4313 | Replacement Cumulated Volume
4314 | Dialysate Cumulated Volume | This variable is associated with ChannelID identifier.
4315 | Effluent Cumulated Volume | This variable is associated with ChannelID identifier.
4316 | Fluid Removal Cumulated Volume
4317 | Base Pressure
4318 | PEEP/CPAP Setting
4319 | Sigh Rate Setting
4320 | Sigh Tidal Volume Setting
4321 | Data Averaging Interval Setting
4322 | Chart Interval Setting
4323 | Static Lung Compliance
4324 | Static Lung Resistance
4325 | Dynamic Lung Compliance
4326 | Dynamic Lung Resistance | NLS_NOM_RES_AWAY_DYN
4327 | Apnea Inspiratory Pressure Setting
4328 | Patient Hose Compliance
4329 | Oximeter Averaging Interval Setting
4330 | System Alarm Status | Overall alert state of the system.
4332 |  | Consult detailed help for more information on this variable
4333 | Patient Alarm Priority | Overall Patient alarm priority. This is the current highest priority value of currently active Patient Alarms.
4335 |  | Consult detailed help for more information on this variable
4336 | Technical Alarm Priority | Overall Technical alarm priority. This is the current highest priority value of currently active Technical Alarms.
4338 |  | Consult detailed help for more information on this variable
4339 | Audio Alarm Priority | Overall audible alarm priority. This is the current highest priority value of currently active audible alarms.
4341 |  | Consult detailed help for more information on this variable
4342 | Technical Alarm Source ID
4343 | Technical Alarm Global State | Consult detailed help for more information on this variable
4344 | Technical Alarm Priority | Consult detailed help for more information on this variable
4345 | Technical Alarm Flag | Consult detailed help for more information on this variable
4346 | Technical Alarm Priority
4347 | Technical Alarm Description
4348 | Patient Alarm Source ID
4349 | Patient Alarm Global State | Consult detailed help for more information on this variable
4350 | Patient Alarm Priority | Consult detailed help for more information on this variable
4351 | Patient Alarm Flag | Consult detailed help for more information on this variable
4352 | Patient Alarm Priority
4353 | Patient Alarm Description
4354 | Device Nomenclature Revision | Format: major.minor (such as 1.0)
4355 | Device Text Catalog Revision | Format: "MM.NN mm.nn", where "MM.NN" is text catalog revision and "mm.nn" is language revision.
4356 | Device Data Format | Always 11 (Unicode).
4357 | Device Operating Area | Possible values: "unspecified", "operating room", "intensive care unit", "neonatal intensive care unit", and "cardiology care unit".
4358 | Patient O2 Consumption (VO2)
4359 | Patient CO2 Production (VCO2)
4360 | Arterial O2 Saturation (Left)
4361 | Arterial O2 Saturation (Right)
4362 | SpO2 Difference
4363 | Perfusion Indicator
4364 | Relative Perfusion (Left)
4365 | Relative Perfusion (Right)
4366 | Ao Systolic
4367 | Ao Diastolic
4368 | Ao Mean
4369 | RAP Systolic
4370 | RAP Diastolic
4371 | RAP Mean
4372 | Intracranial Pressure ICP Systolic
4373 | Intracranial Pressure ICP Diastolic
4374 | ICP Mean
4375 | UAP Systolic
4376 | UAP Diastolic
4377 | UAP Mean
4378 | UVP Systolic
4379 | UVP Diastolic
4380 | UVP Mean
4381 | Unspecific Pressure Systolic
4382 | Unspecific Pressure Diastolic
4383 | Unspecific Pressure Mean
4384 | Cerebral Perfusion Pressure (CPP)
4385 | Stroke Volume Variation
4386 | Index of Left Ventricular Contractility | This variable contains measurement from label "NLS_NOM_GRAD_PRESS_BLD_AORT_POS_MAX" and observed value "NOM_GRAD_PRESS_BLD_AORT_POS_MAX".
4387 | Aortic Blood Pressure (Ao) Wave
4388 | Auxiliary Waveform
4389 | Intermittent Mandatory Ventilation Status
4390 | Left Cardiac Work
4391 | Right Cardiac Work
4392 | Serum Creatinine
4393 | Urine Plasma Osmolarity Ratio
4394 | Urine Sodium Potassium Ratio
4395 | Urine Volume
4396 | Left Cardiac Work Index
4397 | Right Cardiac Work Index
4398 | Light Intensity (SvO2 Module)
4399 | SpO2 SvO2 Difference
4400 | Normalized Post Replacement Volume
4401 | Normalized Effluent Volume | This variable is associated with ChannelID identifier.
4402 | Normalized Total Replacement Volume (PBP) | This variable is associated with ChannelID identifier.
4403 | Normalized Replacement Volume (Pre)
4404 | Normalized Ultrafiltration Volume | This variable is associated with ChannelID identifier.
4405 | Patient Fluid Balance Limit Setting
4406 | Total Replacement Volume (PBP) | This variable is associated with ChannelID identifier.
4407 | Cumulated Liquid Volume (Pre) | This variable is associated with ChannelID identifier.
4408 | Filtration Fraction Setting
4409 | Diff.-Temperature between Tblood and Tskin
4410 | Diff.-Temperature between Tblood and Trect
4411 | Diff.-Temperature between Tblood and Tnaso
4412 | Diff.-Temperature between Tblood and Tesop
4413 | Diff.-Temperature between Tblood and Tcore
4414 | Diff.-Temperature between Tblood and Tart
4415 | Diff.-Temperature between Tblood and Tven
4416 | Diff.-Temperature between Tskin and Trect
4417 | Diff.-Temperature between Tskin and Tnaso
4418 | Diff.-Temperature between Tskin and Tesop
4419 | Diff.-Temperature between Tskin and Tcore
4420 | Diff.-Temperature between Tskin and Tart
4421 | Diff.-Temperature between Tskin and Tven
4422 | Diff.-Temperature between Trect and Tnaso
4423 | Diff.-Temperature between Trect and Tesop
4424 | Diff.-Temperature between Trect and Tcore
4425 | Diff.-Temperature between Trect and Tart
4426 | Diff.-Temperature between Trect and Tven
4427 | Diff.-Temperature between Tnaso and Tesop
4428 | Diff.-Temperature between Tnaso and Tcore
4429 | Diff.-Temperature between Tnaso and Tart
4430 | Diff.-Temperature between Tnaso and Tven
4431 | Diff.-Temperature between Tesop and Tcore
4432 | Diff.-Temperature between Tesop and Tart
4433 | Diff.-Temperature between Tesop and Tven
4434 | Diff.-Temperature between Tcore and Tart
4435 | Diff.-Temperature between Tcore and Tven
4436 | Diff.-Temperature between Tart and Tven
4437 | Newborn Body Length
4438 | Newborn Gestational Age
4439 | Calories
4440 | Input/Output Totals
4441 | Newborn Apgar Scores
4442 | BSA Formula Boyd
4443 | BSA Formula Dubois
4444 | Pupil Reaction (Left Eye) | Light reaction of left eye's pupil.
4445 | Pupil Reaction (Right Eye) | Light reaction of right eye's pupil.
4446 | Pleth Gain
4447 | Airway Temperature
4448 | Ventilation Perfusion Index
4449 | O2 Extraction Index
4450 | Duration Above Base Pressure
4451 | Inspiration Rise Time
4452 | Continuous Positive Airway Pressure (CPAP)
4453 | HFV Amplitude
4454 | Urine Urea
4455 | Urine pH Value
4456 | Total of CO2 (tCO2) | Result of Blood gas Analysis.
4457 | Total Bilirubin
4458 | Serum Glucose
4459 | Urine Glucose
4460 | Direct Bilirubin
4461 | Serum Calcium
4462 | Total Serum Calcium
4463 | Serum Magnesium
4464 | Serum Phosphat
4465 | Serum Kalium (Potassium)
4466 | Serum Clorid
4467 | Serum Albumine
4468 | Urine Clorid
4469 | Serum Globulin
4470 | Serum Protein
4471 | Serum Urea
4472 | Partial Thromboplastin Time
4473 | Prothrombin Time
4474 | Thrombin Time
4475 | Thromboplastine Time (Quick Time)
4476 | Alkalische Phosphatase
4477 | Alpha Amylase
4478 | Cholesterinesterase
4479 | Creatinin Kinase
4480 | Creatine Kinase of Type Muscle-Brain
4481 | Creatine Kinase of Type Muscle
4482 | Gamma Glutamyl Transpeptidase
4483 | Glutamic Oxaloacetic Transaminase
4484 | Glutamic Pyruvic Transaminase
4485 | Cholesterin
4486 | Triglyzeride
4487 | Urine Protein | (Total) Protein in Urine
4488 | Urine Calcium
4489 | Carboxy Hemoglobin
4490 | Total Protein
4491 | Lactate Dehydrogenase
4492 | Aspartin Aminotransferase
4493 | Alveolarproteinose (Alkaline Phosphatase)
4494 | Reticulocyte Count
4495 | Coagulation Time
4496 | Erithrocyte Sedimentation Rate
4497 | Packed Cell Volume
4498 | Kaolin Cephalin Time
4499 | Minimum Pressure During Breathing Cycle
4500 | Infusion Pump Accumulated Volume
4501 | Transthoracic Impedance
4502 | Inspired Expired N2O Difference
4503 | Inspired Expired Halothane Difference
4504 | Inspired Expired Enfluran Difference
4505 | Inspired Expired Isofluran Difference
4506 | Inspired Expired Sevofluran Difference
4507 | Inspired Expired Desfluran Difference
4508 | Inspired Expired Agent Difference
4509 | Inspired Expired O2 Difference
4510 | Patient Circuit O2 | O2 measured in the Patient Circuit
4511 | Barometric Pressure in Sample Chamber | This variable is obsolete that may be removed in a future release. It is duplicated with variable 4082.
4512 | Minute Volume
4513 | Mean Airway Pressure
4514 | Airway Respiration Rate
4515 | Inspiratory Airway Pressure Waveform
4516 | Inspiratory Flow Waveform
4517 | HFV Inspired Mandatory Minute Volume
4518 | HFV Gas Transport Coefficient Value
4519 | HFV Inspired Tidal Volume
4520 | HFV Tidal Volume
4521 | Generic Heart Rate | Heart Rate received from an external device
4522 | PERF from the 2nd SpO2/PLETH module
4523 | Percent Alpha (Left)
4524 | Percent Beta (Left)
4525 | Absolute Alpha (Left)
4526 | Absolute Beta (Left)
4527 | Absolute Delta (Left)
4528 | Absolute Theta (Left)
4529 | Mean Dominant Frequency (Left)
4530 | Median Power Frequency (Left)
4531 | Peak Power Frequency (Left)
4532 | Spectral Edge Frequency (Left)
4533 | Total Power (Left)
4534 | Scale of the Left Channel EEG Wave
4535 | Percent Alpha (Right)
4536 | Percent Beta (Right)
4537 | Absolute Alpha (Right)
4538 | Absolute Beta (Right)
4539 | Absolute Delta (Right)
4540 | Absolute Theta (Right)
4541 | Mean Dominant Frequency (Right)
4542 | Median Power Frequency (Right)
4543 | Peak Power Frequency (Right)
4544 | Spectral Edge Frequency (Right)
4545 | Total Power (Right)
4546 | Scale of the Right Channel EEG Wave
4547 | Duration of Positive Pressure
4548 | Fresh Gas N2O Concentration
4549 | Liquid Level (Anesthetic Agent Bottle)
4550 | Liquid Level (ISOflurane Bottle)
4551 | Liquid Level (ENFlurane Bottle)
4552 | Liquid Level (HALothane Bottle)
4553 | Liquid Level (DESflurane Bottle)
4554 | Liquid Level (SEVoflurane Bottle)
4555 | Invasive Blood Pressure Wave
4556 | Urine Shift Volume
4557 | Urine Flow Rate
4558 | Mean Dominant Frequency (CH1)
4559 | Mean Dominant Frequency (CH2)
4560 | Peak Power Frequency (CH1)
4561 | Peak Power Frequency (CH2)
4562 | % of Power in the Alpha Band (CH1)
4563 | % of Power in the Alpha Band (CH2)
4564 | % of Power in the Beta Band (1st EEG)
4565 | % of Power in the Beta Band (2nd EEG)
4566 | Hemoglobin
4567 | Patient Temperature
4568 | Capillaries pH Value
4569 | Adjusted pH (at Patient Temperature) | Adjusted pH at Patient Temperature
4570 | Adjusted Capillaries pH Value
4571 | Partial O2 in the Capillaries
4572 | Adjusted PO2 (at Patient Temperature) | Adjusted PO2 at Patient Temperature
4573 | Adjusted PcO2 (at Patient Temperature)
4574 | Venous CO2 Partial Pressure (PvCO2)
4575 | Partial CO2 in the Capillaries
4576 | Adjusted PcCO2 (at Patient Temperature)
4577 | Calculated Total CO2 (tCO2)
4578 | Calculated SO2 | Calculated SaO2
4579 | Calculated SvO2
4580 | Calculated ScO2
4581 | Calculated HCO3
4582 | Base Excess of Extra Cellular Fluid
4583 | Calculated Anion Gap
4584 | Urea | Measured by the i-Stat module.
4585 | Manually Entered FiO2 | In the Blood Review window.
4586 | Calculated Arterial Blood Base Excess
4587 | Ionized Magnesium
4588 | Desired Module Temperature | It's the selected/adjusted Temp in i-Stat Window.
4589 | Calculated BUN/Creatinine Ratio | Calculated by the i-Stat module.
4590 | Calculated Urea/Creatinine Ratio | Calculated by the i-Stat module.
4591 | Lactate | Measured value by the i-Stat module.
4592 | Device Simulator Status | Device data with the values: Passed and Failed.
4593 | Pressure of Sensor 1 | Pressure value of sensor 1. Device specific data.
4594 | Pressure of Sensor 2 | Pressure value of sensor 2. Device specific data.
4595 | Module Temperature
4596 | Air Temperature
4597 | Micro Environment Air Temperature Setting
4598 | Weight of the Urine Disposable Bag
4599 | Total Urine Volume of the Current Measurement Period
4600 | Urine Fluid Density
4601 | Urine Fluid Temperature
4602 | Ventilation Activated Status | Parameter to Informs whether the Patient is Ventilated
4603 | Simplified Acuity Physiologic Score
4604 | Value Normal | Missing parameters in SAPS parameter set which are considered in normal range.
4605 | Aspartate Aminotransferase
4606 | Spontaneous Respiration Rate
4607 | Current Fluid in Urine Bag
4608 | Cardiac Function Index (Thermodilution)
4609 | Intrathoracic Blood Volume (Thermodilution)
4610 | Intrathoracic Blood Volume Index (Thermodilution)
4611 | Extravascular Lung Water (Thermodilution)
4612 | Extravascular Lung Water Index (Thermodilution)
4613 | NMT TOF Count
4614 | NMT TOF Ratio
4615 | Twitch Height of the 1Hz/0.1Hz Stimulation Response
4616 | Post Tetatic Count
4617 | Remaining Time to Next Stimulation
4618 | TOF Response TOF1
4619 | TOF Response TOF2
4620 | TOF Response TOF3
4621 | TOF Response TOF4
4622 | Activated Partial Thromboplastin Time (aPTT Whole Blood Time)
4623 | Activated Partial Thromboplastin Time (aPTT Plasma Equivalent Time)
4624 | Prothrombin Time (Whole Blood Time)
4625 | Prothrombin Time (Plasma Equivalent Time)
4626 | Prothrombin Time Ratio
4627 | Spontaneous Inspiration Time
4628 | Overdistension Index
4629 | Lung Time Constant
4630 | Airway Correlation Coefficient
4631 | Rate Volume Ratio
4632 | Normalized Ionized Calcium (Ca)
4633 | Peak Inspiratory Flow Setting
4634 | Trigger Sensitivity Setting
4635 | Plateau Time Setting
4636 | Multiple Sigh Number Setting
4637 | Apnea Tidal Volume Setting
4638 | Apnea Respiration Rate Setting
4639 | Apnea Peak Flow Setting
4640 | Apnea Inspired O2 Setting
4641 | PSV Pressure Setting | Pressure Support Ventilation Setting
4642 | Sigh Activation Status
4643 | Suction O2 Concentration Setting
4644 | PCV Control Pressure Setting
4645 | PCV Inspiratory Time Setting
4646 | PCV Apnea Respiration Rate Setting
4647 | PCV Apnea Inspiratory Time Setting
4648 | PCV Apnea O2 Setting
4649 | PCV Apnea High Airway Pressure Setting
4650 | Inspiratory Component of I:E Ratio Setting | Pressure Ventilation I component of I:E Ratio setting
4651 | Expiratory Component of I:E Ratio Setting | Pressure Ventilation E component of I:E Ratio setting
4652 | Inspiratory Component of Apnea Ventilation I:E Ratio Setting
4653 | Expiratory Component of Apnea Ventilation I:E Ratio Setting
4654 | Setting: Pressure Ventilation Control
4655 | Cycle Time Setting
4656 | IPPV Ventilation Respiration Rate Setting
4657 | IMV Ventilation Respiration Rate Setting
4658 | Pressure Support PEEP Setting
4659 | Apnea Ventilation Delay
4660 | O2 Monitoring Mode
4661 | O2 Calibration Mode
4662 | Inspiratory Time Setting
4663 | Expiratory Time Setting
4664 | Alarm Percentage on Rise Time Setting
4665 | Continuous Positive Airway Pressure (CPAP) Setting | This is the device's measurement with label="NLS_NOM_SETT_CONC_AWAY_O2" and physio_id="NOM_CONC_AWAY_O2".
4666 | Ventilator Flow Setting
4667 | Low Inspiratory Pressure Setting
4668 | HFV Flow Setting
4669 | HFV Respiration Rate Setting
4670 | Mechanical Ventilation Activated Status | Controlled Mechanical Ventilation Setting
4671 | SIMV Mode | Synchronized Intermittent Mandatory Ventilation Setting
4672 | Mandatory Inspired Minute Volume Setting
4673 | Infusion Pump Delivery Rate Setting
4674 | Airway Respiration Rate Setting
4675 | Trigger Flow/Volume Setting | This is the device's measurement with label="NLS_NOM_SETT_VENT_VOL_LUNG_TRAPD" and physio_id="NOM_SETT_VENT_VOL_LUNG_TRAPD".
4676 | Apnea Alarm Delay
4677 | HFV Amplitude Setting
4678 | Volume Alarm Delay
4679 | Pressure Increase Rate Setting
4680 | Volume Warning Setting | Delivered by the Evita 2 VueLink Driver.
4681 | Airway Inspiratory Pressure Lower Alarm Limit
4682 | Apnea Alarm Delay Setting
4683 | Pressure Ramp Setting (ASB Ramp)
4684 | Applied Tidal Volume Setting
4685 | Assist Pressure Sensitivity Setting | Assist Sensitivity. Used by the Bear 1000 ventilator.
4686 | High Inspiratory Pressure Setting
4687 | Background Flow Setting
4688 | Vaporizer Agent Concentration Setting
4689 | Vaporizer Isoflurane Concentration Setting
4690 | Vaporizer Enflurane Concentration Setting
4691 | Vaporizer Halothane Concentration Setting
4692 | Vaporizer Desflurane Concentration Setting
4693 | Vaporizer Sevoflurane Concentration Setting
4694 | Gas Measurement Sensor Position
4695 | O2 Measurement Sensor Position
4696 | Inspired Tidal Volume Setting
4697 | Incubator/Warmer Temperature Setting
4698 | Preset Period of UrVol Setting
4699 | Low Time Setting | Setting variable for Evita 4 Airway Pressure Release Ventilation Mode.
4700 | High Time Setting | Setting variable for Evita 4 Airway Pressure Release Ventilation Mode.
4701 | Low Pressure Setting | Setting variable for Evita 4 Airway Pressure Release Ventilation Mode.
4702 | High Pressure Setting | Setting variable for Evita 4 Airway Pressure Release Ventilation Mode.
4703 | Volume Assist Level Setting
4704 | CPAP Flow Assist Level Setting
4705 | Preset Stimulation Current Setting
4706 | Preset Stimulation Charge Setting
4707 | Preset Stimulation Impulse Duration Setting
4708 | NMT TOF Repetition Time | Preset Train of Four (Slow TOF) Repetition Time Setting
4709 | Panting Limit Setting
4710 | Airway Inspiratory Pressure Upper Alarm Limit | Alarm Limit: High Pressure
4711 | Sigh Volume
4712 | Sustained Pressure Alarm Limit
4713 | Intrathoracic Blood Volume (ITBV)
4714 | Intrathoracic Blood Volume Index (ITBI)
4715 | Extravascular Lung Water
4716 | Extravascular Lung Water Index
4717 | Global End Diastolic Volume
4718 | Global End Diastolic Volume Index
4719 | Cardiac Function Index
4720 | Mixed Venous O2 Saturation
4721 | Generic Term for the Transcutaneous Gases
4722 | Intrinsic PEEP
4723 | Positive Peak Inspiratory Pressure
4724 | Dead Space to Tidal Volume Ratio
4725 | Lung Volume Trapped
4726 | Lung Vital Capacity
4727 | Generic Lung Compliance
4728 | EEG Electromyography
4729 | Suppression Ratio
4730 | EEG Peak Power Frequency
4731 | End Tidal N2
4732 | Inspired N2
4733 | Generic Agent Wave
4734 | End Tidal MAC Concentration
4735 | Left Ventricular Stroke Volume Index
4736 | Left Ventricular Stroke Volume
4737 | Right Ventricular Stroke Volume
4738 | Glasgow Coma Score
4739 | SubScore of the GCS - Eye Response
4740 | SubScore of the GCS - Motoric Response
4741 | SubScore of the GCS - Verbal Response
4742 | Head Circumference
4743 | pH in Arterial Blood
4744 | Arterial Blood Hemoglobin (HB)
4745 | pH in Venous Blood
4746 | Urine Natrium
4747 | Natrium in Serum
4748 | pH in the Blood Plasma
4749 | Blood Natrium (Sodium)
4750 | Kalium (Potassium)
4751 | Generic Glucose Parameter
4752 | Partial CO2
4753 | Partial O2
4754 | Base Excess of Blood
4755 | Percent Alveolarvenous Shunt Qs/Qt
4756 | Extravascular Thermal Volume Index
4757 | VPB Rate
4758 | ECG Lead V1 Wave
4759 | ECG Lead V2 Wave
4760 | ECG Lead V3 Wave
4761 | ECG Lead V4 Wave
4762 | ECG Lead V5 Wave
4763 | ECG Lead V6 Wave
4764 | ECG Lead I Wave
4765 | ECG Lead II Wave
4766 | ECG Lead III Wave
4767 | ECG Lead aVR Wave
4768 | ECG Lead aVL Wave
4769 | ECG Lead aVF Wave
4770 | Mean Dominant Frequency (CH3)
4771 | Mean Dominant Frequency (CH4)
4772 | NMT DBS Count
4773 | Peak Power Frequency (CH3)
4774 | Peak Power Frequency (CH4)
4775 | Pump Infusion Mode
4776 | Remaining Time Counter | Time to Elapse Counter used by the i-Stat module.
4777 | APRV Low Pressure Synchronisation Time Setting
4778 | APRV High Pressure Synchronisation Time Setting
4779 | Acoustic Respiration Rate
4780 | Regional O2 Saturation (rSO2 2)
4781 | Regional O2 Saturation (rSO2 3)
4782 | Regional O2 Saturation (rSO2 4)
4783 | Regional O2 Saturation (rSO2 5)
4784 | Regional O2 Saturation (rSO2 6)
4785 | O2 Delivery Index (Thermodilution)
4786 | Noninvasive Venous Hemoglobin (SpHbv)
4787 | Noninvasive Arterial Hemoglobin (SpHba)
4788 | Generic Pulse Rate
4789 | Corrected Total Hemoglobin
4790 | Corrected Arterial Total Hemoglobin
4791 | Corrected Venous Total Hemoglobin
4792 | Systemic Vascular Resistance (Thermodilution)
4793 | Systemic Vascular Resistance Index (Thermodilution)
4794 | Pulse Rate (Abdominal)
4795 | ECG Lead V3R
4796 | ECG Lead V4R
4797 | ECG Lead V5R
4798 | Inspired NO
4799 | ECG Lead V9
4800 | ECG Lead V7
4801 | ECG Lead V8
4802 | Apnea Ventilation I:E Ratio Setting | This variable is associated with ChannelID identifier. This variable is calculated from variable 4977 and 4978, which is sent as a string in the format "1:xx.xx".
4803 | Assisted Respiration Rate
4804 | Correlation Coefficient Squared
4805 | ECG Lead V6R
4806 | Citrate Solution Concentration
4807 | Calcium Solution Concentration
4808 | Citrate Acid Concentration
4809 | Citrate Load Setting
4810 | Backup Respiration Rate Setting
4811 | Inspired Expired Pressure Difference
4812 | Backup Inspired Expired Pressure Difference Setting
4813 | Femoral Arterial Pressure Mean
4814 | Femoral Arterial Pressure Systolic
4815 | Femoral Arterial Pressure Diastolic
4816 | Tidal Volume Lower Alarm Limit
4817 | Minute Volume Lower Alarm Limit
4818 | ECG Lead V7 Wave
4819 | ECG Lead V8 Wave
4820 | ECG Lead V9 Wave
4821 | Airway Aux Wave
4822 | Fetal DECG Wave
4823 | Patient BSA Formula | Consult detailed help for more information on this variable
4824 | Platelets (Thrombocyte Count)
4825 | Endotracheal Tube Type
4826 | Minute Volume Upper Alarm Limit
4827 | Pulse Rate from Plethysmogram
4828 | ECG Ectopic Status | Consult detailed help for more information on this variable
4829 | ST Lead ES
4830 | ST Lead AS
4831 | ST Lead AI
4832 | ST Index
4833 | Pulse Pressure
4834 | Mean Transit Time (MTt)
4835 | Alveolar Dead Space
4836 | Tympanic Temperature
4837 | Injectate Volume (CO)
4838 | SpO2 (Telemetry)
4839 | Pulse Rate (Telemetry)
4840 | Intracranial Pressure ICP1 (Systolic)
4841 | Intracranial Pressure ICP1 (Diastolic)
4842 | Intracranial Pressure ICP1 (Mean)
4843 | Intracranial Pressure ICP2 Systolic
4844 | Intracranial Pressure ICP2 Diastolic
4845 | Intracranial Pressure ICP2 Mean
4846 | Brachial Arterial Pressure (BA) Systolic
4847 | Brachial Arterial Pressure (BA) Diastolic
4848 | Brachial Arterial Pressure (BA) Mean
4849 | Vesical Temperature (Bladder)
4850 | Cerebral Temperature
4851 | ST Level for Lead I
4852 | ST Level for Lead II
4853 | ST Level for Lead III
4854 | ST Level for Lead aVR
4855 | ST Level for Lead aVL
4856 | ST Level for Lead aVF
4857 | ST Level for Lead V1
4858 | ST Level for Lead V2
4859 | ST Level for Lead V3
4860 | ST Level for Lead V4
4861 | ST Level for Lead V5
4862 | ST Level for Lead V6
4863 | Prothrombin Time (International Normalized Ratio)
4864 | PSV Tidal Volume | Tidal Volume in Pressure Support Ventilation Mode
4865 | Negative Inspiratory Force Index
4866 | A-Line ARX Index
4867 | Impedance Cardiography (ICG) Wave
4868 | Thoracic Fluid Content
4869 | Thoracic Fluid Content Index
4870 | Accelerated Cardiac Index
4871 | Heart Contractility Index
4872 | O2 Saturation (Left)
4873 | O2 Saturation (Right)
4874 | Pulse Pressure Variation
4875 | Downslope Time
4876 | Pulmonary Vascular Permeability Index | This is the device's measurement with label="NLS_NOM_PERM_VASC_PULM_INDEX" and physio_id="NOM_PERM_VASC_PULM_INDEX". See also variable 8791.
4877 | ART Pulse Rate | ART pulse rate
4878 | Umbilical Arterial Pressure Pulse Rate
4879 | Patient Visit Number | Patient encounter ID contains the ID of the current visit of the patient
4880 | Patient Height at Entered Time
4881 | Patient Weight at Entered Time
4882 | Patient Body Surface Area (BSA)
4883 | Inspired O2 Setting (FiO2)
4884 | Inspiratory Component of I:E Ratio Setting
4885 | Expiratory Component of I:E Ratio Setting
4886 | Lung Elastance
4887 | Backup Inspiratory Time Setting
4888 | Pressure Support Level Above Phigh
4889 | Total Respiration Duration Setting
4890 | Heater Power
4891 | Capillary Refill Time
4892 | Estimated Patient Circuit Leakage
4893 | Estimated Ventilation Total Leakage
4894 | Ventilation Mode Additional Information
4895 | Backup Plateau Time Setting
4896 | ECG Lead V3R Wave
4897 | ECG Lead V4R Wave
4898 | ECG Lead V5R Wave
4899 | ECG Lead V6R Wave
4900 | Signal Quality Index (BIS) | This variable is associated with ChannelID identifier.
4901 | Continuous NIBP Systolic
4902 | Continuous NIBP Diastolic
4903 | Continuous NIBP Mean
4904 | Next Intervention Time
4905 | Total Power (CH3)
4906 | Total Power (CH4)
4907 | Total Power (Left)
4908 | Total Power (Right)
4909 | PCV I:E Ratio Setting | This variable is associated with ChannelID identifier. This variable is calculated from variable 4975 and 4976, which is sent as a string in the format "1:xx.xx".
4910 | Inspiratory CO2 (FiCO2)
4911 | End Tidal NO2
4912 | Fresh Gas Total Flow
4913 | Bolus Volume Setting
4914 | Generic NIBP Pressure | Available on IntelliVue devices with firmware revision G or later.
4915 | Endotracheal Tube Diameter
4916 | Mandatory Respiration Rate
4917 | Mandatory Minute Volume
4918 | Mandatory Tidal Volume
4919 | N2O Consumption
4920 | Device Power Up Time | Number of seconds since the Monitor device's last power-on.
4921 | O2 Consumption
4922 | Air Consumption
4923 | Arterial Blood Pressure (ABP) | Available on IntelliVue devices with firmware revision G or later. Arterial Blood Pressure (NOM_PRESS_BLD_ART_ABP).
4924 | ABP Pulse Rate | ABP pulse rate
4925 | Signal Noise Ratio (CCO)
4926 | Spontaneous Airway Respiration Rate
4927 | Synchronized Respiration Rate
4928 | Mechanical Respiration Rate
4929 | Daily Urine Output
4930 | Activated Partial Thromboplastin Time Ratio
4931 | Global Ejection Fraction | This is the device's measurement with label="NLS_NOM_FRACT_EJECT" and physio_id="NOM_FRACT_EJECT". See also 8790.
4932 | Generic CO2 Parameter | This variable contains "PECO2, Partial expired CO2", which is measured from label="NLS_NOM_VENT_AWAY_CO2_EXP", physio_id="NOM_VENT_AWAY_CO2_EXP".
4933 | Central Venous O2 Saturation (ScvO2)
4934 | Suppression Ratio (BIS) | This variable is associated with ChannelID identifier.
4935 | Right Ventricular Stroke Work Index
4936 | Base Excess of Extracellular Fluid
4937 | Rise Time Setting
4938 | Time Duration at Low Pressure (Tlow)
4939 | Time Duration at High Pressure (Thigh)
4940 | Spontaneous Tidal Volume
4941 | Cardiac Troponin l
4942 | Cardiopulmonary Bypass Status
4943 | Cardiac Brain Natriuretic Peptide
4944 | Mean Dominant Frequency
4945 | Filtration Current Volume
4946 | Calcium Flow Rate Setting
4947 | Calcium Compensation Setting
4948 | Cerebral State Index (CSI)
4949 | EEG Amplitude | This variable is associated with ChannelID identifier.
4950 | Patient Middle Names | The Middle Name attribute contains the middle name of the patient.
4951 | Ao Pulse Rate
4952 | PAP Pulse Rate
4953 | CVP Pulse Rate
4954 | RAP Pulse Rate
4955 | LAP Pulse Rate
4956 | ICP Pulse Rate
4957 | Umbilical Venous Pressure Pulse Rate
4958 | FAP Pulse Rate
4959 | BAP Pulse Rate
4960 | IC1 Pulse Rate
4961 | Intracranial Pressure ICP2 Pulse Rate
4962 | Pulse Rate (P1)
4963 | Pulse Rate (P2)
4964 | Pulse Rate (P3)
4965 | Pulse Rate (P4)
4966 | Pulse Rate from Plethysmogram (Left)
4967 | Pleth Wave (Right)
4968 | Pleth Wave (Left)
4969 | Calculated SO2
4970 | Airway Respiration Rate (Ohmeda)
4971 | Airway Respiration Rate Setting (Ohmeda)
4972 | Fresh Gas O2 Concentration
4973 | Manually entered FIO2 or Flow | This variable contains measurement from label "NLS_NOM_CONC_AWAY_O2" and observed value "NOM_CONC_AWAY_O2".
4974 | Generic O2 Parameter | This is the device's measurement with label="NLS_NOM_SETT_CONC_AWAY_O2" and physio_id="NOM_CONC_AWAY_O2".
4975 | Fresh Gas O2 Flow Setting | Mixer parameter.
4976 | Airway Pressure Numeric Value (Ohmeda)
4977 | Inspiratory Peak Flow
4978 | Maximum Pressure During Breathing Cycle
4979 | Peak Pressure
4980 | Peak Inspiratory Flow Setting
4981 | Airway Maximum Pressure Setting
4982 | Peak Inspiratory Pressure Setting
4983 | Setting: Inspiratory Pressure (Ohmeda)
4984 | Mean Airway Pressure (Ohmeda)
4985 | EEG Wave (Left)
4986 | EEG Wave (Right)
4987 | % of Power in the Delta Band - 1st EEG
4988 | % of Power in the Delta Band - 2nd EEG
4989 | Percent Delta - Left Side
4990 | Percent Delta - Right Side
4991 | % of Power in the Theta Band (1st EEG)
4992 | % of Power in the Theta Band (2nd EEG)
4993 | Percent Theta - Left Side
4994 | Percent Theta - Right Side
4995 | Generic Calculation Constant | This variable contains measurement from Philips label "NLS_NOM_NUM_CALC_CONST" and observed value "NOM_METRIC_NOS". See also variable 8493.
4996 | Pressure Ventilation Control Pressure
4997 | Right Heart Ejection Fraction
4998 | White Blood Count (Leucocyte Count)
4999 | Red Blood Count (Erithrocyte Count)
5000 | Mean Corpuscular Volume
5001 | Corpuscular Mean Hemoglobin Content
5002 | Ferrum
5003 | Fetal Hemoglobin
5004 | MetHemoglobin
5005 | Airway Volume Wave
5006 | Expiratory Peak Flow
5007 | Expiratory Peak Flow
5008 | The V stands for Volume
5009 | The R stands for Rate
5010 | Adjusted PaCO2 (at Patient Temperature)
5011 | Incubator Heating Power | Power Required to set Air&Pat Temp in the Incubator
5012 | Actual Patient Age | Actual patient age, provided by a module, measured in years. Please note that this Patient Age will be available only when the corresponding module is connected to the Monitor. Variable 169 contains the patient age as provided by the Monitor itself (calculated from the birth date entered on the Monitor device.)
5013 | The P stands for Pressure
5014 | PCV Apnea Control Pressure Setting
5015 | Fresh Gas Air Flow Setting | Mixer parameter.
5016 | PEEP/CPAP Lower Alarm Limit
5017 | Tidal Volume Upper Alarm Limit
5018 | T-Tre Temperature difference between Temp and Trect
5019 | T-Tbl Temperature difference between Temp and Tblood
5020 | T-Tco Temperature difference between Temp and Tcore
5021 | T-Tsk Temperature difference between Temp and Tskin
5022 | T-Tes Temperature difference between Temp and Tesoph
5023 | T-Tna Temperature difference between Temp and Tnaso
5024 | T-Tar Temperature difference between Temp and Tart
5025 | T-Tve Temperature difference between Temp and Tven
5026 | T-Taw Temperature difference between Temp and Tairwy
5027 | Tre-T Temperature difference between Trect and Temp
5028 | Tre-bl Temperature difference between Trect and Tblood
5029 | Tre-sk Temperature difference between Trect and Tskin
5030 | Tre-aw Temperature difference between Trect and Tairwy
5031 | Temperature difference between Tblood and Temp
5032 | Tbl-aw Temperature difference between Tblood and Tairwy
5033 | Tco-T Temperature difference between Tcore and Temp
5034 | Tco-re Temperature difference between Tcore and Trect
5035 | Tco-bl Temperature difference between Tcore and Tblood
5036 | Tco-sk Temperature difference between Tcore and Tskin
5037 | Tco-es Temperature difference between Tcore and Tesoph
5038 | Tco-na Temperature difference between Tcore and Tnaso
5039 | Tco-aw Temperature difference between Tcore and Tairwy
5040 | Tsk-T Temperature difference between Tskin and Temp
5041 | Tsk-bl Temperature difference between Tskin and Tblood
5042 | Tsk-aw Temperature difference between Tskin and Tairwy
5043 | Tes-T Temperature difference between Tesoph and Temp
5044 | Tes-re Temperature difference between Tesoph and Trect
5045 | Tes-bl Temperature difference between Tesoph and Tblood
5046 | Tes-sk Temperature difference between Tesoph and Tskin
5047 | Tes-na Temperature difference between Tesoph and Tnaso
5048 | Tes-aw Temperature difference between Tesoph and Tairwy
5049 | Tna-T Temperature difference between Tnaso and Temp
5050 | Tna-re Temperature difference between Tnaso and Trect
5051 | Tna-bl Temperature difference between Tnaso and Tblood
5052 | Tna-sk Temperature difference between Tnaso and Tskin
5053 | Tna-aw Temperature difference between Tnaso and Tairwy
5054 | Tar-T Temperature difference between Tart and Temp
5055 | Tar-re Temperature difference between Tart and Trect
5056 | Tar-bl Temperature difference between Tart and Tblood
5057 | Tar-co Temperature difference between Tart and Tcore
5058 | Tar-sk Temperature difference between Tart and Tskin
5059 | Tar-es Temperature difference between Tart and Tesoph
5060 | Tar-na Temperature difference between Tart and Tnaso
5061 | Tar-aw Temperature difference between Tart and Tairwy
5062 | Tve-T Temperature difference between Tven and Temp
5063 | Tve-re Temperature difference between Tven and Trect
5064 | Tve-bl Temperature difference between Tven and Tblood
5065 | Tve-co Temperature difference between Tven and Tcore
5066 | Tve-sk Temperature difference between Tven and Tskin
5067 | Tve-es Temperature difference between Tven and Tesoph
5068 | Tve-na Temperature difference between Tven and Tnaso
5069 | MAC Concentration (Awake)
5070 | Tve-ar Temperature difference between Tven and Tart
5071 | Tve-aw Temperature difference between Tven and Tairwy
5072 | Taw-T Temperature difference between Tairwy and Temp
5073 | Taw-re Temperature difference between Tairwy and Trect
5074 | Taw-bl Temperature difference between Tairwy and Tblood
5075 | Taw-co Temperature difference between Tairwy and Tcore
5076 | Taw-sk Temperature difference between Tairwy and Tskin
5077 | Taw-es Temperature difference between Tairwy and Tesoph
5078 | Taw-na Temperature difference between Tairwy and Tnaso
5079 | Taw-ar Temperature difference between Tairwy and Tart
5080 | Taw-ve Temperature difference between Tairwy and Tven
5081 | Generic Agent 1 Wave
5082 | End Tidal Generic Agent 1
5083 | Inspired Generic Agent 1
5084 | Generic Agent 2 Wave
5085 | End Tidal Generic Agent 2
5086 | Inspired Generic Agent 2
5087 | Adjusted Arterial Blood pH Value
5088 | Adjusted Venous Blood pH Value
5089 | Adjusted PvCO2 (at Patient Temperature)
5090 | Adjusted PvO2 (at Patient Temperature)
5091 | Generic Calculation Constant | This variable contains parameter from Philips label "NLS_NOM_NUM_CATHETER_CONST" and observed value "NOM_NUM_CATHETER_CONST". See also variable 8395.
5092 | ECG Lead V Wave
5093 | ECG Lead MCL Wave
5094 | ECG Lead MCL1 Wave
5095 | MAC Concentration
5096 | Pleth Wave (Telemetry)
5097 | NIBP Systolic (Telemetry)
5098 | NIBP Diastolic (Telemetry)
5099 | NIBP Mean (Telemetry)
5100 | Generic Pressure 1 (P1) Systolic
5101 | Generic Pressure 1 (P1) Diastolic
5102 | Generic Pressure 1 (P1) Mean
5103 | Generic Pressure 2 (P2) Systolic
5104 | Generic Pressure 2 (P2) Diastolic
5105 | Generic Pressure 2 (P2) Mean
5106 | Generic Pressure 3 (P3) Systolic
5107 | Generic Pressure 3 (P3) Diastolic
5108 | Generic Pressure 3 (P3) Mean
5109 | Generic Pressure 4 (P4) Systolic
5110 | Generic Pressure 4 (P4) Diastolic
5111 | Generic Pressure 4 (P4) Mean
5112 | Incubator Humidity
5113 | Predictive Oral Temperature
5114 | Predictive Rectal Temperature
5115 | Predictive Axillary Temperature
5116 | Perfusion (Telemetry)
5117 | Right to Left Heart Shunt
5118 | Adjusted PcO2 (at Patient Temperature)
5119 | Adjusted Arterial PaO2 (at Patient Temperature)
5120 | Pulse Rate from Plethysmogram (Right)
5121 | Patient O2 Consumption per Body Surface
5122 | Patient O2 Consumption per Body Weight
5123 | Ideal Patient Weight
5124 | End Flow Setting
5125 | PSV Backup Delay Setting
5126 | Technical Alarm Internal Source ID | Consult detailed help for more information on this variable
5127 | Technical Alarm Source Type
5128 | Device Technical Alarm | Consult detailed help for more information on this variable
5129 | Patient Alarm Internal Source ID | Consult detailed help for more information on this variable
5130 | Patient Alarm Source Type
5131 | Patient Alarm | Consult detailed help for more information on this variable
5132 | Accoustic Respiration Rate Wave
5133 | Blood Temperature Wave
5134 | Non-specific Pressure 1 Wave
5135 | Generic Invasive Blood Pressure P1 | Available on IntelliVue devices.
5136 | Generic Invasive Blood Pressure P1 Systolic
5137 | Generic Invasive Blood Pressure P1 Diastolic
5138 | Generic Invasive Blood Pressure P1 Mean
5139 | Non-specific Pressure 2 Wave
5140 | Generic Invasive Blood Pressure P2 | Available on IntelliVue devices.
5141 | Generic Invasive Blood Pressure P2 Systolic
5142 | Generic Invasive Blood Pressure P2 Diastolic
5143 | Generic Invasive Blood Pressure P2 Mean
5144 | Non-specific Pressure 3 Wave
5145 | Generic Invasive Blood Pressure P3 | Available on IntelliVue devices.
5146 | Generic Invasive Blood Pressure P3 Systolic
5147 | Generic Invasive Blood Pressure P3 Diastolic
5148 | Generic Invasive Blood Pressure P3 Mean
5149 | Non-specific Pressure 4 Wave
5150 | Generic Invasive Blood Pressure P4 | Available on IntelliVue devices.
5151 | Generic Invasive Blood Pressure P4 Systolic
5152 | Generic Invasive Blood Pressure P4 Diastolic
5153 | Generic Invasive Blood Pressure P4 Mean
5154 | Non-specific Pressure 5 Wave
5155 | Generic Invasive Blood Pressure P5 | Available on IntelliVue devices.
5156 | Generic Invasive Blood Pressure P5 Systolic
5157 | Generic Invasive Blood Pressure P5 Diastolic
5158 | Generic Invasive Blood Pressure P5 Mean
5159 | Non-specific Pressure 6 Wave
5160 | Generic Invasive Blood Pressure P6 | Available on IntelliVue devices.
5161 | Generic Invasive Blood Pressure P6 Systolic
5162 | Generic Invasive Blood Pressure P6 Diastolic
5163 | Generic Invasive Blood Pressure P6 Mean
5164 | Non-specific Pressure 7 Wave
5165 | Generic Invasive Blood Pressure P7 | Available on IntelliVue devices.
5166 | Generic Invasive Blood Pressure P7 Systolic
5167 | Generic Invasive Blood Pressure P7 Diastolic
5168 | Generic Invasive Blood Pressure P7 Mean
5169 | Non-specific Pressure 8 Wave
5170 | Generic Invasive Blood Pressure P8 | Available on IntelliVue devices.
5171 | Generic Invasive Blood Pressure P8 Systolic
5172 | Generic Invasive Blood Pressure P8 Diastolic
5173 | Generic Invasive Blood Pressure P8 Mean
5174 | Intra-Uterine Pressure (IUP) Wave
5175 | Intra-Uterine Pressure
5176 | Intra-Uterine Pressure Systolic
5177 | Intra-Uterine Pressure Diastolic
5178 | Intra-Uterine Pressure Mean
5179 | Unspecified Blood Pressure
5180 | Unspecified Blood Pressure Systolic
5181 | Unspecified Blood Pressure Diastolic
5182 | Unspecified Blood Pressure Mean
5183 | Airway Flow Wave
5184 | Auxiliary Wave/Parameter
5185 | Global Ejection Fraction (2) | This is the device's measurement with label="NLS_NOM_EMFC_GEF" and physio_id="NOM_METRIC_NOS". See also 7810.
5186 | Pulmonary Vascular Permeability Index (2) | This is the device's measurement with label="NLS_NOM_EMFC_PVPI" and physio_id="NOM_METRIC_NOS". See also variable 6493.
5187 | Pressure Drop
5188 | Replacement Pre-delivered Rate Setting
5189 | Replacement Post-delivered Rate Setting
5190 | Pre-blood Flow Rate Setting
5191 | Patient Weight Balance Setting
5192 | Anticoagulant Bolus Interval Setting
5193 | Anticoagulant Infusion Rate Setting
5194 | Ultrafiltration Ratio Setting
5195 | Therapy Elapsed Time
5196 | Airway Volume
5197 | Processed Cumulated Blood Volume
5198 | Replacement Pre-delivered Cumulated Volume
5199 | Replacement Post-delivered Cumulated Volume | This variable is associated with ChannelID identifier.
5200 | Pre-blood Cumulated Volume | This variable is associated with ChannelID identifier.
5201 | Syringe Speed
5202 | Maximum Volume Gain Setting
5203 | Patient Fluid Balance
5204 | Arterial Blood Pressure (ABP) Wave
5205 | Arterial Blood Pressure (ART) Wave
5206 | Pulmonary Arterial Pressure (PAP) Wave
5207 | Central Venous Pressure (CVP) Wave
5208 | Right Atrial Pressure (RAP) Wave
5209 | Left Atrial Pressure (LAP) Wave
5210 | Intra-cranial Pressure (ICP) Wave
5211 | Umbilical Arterial Pressure (UAP) Wave
5212 | Umbilical Venous Pressure (UVP) Wave
5213 | Airway CO2 Level Waveform
5214 | Airway O2 Wave
5215 | Airway N2 Wave
5216 | Generic Agent Wave (2nd)
5217 | Impedance Respiration Wave
5218 | Brachial Arterial Pressure (BAP) Wave
5219 | Intracranial Pressure 1 (IC1) Wave
5220 | Intracranial Pressure 2 (IC2) Wave
5221 | Femoral Arterial Pressure (FAP) Wave
5222 | Vector ECG Wave
5223 | Transcutaneous Heater Power
5224 | BIS Standard Deviation | This variable is associated with ChannelID identifier.
5225 | EMG Standard Deviation (BIS) | This variable is associated with ChannelID identifier.
5226 | Asymmetric Index | This variable is associated with ChannelID identifier.
5227 | O2 Consumption Index Estimate
5228 | Aortic Blood Pressure (Ao) | Available on IntelliVue devices with firmware revision G or later. Arterial Blood Pressure in the Aorta (NOM_PRESS_BLD_AORT).
5229 | Umbilical Arterial Pressure (UAP) | Available on IntelliVue devices with firmware revision G or later.
5230 | Umbilical Venous Pressure (UVP) | Available on IntelliVue devices with firmware revision G or later.
5231 | Humidifier Activation Status
5232 | Nebulizer Time Setting
5233 | Edi Diaphragm Electrical Activity Waveform
5234 | Peak Diaphragm Electrical Activity (Edi)
5235 | Minimum Diaphragm Electrical Activity (Edi)
5236 | % Edi Trigger In NAVA
5237 | % Edi Cycle Off In NAVA
5238 | NAVA Level Setting
5239 | Edi Trigger Setting
5240 | Patient Body Mass Index (BMI)
5241 | Maximum Positive Auxiliary Pressure
5242 | Mean Positive Auxiliary Pressure
5243 | Minimum Positive Auxiliary Pressure
5244 | Patient CO2 Production Indexed By Body Surface
5245 | Patient CO2 Production Indexed By Body Weight
5246 | Tidal Volume Indexed By Body Weight
5247 | Functional Residual Capacity
5248 | Corrected QT Interval (QTc)
5249 | Airway N2 Concentration | Measurement from VueLink
5250 | Airway Agent Concentration | Measurement from VueLink. This variable is associated with a SourceID identifier.
5251 | Respiration Rate (Spirometry)
5252 | QT Heart Rate (QT-HR) | QT heart rate
5253 | Delta QTc
5254 | Humidification Type Setting
5255 | Normalized Rapid Shallow Breathing Index
5256 | Percent of Spontaneous Inspiratory Time
5257 | Energy Balance Calculated by Metabolics
5258 | Alveolar Volume
5259 | Apparatus Dead Space
5260 | Pleth (Pre-Ductal) | Pulse Rate from Plethysmogram (pre ductal)
5261 | Pleth (Post-Ductal) | Pulse Rate from Plethysmogram (post ductal)
5262 | Relative Perfusion (Pre-Ductal)
5263 | Relative Perfusion (Post-Ductal)
5264 | Cardioplegia Main Pump Time of Current Plegia Session
5265 | Cardioplegia Main Pump Time Since Stop of Plegia Session
5266 | Cardioplegia Main Pump Total Plegia Time
5267 | Cardioplegia Slave Pump Time Since Start of Current Plegia Session
5268 | Cardioplegia Slave Pump Time Since Stop of Plegia Session
5269 | Cardioplegia Slave Pump Total Plegia Time
5270 | PAV-based Lung Compliance
5271 | PAV-based Patient Resistance
5272 | PAV-based Total Resistance Estimate
5273 | PAV-based Lung Elastance
5274 | Backup Inspiratory Pressure Setting
5275 | Percent of Leakage
5276 | Maximum Inspiratory Time Setting
5277 | Plateau Time
5278 | QT Baseline
5279 | QT Heart Rate Baseline
5280 | High Density Lipoprotein
5281 | Low Density Lipoprotein
5282 | C-reactive Protein
5283 | Urine Hemoglobin
5284 | Noninvasive Total Hemoglobin (SpHb)
5285 | Noninvasive Oxygen Content (SpOC)
5286 | Noninvasive Carboxyhemoglobin (SpCO)
5287 | Noninvasive Methemoglobin (SpMet)
5288 | Pleth Variability Index (PVI)
5289 | Regional O2 Saturation (rSO2)
5290 | rSO2 Area Under Curve (AUC) | This variable is associated with ChannelID identifier.
5291 | Generic O2 Saturation (3)
5292 | Generic O2 Saturation (4)
5293 | Tissue O2 Saturation (StO2)
5294 | Nitrogen Balance (NsLoss)
5295 | Inspired O2 Concentration | See variable 52527 for partial pressure
5296 | Inspired O2 Partial Pressure | This variable is calculated from variable 52523 and will output inspired O2 partial pressure.
5298 | Heart Rate
5299 | SpO2
5300 | Cardiac Output
5301 | Central Venous Pressure (CVP) | This variable transmits either a continuous measurement or a manual input value.
5302 | Patient Category | Consult detailed help for more information on this variable
5303 | Cardiac Index
5304 | Stroke Volume | Beat to beat stroke volume.
5305 | Systemic Vascular Resistance
5306 | Patient Height
5307 | Patient Weight
5308 | Generic Error Code | TD measurement error, "0" is transmitted if no error occurs.
5309 | Hematocrit
5310 | Patient Body Surface Area
5311 | Systemic Vascular Resistance Index
5312 | O2 Consumption (VO2)
5313 | O2 Consumption Index
5314 | Cardiac Power Index (CCO)
5315 | Cardiac Power Output (CCO)
5316 | ICG Measurement Counter | Counter for signalizing update of ICG measurement.
5317 | Plasma Disappearance Rate
5318 | ICG Retention Rate
5319 | Cardiac Index ((Thermodilution)
5320 | Cardiac Output (Thermodilution)
5321 | Global Ejection Fraction (Thermodilution)
5322 | Global End-Diastolic Volum (Thermodilution)
5323 | Global End-Diastolic Volum Index (Thermodilution)
5324 | Pulmonary Vascular Permeability Index (Thermodilution)
5325 | Generic ICG Error | ICG measurement error, "0" is transmitted if no error occurs.
5326 | Patient ID | 9 digits maximum, only contain numbers from 0 to 9.
5327 | Patient Name | Allowed characters include A..Z, 0..9, ' ', '.', and '-'. The device sends a single character each 400 ms, so it takes time to transmit the full patient name.
5328 | Predicted Patient Body Surface Area
5329 | Predicted Patient Weight
5330 | Device Date/Time | This variable contains the current time of the device in the form of "MM/dd/yyyy - hh:mm:ss". It is also sent in the "DeviceTime" attribute of other variables.
5331 | Arterial BP Systolic
5332 | Arterial BP Diastolic
5333 | Arterial BP Mean
5334 | Patient Blood Temperature | A offset of 30°C is already taken into account by the DDI and consequently, the user receives the blood temperature real value, without having to interpret it.
5335 | Injectate Temperature | A offset of 15°C is already taken into account by the DDI and consequently, the user receives the injectate temperature real value, without having to interpret it.
5336 | O2 Delivery (DO2)
5337 | SvO2 Alarm | High/Low Alarm of ScvO2 Consult detailed help for more information on this variable
5338 | SvO2 Calibration Status | ScvO2Cal: Point of time of calibration. "1" once during “Bloodsample drawn”, otherwise "0".
5339 | Device Software Revision | Release PiCCO software version, composed by one letter (type of protocol: V for usual protocol, D for demo version and X for unreleased version) and from one to three digits seperated by dots (version number). For example: D5.1
5340 | Ejection Fraction | Global Ejection Fraction
5341 | Alternate Patient ID | Alternate display of Patient ID. Allowed characters include A..Z, 0..9, ' ', '.', and '-'. The device sends a single character each 400 ms, so it takes time to transmit the full patient ID.
5342 | Arterial O2 Saturation (SaO2) | Arterial oxygenation saturation (input value for display of parameters DO2 and VO2).
5343 | Device Protocol Revision | Release PiCCO2 software version, composed by one letter (type of protocol: V for usual protocol, D for demo version and X for unreleased version) and from one to three digits seperated by dots (version number). For example: D5.1
5344 | Stroke Volume Variation
5345 | O2 Delivery Index (DO2I)
5346 | Cardiac Function Index (Thermodilution)
5347 | Intrathoracic Blood Volume (Thermodilution)
5348 | Intrathoracic Blood Volume Index (Thermodilution)
5349 | Extravascular Lung Water (Thermodilution)
5350 | Extravascular Lung Water Index (Thermodilution)
5351 | Intrathoracic Blood Volume (ITBV)
5352 | Intrathoracic Blood Volume Index (ITBI)
5353 | Extravascular Lung Water
5354 | Extravascular Lung Water Index
5355 | Global End Diastolic Volume
5356 | Global End Diastolic Volume Index
5357 | Cardiac Function Index
5358 | SvO2 Calibration Data | ScvO2Cal: ScvO2 Calibration input.
5359 | Continuous Cardiac Output (CCO) Alarm | High/Low Alarm of PCCO Consult detailed help for more information on this variable
5360 | Pulse Contour Cardiac Output
5361 | Thermodilution Measurement Status | Consult detailed help for more information on this variable
5362 | Index Of Left Ventricular Contractility (dPmx)
5363 | Mean Transit Time (MTt)
5364 | Down Slope Time
5365 | Temperature Difference (Thermodilution)
5366 | Injectate Temperature (Thermodilution)
5367 | PICCO Catheter | Range 0-14 Consult detailed help for more information on this variable
5368 | Pulse Pressure Variation
5369 | Pulmonary Vascular Permeability Index
5370 | Patient Gender | Consult detailed help for more information on this variable
5371 | Pulse Contour Cardiac Index
5372 | Thermodilution Measurement Counter | Counter for signalizing when all discontinuous thermodilution parameters have been updated.
5373 | Right to Left Heart Shunt Fraction
5374 | Arterial Blood Pressure Wave
5375 | Injectate Volume Setting
5376 | Stroke Volume Min
5377 | Stroke Volume Max
5378 | Stroke Volume Index
5379 | Stroke Volume Mean
5380 | Central Venous O2 Saturation (ScvO2)
5381 | SpO2 Alarm | High/Low Alarm of SpO2 Consult detailed help for more information on this variable
5383 | Maximum Tidal Volume Limit | Possible units: mL, L
5384 | Battery Running Alarm | Consult detailed help for more information on this variable
5385 | Volume Not Delivered Alarm | Consult detailed help for more information on this variable
5386 | Patient Rebreathing Alarm | Consult detailed help for more information on this variable
5387 | General Alarm Status | Contains every alarm code that is currently active
5389 |  | Consult detailed help for more information on this variable
5390 | ASB Ramp Time Setting | Possible units: msec, sec, min
5391 | Maximum Lung Elastance Setting | Possible units: cmH2O/L, mBar/L
5392 | Maximum Airway Resistance Setting | Possible units: cmH2O/L/s, mBar/L/s
5393 | Low Airway Flow Alarm | Consult detailed help for more information on this variable
5394 | Pressure Alarm Delay
5395 | Patient Alarm Global State | Alarm disabled status
5397 |  | Consult detailed help for more information on this variable
5398 | Trigger Sensitivity Setting
5399 | Percent of Patient Triggered Breaths
5400 | Estimated Patient Circuit Leakage | Possible units: mL/min, L/min
5401 | Ventilation Comfort Mode Flex Setting
5402 | IPAP Minimum Pressure Setting | Possible units: cmH2O, hPa, kPa, mBar, mmHg
5403 | Patient Expiratory Port Type | Consult detailed help for more information on this variable
5404 | Vt Not Achieved Alarm | Consult detailed help for more information on this variable
5405 | Patient Interface Type | Consult detailed help for more information on this variable
5406 | Tidal Volume Exceeded Alarm | Consult detailed help for more information on this variable
5407 | Check Maximum Pressure Alarm | Consult detailed help for more information on this variable
5408 | Maximum IPAP Setting | Possible units: cmH2O, hPa, kPa, mBar, mmHg
5409 | Low Ventilator Total Respiration Rate Alarm | Consult detailed help for more information on this variable
5410 | Low Expired Tidal Volume Alarm | Consult detailed help for more information on this variable
5411 | Low Expired Minute Volume Alarm | Consult detailed help for more information on this variable
5412 | Low Peak Inspiratory Pressure Alarm | Consult detailed help for more information on this variable
5413 | Device Model
5414 | Device Software Revision
5415 | Ventilator Total Respiration Rate
5416 | Inspiration Time Percentage
5417 | Expired Tidal Volume
5418 | Expired Minute Volume
5419 | Peak Inspiratory Pressure
5420 | Airway Flow Waveform
5421 | Airway Volume Waveform
5422 | Airway Pressure Waveform
5423 | Ventilation Standby State
5424 | Ventilation Mode | Consult detailed help for more information on this variable
5425 | Ventilation Mode Primary | Consult detailed help for more information on this variable
5426 | Ventilation Mode Options | Consult detailed help for more information on this variable
5427 | Inspired O2 (FiO2) Setting
5428 | Ventilation Leakage Minute Volume
5429 | Mandatory Respiratory Rate Setting
5430 | Inspiratory Time Setting
5431 | lnspiratory Rise Time Setting
5432 | Inspired Tidal Volume Setting
5433 | CPAP Setting
5434 | Inspiratory Flow Termination Setting | Auto-Trak sensitivity: set flow for cycling to exhalation
5435 | Proportional Support Level Setting
5436 | EPAP Setting
5437 | IPAP Setting
5438 | Fresh Gas Total Flow Rate Setting
5439 | High Ventilator Total Respiration Rate Alarm | Consult detailed help for more information on this variable
5440 | High Expired Tidal Volume Alarm | Consult detailed help for more information on this variable
5441 | High Peak Inspiratory Pressure Alarm | Consult detailed help for more information on this variable
5442 | Battery Low Alarm | This variable is associated with a ChannelID identifier.
5444 |  | Consult detailed help for more information on this variable
5445 | Main Power Disconnected Alarm | Consult detailed help for more information on this variable
5446 | Ventilator Patient Disconnected Alarm | Consult detailed help for more information on this variable
5447 | Ventilator Inoperative Alarm | Consult detailed help for more information on this variable
5448 | Ventilator Obstruction Alarm | Consult detailed help for more information on this variable
5449 | Pressure Sensor Disconnected Alarm | Consult detailed help for more information on this variable
5450 | O2 Supply Alarm | Consult detailed help for more information on this variable
5451 | O2 Supply High Alarm | Consult detailed help for more information on this variable
5452 | O2 Supply Low Alarm | Consult detailed help for more information on this variable
5454 | End Tidal CO2 (EtCO2) | This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. Possible units: cmH2O, mbar, bar, hPa, kPa, mmHg, %
5455 | Respiration Rate | This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. Possible units: 1/min
5456 | Minimum Respiration Rate Threshold Setting | Possible units: 1/min
5457 | Ventilation Mode | Consult detailed help for more information on this variable
5458 | Inspired O2 (FiO2) | This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. Possible units: %
5459 | Device Name | ASCII string up to 24 characters long.
5460 | Inspiration Time Percent Setting | Possible units: %
5461 | Battery Running Alarm
5462 | Low Inspired Tidal Volume Alarm
5463 | Low Expired Tidal Volume Alarm
5464 | Positive End Expiratory Pressure (PEEP) | This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. Possible units: cmH2O, mbar, bar, hPa, kPa, mmHg
5465 | High Inspired Tidal Volume Alarm
5466 | Tidal Volume Setting | Possible units: mL, L
5467 | Sigh Interval Setting | Possible units: ms, sec, min, h
5468 | Sigh Pressure Limit Setting | Possible units: cmH2O, mbar, bar, hPa, kPa, mmHg
5469 | Apnea Tidal Volume Setting | Possible units: mL, L
5470 | Apnea Respiration Rate Setting | Possible units: 1/min
5471 | Battery Low Alarm
5472 | Inspiratory Flow Trigger Setting | Possible units: mL/min, mL/h, L/s, L/min, L/h
5473 | Apnea Time Setting | Possible units: ms, sec, min, h
5474 | Mean Airway Pressure | This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. Possible units: cmH2O, mbar, bar, hPa, kPa, mmHg
5475 | O2 Sensor Alarm
5476 | Expiratory Trigger Level Setting | Possible units: %
5477 | Device Power Supply Status | Value is Sector, Internal Battery, or External Battery
5478 | Device Serial Number | ASCII string up to 24 characters long.
5479 | SIMV Respiration Rate Setting | Possible units: 1/min
5480 | Inspired Minute Volume | This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. Possible units: mL, L
5481 | Inspiratory Time Setting | Possible units: ms, sec, min, h
5482 | Positive End Expiratory Pressure (PEEP) Setting | Possible units: cmH2O, mbar, bar, hPa, kPa, mmHg
5483 | Inspired O2 (FiO2) Setting | Possible units: %
5484 | Battery Alarm | This alarm is raised if the internal battery is discharged.
5485 | Inspired Tidal Volume | This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. Possible units: mL, L
5486 | Expired Tidal Volume | This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. Possible units: mL, L
5487 | Expired Minute Volume | This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. Possible units: mL, L
5488 | Patient Rebreathing Alarm
5489 | Low Expired Minute Volume Alarm
5490 | Airway Peak Pressure Alarm
5491 | Ventilator Control Mode | The variable indicates if the ventilator is in Non-Invasive Ventilation. ON means non-invasive ventilation and OFF means invasive ventilation.
5492 | Target Respiration Rate | Possible units: 1/min
5493 | I:E Ratio Setting
5494 | Generic Device Alarm | This is a multiset variable containing the text messages for all active alarms. Associated with variable 6164 for the matching alarm codes.
5495 | Support Pressure Setting | Possible units: cmH2O, mbar, bar, hPa, kPa, mmHg
5496 | Device Software Revision | ASCII string up to 24 characters long.
5497 | Tidal Volume Target | Possible units: mL, L
5498 | I:E Ratio
5499 | Plateau Airway Pressure | This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. Possible units: cmH2O, mbar, bar, hPa, kPa, mmHg
5500 | Patient Disconnect Alarm
5501 | Patient Type | Value is ADULT, CHILD, INFANT, or BABY.
5502 | Volume Not Achieved Alarm
5503 | Plateau Time Setting | Possible units: ms, sec, min, h
5504 | Breathing System Leakage | This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. Possible units: %
5505 | Respiration Rate Setting | Possible units: 1/min
5506 | Sigh Tidal Volume Setting | Possible units: mL, L
5507 | Apnea Ventilation Alarm
5508 | Maximum Inspired Pressure Setting | Possible units: cmH2O, mbar, bar, hPa, kPa, mmHg
5509 | Inspiratory Pressure Setting | Possible units: cmH2O, mbar, bar, hPa, kPa, mmHg
5510 | Airway Peak Flow Setting | Possible units: mL/min, mL/h, L/s, L/min, L/h
5511 | Low EtCO2 Alarm
5512 | High EtCO2 Alarm
5513 | Waveform Shape Setting | Consult detailed help for more information on this variable
5514 | PEEP/CPAP Alarm
5515 | ET Tube Blocked Alarm
5516 | General Alarm Code | This is a multiset variable containing the alarm codes for all active alarms. Associated with variable 2590 for the matching alarm text messages.
5518 |  | Consult detailed help for more information on this variable
5519 | O2 Supply Alarm
5520 | High Leak Alarm
5521 | Ratio of Inspiration Time to Total Breath Cycle | This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. Possible units: %
5522 | Expiratory Flow Sensor Failure Alarm
5523 | Patient Demand High Alarm
5524 | High Inspired Minute Volume Alarm
5525 | Low Inspired Minute Volume Alarm
5526 | Plateau Airway Pressure Alarm
5527 | Nebulizer Error
5528 | Peak Pressure | This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. Possible units: cmH2O, mbar, bar, hPa, kPa, mmHg
5529 | Plimit Reached Alarm
5530 | High Expired Tidal Volume Alarm
5531 | Low FiO2 Alarm
5532 | High FiO2 Alarm
5533 | Low Respiration Rate Alarm
5534 | High Respiration Rate Alarm
5535 | CO2 Apnea Alarm
5536 | Battery Failure Alarm
5537 | Ventilator Failure Alarm
5538 | Pressure Rise Time Setting
5539 | Maximum Inspiratory Time Setting | Possible units: ms, sec, min, h
5540 | High Expired Minute Volume Alarm
5541 | CO2 Sensor Failure Alarm
5543 | Air Bubble Detection State | Consult detailed help for more information on this variable
5544 | Module State | This variable is 16 bits long.
5546 |  | Consult detailed help for more information on this variable
5547 | Device Power Supply Status | This variable is 24 bits long.
5549 |  | Consult detailed help for more information on this variable
5550 | AC Power Supply Failure Alarm | This variable is 32 bits long.
5552 |  | Consult detailed help for more information on this variable
5553 | Generic Device Alarm | This variable is 40 bits long.
5555 |  | Consult detailed help for more information on this variable
5556 | Flow Sensor Type | Consult detailed help for more information on this variable
5557 | Patient Temperature | This variable is associated with LowerAlarmLimit and UpperAlarmLimit attributes. This variable is associated with ChannelID identifier.
5558 | Device Screen State | Possible values :
5559 |  | 0 : GUI Unlocked
5560 |  | 1 : GUI Locked
5561 | Pulsatile Pulse Percentage Setting | This variable is for pulsatile mode.
5562 | Generic Pump Rotation Speed | This variable is associated with LowerAlarmLimit and UpperAlarmLimit attributes.
5563 | Pulsatile Simulated Frequency Setting | This variable is for pulsatile mode.
5564 | Device Alarm Status | This variable is 8 bits long.
5566 |  | Consult detailed help for more information on this variable
5567 | HLM System Control Module Technical Alarm | This variable is 24 bits long.
5569 |  | Consult detailed help for more information on this variable
5570 | Generic Pump Running Status | This variable is 32 bits long.
5572 |  | Consult detailed help for more information on this variable
5573 | Generic Pump Flow Rate Setting
5574 | Generic Pump Rotation Speed Setting
5575 | Generic Pump Flow Alarm | This variable is 16 bits long.
5577 |  | Consult detailed help for more information on this variable
5578 | Blood Level Measurement Status | Consult detailed help for more information on this variable
5579 | Blood Flow Rate | This variable is associated with LowerAlarmLimit and UpperAlarmLimit attributes.
5580 | Generic Pressure Line Value | This variable is associated with LowerAlarmLimit and UpperAlarmLimit attributes. This variable is associated with ChannelID identifier.
5581 | Pulsatile Pump Speed Amplitude Setting | This variable is for pulsatile mode.
5582 | Generic ECG Alarm | This variable is 8 bits long.
5584 |  | Consult detailed help for more information on this variable