## Sheet1
ID | Label | Len_Occurrences | Files_Occurences | Unit | OBX_Units | Type | Group | Description | Device | Source
113 | Pleth Waveform | 2 | ['df_MX800_HL7_waveform', 'df_Test1_hl7'] | - | 139 | Double | Waveforms |  | IntelliVue X3 | Waves: Pleth
5141 | ECG Lead II Wave | 2 | ['df_MX800_HL7_waveform', 'df_Test1_hl7'] | - | 47 | Double | Waveforms |  | IntelliVue X3 | Waves: II
5142 | ECG Lead III Wave | 2 | ['df_MX800_HL7_waveform', 'df_Test1_hl7'] | - | 47 | Double | Waveforms |  | IntelliVue X3 | Waves: III
8857 | Arterial Blood Pressure (ABP) Wave | 2 | ['df_MX800_HL7_waveform', 'df_Test1_hl7'] | - | 9 | Double | Waveforms |  | IntelliVue X3 | Waves: ABP
8870 | Impedance Respiration Wave | 2 | ['df_MX800_HL7_waveform', 'df_Test1_hl7'] | - | 84 | Double | Waveforms |  | IntelliVue X3 | Waves: Resp
1364 |  | 3 | ['df_MX800_HL7', 'df_V60_OPptiflow_HL7', 'df_V60_ST_HL7'] |  | 139
2975 |  | 3 | ['df_MX800_HL7', 'df_V60_OPptiflow_HL7', 'df_V60_ST_HL7'] |  | 139
2978 |  | 3 | ['df_MX800_HL7', 'df_V60_OPptiflow_HL7', 'df_V60_ST_HL7'] |  | 139
4377 |  | 3 | ['df_MX800_HL7', 'df_V60_OPptiflow_HL7', 'df_V60_ST_HL7'] |  | 139
4378 |  | 3 | ['df_MX800_HL7', 'df_V60_OPptiflow_HL7', 'df_V60_ST_HL7'] |  | 139
5973 |  | 3 | ['df_MX800_HL7', 'df_V60_OPptiflow_HL7', 'df_V60_ST_HL7'] |  | 139
5975 |  | 3 | ['df_MX800_HL7', 'df_V60_OPptiflow_HL7', 'df_V60_ST_HL7'] |  | 139
6289 |  | 3 | ['df_MX800_HL7', 'df_V60_OPptiflow_HL7', 'df_V60_ST_HL7'] |  | 139
6290 |  | 3 | ['df_MX800_HL7', 'df_V60_OPptiflow_HL7', 'df_V60_ST_HL7'] |  | 139
6291 |  | 3 | ['df_MX800_HL7', 'df_V60_OPptiflow_HL7', 'df_V60_ST_HL7'] |  | 139
6292 |  | 3 | ['df_MX800_HL7', 'df_V60_OPptiflow_HL7', 'df_V60_ST_HL7'] |  | 139
6293 |  | 3 | ['df_MX800_HL7', 'df_V60_OPptiflow_HL7', 'df_V60_ST_HL7'] |  | 139
6294 |  | 3 | ['df_MX800_HL7', 'df_V60_OPptiflow_HL7', 'df_V60_ST_HL7'] |  | 139
6296 | O2 Supply Alarm | 3 | ['df_MX800_HL7', 'df_V60_OPptiflow_HL7', 'df_V60_ST_HL7'] | unitless | 139 | Boolean | Alarms |  | Monnal T60
6299 |  | 3 | ['df_MX800_HL7', 'df_V60_OPptiflow_HL7', 'df_V60_ST_HL7'] |  | 139
6300 |  | 3 | ['df_MX800_HL7', 'df_V60_OPptiflow_HL7', 'df_V60_ST_HL7'] |  | 139
6303 |  | 3 | ['df_MX800_HL7', 'df_V60_OPptiflow_HL7', 'df_V60_ST_HL7'] |  | 139
6305 |  | 3 | ['df_MX800_HL7', 'df_V60_OPptiflow_HL7', 'df_V60_ST_HL7'] |  | 139
6376 | High Leak Alarm | 3 | ['df_MX800_HL7', 'df_V60_OPptiflow_HL7', 'df_V60_ST_HL7'] | unitless | 139 | Boolean | Alarms |  | Monnal T60
92001 | Ventilator Patient Disconnected Alarm | 3 | ['df_MX800_HL7', 'df_V60_OPptiflow_HL7', 'df_V60_ST_HL7'] | unitless | 139 | Long | Alarms | Consult detailed help for more information on this variable | Common
92011 | Ventilator Inoperative Alarm | 3 | ['df_MX800_HL7', 'df_V60_OPptiflow_HL7', 'df_V60_ST_HL7'] | unitless | 139 | Long | Alarms | Consult detailed help for more information on this variable | Common
26 | Device Inactivation Status | 3 | ['df_MX800_HL7', 'df_V60_OPptiflow_HL7', 'df_V60_ST_HL7'] | unitless | 139 | String | DeviceInfo | Possible values: "enabled" and "audio-paused". See "Multiset Alarms" on help file. | Common | Alarms
171 |  | 3 | ['df_MX800_HL7', 'df_V60_OPptiflow_HL7', 'df_V60_ST_HL7'] |  | 139
172 |  | 3 | ['df_MX800_HL7', 'df_V60_OPptiflow_HL7', 'df_V60_ST_HL7'] |  | 139
4390 |  | 3 | ['df_MX800_HL7', 'df_V60_OPptiflow_HL7', 'df_V60_ST_HL7'] |  | 139
6216 |  | 3 | ['df_MX800_HL7', 'df_V60_OPptiflow_HL7', 'df_V60_ST_HL7'] |  | 139
6307 |  | 3 | ['df_MX800_HL7', 'df_V60_OPptiflow_HL7', 'df_V60_ST_HL7'] |  | 139
6266 |  | 3 | ['df_MX800_HL7', 'df_V60_OPptiflow_HL7', 'df_V60_ST_HL7'] |  | 139
6277 |  | 3 | ['df_MX800_HL7', 'df_V60_OPptiflow_HL7', 'df_V60_ST_HL7'] |  | 19
3326 | Device Average Gas Flow | 1 | ['df_V500_HL7'] | L/min | 32 | Double | OtherMeasuredData | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. Average device flow, FlowDev. | Evita Infinity V500 [MEDIBUS.X]
9476 | Percent of Leakage | 1 | ['df_V500_HL7'] | % | 19 | Double | OtherMeasuredData | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Evita V300. Leakage minute volume in % of inspiratory minute volume, % leak Leakage rel. % leak - Infinity V500 determines the difference between the delivered flow on the inspiratory side and the measured flow on the expiratory side. This difference provides a measure of the amount of leakage and is displayed by Infinity V500 as the leakage minute volume MVleak and relative as % leak (MVleak to MV). For Evita Infinity V500, Babylog VN500: Leakage rel. % leak. | Evita Infinity V500 [MEDIBUS.X]
60 | Tidal Volume | 1 | ['df_V500_HL7'] | mL | 26 | Double | VitalSigns | For Evita 4, Evita 2 dura, Evita XL: Expiratory tidal volume. This variable is associated with UpperAlarmLimit attribute, which is also sent in separate variable 2763. Values below 1 L will be transmitted in units of mL, otherwise in units of L. Device firmware version 04.00 and higher: Limit is not transmitted, if alarm is switched off. For Savina: This variable is associated with UpperAlarmLimit attribute.Device SW version 02.00 and higher: Limit is not transmitted, if alarm is switched off. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300, Fabius Family. This variable is associated with UpperAlarmLimit attribute, which is available for Perseus A500, Evita Infinity V500, Savina 300, Evita V300. This variable is associated with LowerAlarmLimit attribute, which is available for Evita Infinity V500, Evita V300. For Evita Infinity V500, Babylog VN500: This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. For Zeus, Zeus IE, Primus, Primus IE, Apollo, Pallas: Exp. Tidal Volume. | Evita Infinity V500 [MEDIBUS.X]
61 | Dynamic Compliance | 1 | ['df_V500_HL7'] | L/bar | 105 | Double | VitalSigns | For Evita 4, Evita 2 dura, Evita XL: Values below 100 will be transmitted with one decimal place, otherwise without a decimal place. For Savina: Available since device SW version 02.00. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Evita V300. | Evita Infinity V500 [MEDIBUS.X]
582 | Resistance | 1 | ['df_V500_HL7'] | mbar/L/s | 106 | Double | VitalSigns | For Evita 4, Evita 2 dura, Evita XL: Values below 100 will be transmitted with one decimal place, otherwise without a decimal place. For Savina: Available since device SW version 02.00. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Zeus IE, Evita V300. | Evita Infinity V500 [MEDIBUS.X]
635 | Inspired O2 (FiO2) | 1 | ['df_V500_HL7'] | % | 19 | Double | VitalSigns | For Savina: This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. Available since device SW version 03.00. Limits are transmitted in LPO-Mode (Option) only. For Fabius GS, Fabius GS premium, Fabius Tiro, Fabius plus, Fabius plus XL, Fabius OS, Fabius MRI: This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Evita V300, Fabius Family. This variable is associated with UpperAlarmLimit attribute, which is available for Perseus A500, Atlan, Primus, Primus IE, Apollo, Zeus IE, Fabius Family. This variable is associated with LowerAlarmLimit attribute, which is available for Perseus A500, Atlan, Primus, Primus IE, Apollo, Zeus IE, Fabius Family. Inspiratory oxygen fraction, FiO2. For Zeus, Zeus IE: This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. For Primus, Primus IE, Apollo, Pallas: This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. | Evita Infinity V500 [MEDIBUS.X]
775 | Inspiratory Component of I:E Ratio | 1 | ['df_V500_HL7'] | unitless | 139 | Double | VitalSigns | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: Obsolete variable that may be removed in future release. Partial information of I:E ratio. For Evita Infinity V500, Savina: Partial information. Please use variable 1190 instead. Obsolete variable that may be removed in future release. | Evita Infinity V500 [MEDIBUS.X]
776 | Expiratory Component of I:E Ratio | 1 | ['df_V500_HL7'] | unitless | 139 | Double | VitalSigns | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: Obsolete variable that may be removed in future release. Partial information of I:E ratio. For Evita Infinity V500, Savina: Partial information. Please use variable 1190 instead. Obsolete variable that may be removed in future release. | Evita Infinity V500 [MEDIBUS.X]
1189 | Positive End Expiratory Pressure (PEEP) | 1 | ['df_V500_HL7'] | mbar | 44 | Double | VitalSigns | For Fabius GS, Fabius GS premium, Fabius Tiro, Fabius plus, Fabius plus XL, Fabius OS, Fabius MRI: 1mbar = 1cmH2O. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300, Fabius Family. | Evita Infinity V500 [MEDIBUS.X]
1190 | I:E Ratio | 1 | ['df_V500_HL7'] | unitless | 139 | String | VitalSigns | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Evita V300. Format: 1:x if E > I, x:1 if I > E, or 1:1 if I = E. The useless leading or ending “0” or “.” is not transmitted, and the maximum digits after the floating point is 2. For Evita Infinity V500, Babylog VN500: Format: 1:x if E > I, x:1 if I > E, or 1:1 if I = E. The useless leading or ending “0” or “.” is not transmitted, and the maximum digits after the floating point is 2. | Evita Infinity V500 [MEDIBUS.X]
1306 | Tidal Volume | 2 | ['df_V500_HL7', 'df_V60_ST_HL7'] | mL | 26 | Double | VitalSigns | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: Obsolete variable that may be removed in future release. This variable is marked as deprecated in order to synchronize with other models. Please use variable 60 instead. For Evita Infinity V500, Babylog VN500: Deprecated in order to synchronize with other models. Please use variable 60 instead. | Evita Infinity V500 [MEDIBUS.X]
1307 | Minute Volume | 2 | ['df_V500_HL7', 'df_V60_ST_HL7'] | L/min | 32 | Double | VitalSigns | For Evita 4, Evita 2 dura, Evita XL: This variable is associated with UpperAlarmLimit attribute, which is also sent in separate variable 1400. This variable is associated with LowerAlarmLimit attribute, which is also sent in separate variable 1399. For device firmware version 04.00 and higher: Lower alarm limit is not transmitted, if alarm is switched off. Values may be transmitted with one or two decimal place, depended on firmware version. Note that the value sent by the device is less precise than what is displayed on device screen. For example, a displayed value of 5.12 is sent by the device and thus by DDI as 5.1. For Evita 2, Evita, Fabius GS, Fabius GS premium, Fabius Tiro, Fabius plus, Fabius plus XL, Fabius OS, Fabius MRI, Evita Infinity V500, Babylog VN500, Carina: This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. This variable is available only with a precision to the tenth. For example: for a minute volume value of 7.43, the device displays 7.43 L/min but the DDI will output it as 7.4 L/min. For Savina: This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300, Fabius Family. This variable is associated with UpperAlarmLimit attribute, which is available for Perseus A500, Atlan, Savina 300, Primus, Primus IE, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Fabius Family. This variable is associated with LowerAlarmLimit attribute, which is available for Perseus A500, Atlan, Savina 300, Primus, Primus IE, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Fabius Family. This variable is available only with a precision to the tenth. For example: for a minute volume value of 7.43, the device displays 7.43 L/min but the DDI will output it as 7.4 L/min. For Zeus, Zeus IE, Primus, Primus IE, Apollo, Pallas: This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. Respiratory Minute Volume This variable is available only with a precision to the tenth. For example: for a minute volume value of 7.43, the device displays 7.43 L/min but the DDI will output it as 7.4 L/min. For Evita Infinity V500 and Babylog VN500: This variable is only output if the device version is less than 2.30, otherwise variable 2326 is exported. | Evita Infinity V500 [MEDIBUS.X]
1314 | Spontaneous Respiration Rate | 1 | ['df_V500_HL7'] | 1/min | 1 | Double | VitalSigns | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Zeus IE, Oxylog 3000 plus, Evita V300. This variable is associated with UpperAlarmLimit attribute, which is available for Oxylog 3000 plus. | Evita Infinity V500 [MEDIBUS.X]
1415 | Mean Airway Pressure | 1 | ['df_V500_HL7'] | mbar | 44 | Double | VitalSigns | For Fabius GS, Fabius GS premium, Fabius Tiro, Fabius plus, Fabius plus XL, Fabius OS, Fabius MRI: This variable is associated with LowerAlarmLimit attribute. 1mbar = 1cmH2O. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300, Fabius Family. For Zeus, Zeus IE, Primus, Primus IE, Apollo, Pallas: Mean Breathing Pressure. | Evita Infinity V500 [MEDIBUS.X]
1426 | Intrinsic PEEP | 1 | ['df_V500_HL7'] | mbar | 44 | Double | VitalSigns | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. | Evita Infinity V500 [MEDIBUS.X]
1569 | Minimum Airway Pressure | 1 | ['df_V500_HL7'] | mbar | 44 | Double | VitalSigns | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. | Evita Infinity V500 [MEDIBUS.X]
1570 | P0.1 Occlusion Pressure | 1 | ['df_V500_HL7'] | mbar | 44 | Double | VitalSigns | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. Occlusion pressure in P0.1 maneuver. For Evita Infinity V500, Babylog VN500: Only available for Evita Infinity V500 | Evita Infinity V500 [MEDIBUS.X]
1571 | Lung Volume Trapped (TrpVol) | 1 | ['df_V500_HL7'] | mL | 26 | Double | VitalSigns | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. Volume trapped in the lung by intrinsic PEEP, Vtrap. For Evita Infinity V500: Trapped Volume Vtrap. | Evita Infinity V500 [MEDIBUS.X]
1761 | Airway Resistance | 1 | ['df_V500_HL7'] | mbar/L/s | 106 | Double | VitalSigns | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. | Evita Infinity V500 [MEDIBUS.X]
2092 | Inspired Minute Volume | 1 | ['df_V500_HL7'] | L/min | 32 | Double | VitalSigns | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Carina, Evita V300. This variable is associated with UpperAlarmLimit attribute, which is available for Carina. This variable is associated with LowerAlarmLimit attribute, which is available for Carina. | Evita Infinity V500 [MEDIBUS.X]
2326 | Expired Minute Volume | 1 | ['df_V500_HL7'] | L/min | 32 | Double | VitalSigns | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. This variable is associated with UpperAlarmLimit attribute, which is available for Evita Infinity V500, Babylog VN500, Evita V300. This variable is associated with LowerAlarmLimit attribute, which is available for Evita Infinity V500, Babylog VN500, Evita V300. For Evita Infinity V500, Babylog VN500: This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. | Evita Infinity V500 [MEDIBUS.X]
2416 | Spontaneous Inspired Tidal Volume | 1 | ['df_V500_HL7'] | mL | 26 | Double | VitalSigns | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. For Zeus, Zeus IE: This variable is associated with UpperAlarmLimit attribute. Insp. Tidal volume. | Evita Infinity V500 [MEDIBUS.X]
2432 | Expired Spontaneous Tidal Volume | 1 | ['df_V500_HL7'] | mL | 26 | Double | VitalSigns | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. | Evita Infinity V500 [MEDIBUS.X]
2433 | Mechanical Expired Minute Volume | 1 | ['df_V500_HL7'] | L/min | 32 | Double | VitalSigns | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Primus, Primus Infinity Empowered, Apollo, Evita V300. Mandatory expiratory minute volume, MVemand | Evita Infinity V500 [MEDIBUS.X]
2434 | Spontaneous Expired Minute Volume | 1 | ['df_V500_HL7'] | L/min | 32 | Double | VitalSigns | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Primus, Primus Infinity Empowered, Apollo, Evita V300. | Evita Infinity V500 [MEDIBUS.X]
2564 | Peak Airway Pressure | 1 | ['df_V500_HL7'] | mbar | 44 | Double | VitalSigns | For Evita 4, Evita 2 dura, Evita XL: This variable is associated with UpperAlarmLimit attribute, which is also sent in separate variable 7040. For Evita 2, Evita, Savina: This variable is associated with UpperAlarmLimit attribute. For Fabius GS, Fabius GS premium, Fabius Tiro, Fabius plus, Fabius plus XL, Fabius OS, Fabius MRI: 1mbar = 1cmH2O. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300, Fabius Family. This variable is associated with UpperAlarmLimit attribute, which is available for Perseus A500, Atlan, Evita Infinity V500, Babylog VN500, Savina 300, Zeus IE, Oxylog 3000 plus, Carina, Evita V300, Fabius Family. This variable is associated with LowerAlarmLimit attribute, which is available for Zeus IE, Fabius Family. Peak inspiratory pressure, PIP. For Evita Infinity V500, Babylog VN500, Carina: This variable is associated with UpperAlarmLimit attribute. For Zeus, Zeus IE: This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. Peak Breathing Press. For Primus, Primus IE, Apollo, Pallas: Peak Breathing Pressure. | Evita Infinity V500 [MEDIBUS.X]
2570 | Spontaneous Tidal Volume | 1 | ['df_V500_HL7'] | - | 31 | Double (Multiset) | VitalSigns | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Zeus IE, Evita V300. Possible units: L, mL. For Evita Infinity V500, Babylog VN500: Possible units: L, mL. | Evita Infinity V500 [MEDIBUS.X]
3966 | Inspired Mandatory Tidal Volume | 1 | ['df_V500_HL7'] | mL | 26 | Double | VitalSigns | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. | Evita Infinity V500 [MEDIBUS.X]
3968 | Spontaneous Fraction of Minute Volume | 1 | ['df_V500_HL7'] | % | 19 | Double | VitalSigns | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. Spontaneous breathing portion of minute volume, % Mvspon. For Evita Infinity V500, Babylog VN500: Spontaneous Fraction Min. Vol. %MVspon (MVspon : MVtotal). | Evita Infinity V500 [MEDIBUS.X]
3971 | Negative Inspiratory Force | 1 | ['df_V500_HL7'] | mbar | 44 | Double | VitalSigns | For Evita 4, Evita 2 dura, Evita XL: The value is not transmitted, if it is displayed on Evita 4 / Evita 2 dura as "[less than] –45 mbar". Firmware version 04.00 or higher For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. | Evita Infinity V500 [MEDIBUS.X]
4003 | Generic Respiration Rate | 1 | ['df_V500_HL7'] | 1/min | 1 | Double | VitalSigns | For Evita 4, Evita 2 dura, Evita XL: This variable is associated with UpperAlarmLimit attribute, which is also sent in separate variable 1339. Total Respiratory Rate (RRtotal). For Evita 2, Evita, Savina, Carina: This variable is associated with UpperAlarmLimit attribute. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300, Fabius Family. This variable is associated with UpperAlarmLimit attribute, which is available for Evita Infinity V500, Babylog VN500, Savina 300, Carina, Evita V300. Respiratory rate, RR. For Evita Infinity V500, Babylog VN500: This variable is associated with UpperAlarmLimit attribute. Respiratory Rate (RRtotal) RR. | Evita Infinity V500 [MEDIBUS.X]
4892 | Airway Correlation Coefficient | 1 | ['df_V500_HL7'] | unitless | 139 | Double | VitalSigns | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. Correlation factor, r2. For Evita Infinity V500, Babylog VN500: r². | Evita Infinity V500 [MEDIBUS.X]
5340 | Expired Mandatory Tidal Volume | 1 | ['df_V500_HL7'] | mL | 26 | Double | VitalSigns | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. | Evita Infinity V500 [MEDIBUS.X]
7640 | Mandatory Respiration Rate | 1 | ['df_V500_HL7'] | 1/min | 1 | Double | VitalSigns | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Zeus IE, Evita V300. | Evita Infinity V500 [MEDIBUS.X]
7642 | Mandatory Tidal Volume | 1 | ['df_V500_HL7'] | - | 26 | Double (Multiset) | VitalSigns | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Babylog VN500, Babylog VN600, Babylog VN800, Zeus IE, Evita V300(added with SW 2.30 for Babylog VN500, added with SW 2.30 for Evita Infinity V500). Possible units: L, mL. Unit "L" available for Evita Infinity V500, Babylog VN500, Evita V300 (added with SW 2.30 for Babylog VN500, added with SW 2.30 for Evita Infinity V500). Unit "mL" available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Zeus IE, Evita V300. | Evita Infinity V500 [MEDIBUS.X]
9466 | Mandatory Triggered Respiration Rate | 1 | ['df_V500_HL7'] | 1/min | 1 | Double | VitalSigns | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300(added with SW 2.42 for Evita Infinity V500, Evita V300 and Babylog VN500). Respiratory rate of triggered, mandatory breaths, RRtrig For Evita 4, Evita 2 dura, Evita XL: Available only with Evita XL firmware version 7.00 or higher. | Evita Infinity V500 [MEDIBUS.X]
1312 |  | 1 | ['df_V60_ST_HL7'] |  | 32
6288 | Percent of Patient Triggered Breaths | 1 | ['df_V60_ST_HL7'] | % | 19 | Double | OtherMeasuredData |  | Common
6268 |  | 1 | ['df_V60_ST_HL7'] |  | 1
6269 |  | 1 | ['df_V60_ST_HL7'] |  | 39
6270 |  | 1 | ['df_V60_ST_HL7'] |  | 39
6271 |  | 1 | ['df_V60_ST_HL7'] |  | 4
6272 |  | 1 | ['df_V60_ST_HL7'] |  | 4
6273 |  | 1 | ['df_V60_ST_HL7'] |  | 139
6279 |  | 1 | ['df_V60_ST_HL7'] |  | 39
6281 | Tidal Volume Lower Alarm Limit | 1 | ['df_V60_ST_HL7'] | - | 26 | Double | Settings |  | IntelliVue X3 | AlarmLimit: lowTV
6282 |  | 1 | ['df_V60_ST_HL7'] |  | 1
6283 | Minute Volume Lower Alarm Limit | 1 | ['df_V60_ST_HL7'] | - | 31 | Double | Settings |  | IntelliVue X3 | AlarmLimit: lowMV
9223 |  | 1 | ['df_V60_ST_HL7'] |  | 139
2327 |  | 1 | ['df_V60_ST_HL7'] |  | 39
2761 |  | 1 | ['df_V60_ST_HL7'] |  | 19
5945 |  | 1 | ['df_V60_ST_HL7'] |  | 1
6379 | Ratio of Inspiration Time to Total Breath Cycle | 1 | ['df_V60_ST_HL7'] | - | 139 | Double | VitalSigns | This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. Possible units: % | Common