## Sheet1
ID | Label_icca | Driver_icca | Label | Unit | Description | Type | Group | Device | Source | Nom_Fichier
60 | VT exp mesuré | DragerMedibus | Tidal Volume | mL | For Evita 4, Evita 2 dura, Evita XL: Expiratory tidal volume. This variable is associated with UpperAlarmLimit attribute, which is also sent in separate variable 2763. Values below 1 L will be transmitted in units of mL, otherwise in units of L. Device firmware version 04.00 and higher: Limit is not transmitted, if alarm is switched off. For Savina: This variable is associated with UpperAlarmLimit attribute.Device SW version 02.00 and higher: Limit is not transmitted, if alarm is switched off. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300, Fabius Family. This variable is associated with UpperAlarmLimit attribute, which is available for Perseus A500, Evita Infinity V500, Savina 300, Evita V300. This variable is associated with LowerAlarmLimit attribute, which is available for Evita Infinity V500, Evita V300. For Evita Infinity V500, Babylog VN500: This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. For Zeus, Zeus IE, Primus, Primus IE, Apollo, Pallas: Exp. Tidal Volume. | Double | VitalSigns | ['Evita Infinity V500 [MEDIBUS.X]\xa0', 'Evita Infinity V800\xa0'] | [] | ['Evita V800', 'Evita V500']
61 | Compliance | DragerMedibus | Dynamic Compliance | L/bar | For Evita 4, Evita 2 dura, Evita XL: Values below 100 will be transmitted with one decimal place, otherwise without a decimal place. For Savina: Available since device SW version 02.00. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Evita V300. | Double | VitalSigns | ['Evita Infinity V500 [MEDIBUS.X]\xa0', 'Evita Infinity V800\xa0'] | [] | ['Evita V800', 'Evita V500']
582 | Resistance | DragerMedibus | Resistance | mbar/L/s | For Evita 4, Evita 2 dura, Evita XL: Values below 100 will be transmitted with one decimal place, otherwise without a decimal place. For Savina: Available since device SW version 02.00. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Zeus IE, Evita V300. | Double | VitalSigns | ['Evita Infinity V500 [MEDIBUS.X]\xa0', 'Evita Infinity V800\xa0'] | [] | ['Evita V800', 'Evita V500']
635 | FiO2 mesurée | DragerMedibus | Inspired O2 (FiO2) | ['%', '-'] | ['This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. Possible units: %\xa0', 'For Savina: This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. Available since device SW version 03.00. Limits are transmitted in LPO-Mode (Option) only. For Fabius GS, Fabius GS premium, Fabius Tiro, Fabius plus, Fabius plus XL, Fabius OS, Fabius MRI: This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Evita V300, Fabius Family. This variable is associated with UpperAlarmLimit attribute, which is available for Perseus A500, Atlan, Primus, Primus IE, Apollo, Zeus IE, Fabius Family. This variable is associated with LowerAlarmLimit attribute, which is available for Perseus A500, Atlan, Primus, Primus IE, Apollo, Zeus IE, Fabius Family. Inspiratory oxygen fraction, FiO2. For Zeus, Zeus IE: This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. For Primus, Primus IE, Apollo, Pallas: This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes.\xa0'] | Double | VitalSigns | ['Common\xa0', 'Evita Infinity V500 [MEDIBUS.X]\xa0', 'Evita Infinity V800\xa0'] | [] | ['Monnal T60', 'Evita V800', 'Evita V500']
775 | Inspiratory Component I:E ratio | DragerMedibus | Inspiratory Component of I:E Ratio | unitless | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: Obsolete variable that may be removed in future release. Partial information of I:E ratio. For Evita Infinity V500, Savina: Partial information. Please use variable 1190 instead. Obsolete variable that may be removed in future release. | Double | VitalSigns | ['Evita Infinity V500 [MEDIBUS.X]\xa0', 'Evita Infinity V800\xa0'] | [] | ['Evita V800', 'Evita V500']
776 | Expiratory Component I:E ratio | DragerMedibus | Expiratory Component of I:E Ratio | unitless | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: Obsolete variable that may be removed in future release. Partial information of I:E ratio. For Evita Infinity V500, Savina: Partial information. Please use variable 1190 instead. Obsolete variable that may be removed in future release. | Double | VitalSigns | ['Evita Infinity V500 [MEDIBUS.X]\xa0', 'Evita Infinity V800\xa0'] | [] | ['Evita V800', 'Evita V500']
1189 | PEEP mesure | DragerMedibus | Positive End Expiratory Pressure (PEEP) | ['mbar', '-'] | ['This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. Possible units: cmH2O, mbar, bar, hPa, kPa, mmHg\xa0', 'For Fabius GS, Fabius GS premium, Fabius Tiro, Fabius plus, Fabius plus XL, Fabius OS, Fabius MRI: 1mbar = 1cmH2O. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300, Fabius Family.\xa0'] | Double | VitalSigns | ['Common\xa0', 'Evita Infinity V500 [MEDIBUS.X]\xa0', 'Evita Infinity V800\xa0'] | [] | ['Monnal T60', 'Evita V800', 'Evita V500']
1190 | I/E Ratio | DragerMedibus | I:E Ratio | ['unitless', '-'] | ['For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Evita V300. Format: 1:x if E > I, x:1 if I > E, or 1:1 if I = E. The useless leading or ending “0” or “.” is not transmitted, and the maximum digits after the floating point is 2. For Evita Infinity V500, Babylog VN500: Format: 1:x if E > I, x:1 if I > E, or 1:1 if I = E. The useless leading or ending “0” or “.” is not transmitted, and the maximum digits after the floating point is 2.\xa0', 'This variable contains expiratory component of I:E Ratio assuming that the inspiratory component is "1".\xa0'] | ['Double\xa0', 'String\xa0'] | VitalSigns | ['IntelliVue MX800\xa0', 'Evita Infinity V500 [MEDIBUS.X]\xa0', 'IntelliVue X3\xa0', 'Evita Infinity V800\xa0', 'IntelliVue MP70\xa0'] | I:E 1: | ['X3', 'MX800', 'Evita V800', 'MP70', 'Evita V500']
1307 | VM inspire mesure | DragerMedibus | Minute Volume | L/min | For Evita 4, Evita 2 dura, Evita XL: This variable is associated with UpperAlarmLimit attribute, which is also sent in separate variable 1400. This variable is associated with LowerAlarmLimit attribute, which is also sent in separate variable 1399. For device firmware version 04.00 and higher: Lower alarm limit is not transmitted, if alarm is switched off. Values may be transmitted with one or two decimal place, depended on firmware version. Note that the value sent by the device is less precise than what is displayed on device screen. For example, a displayed value of 5.12 is sent by the device and thus by DDI as 5.1. For Evita 2, Evita, Fabius GS, Fabius GS premium, Fabius Tiro, Fabius plus, Fabius plus XL, Fabius OS, Fabius MRI, Evita Infinity V500, Babylog VN500, Carina: This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. This variable is available only with a precision to the tenth. For example: for a minute volume value of 7.43, the device displays 7.43 L/min but the DDI will output it as 7.4 L/min. For Savina: This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300, Fabius Family. This variable is associated with UpperAlarmLimit attribute, which is available for Perseus A500, Atlan, Savina 300, Primus, Primus IE, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Fabius Family. This variable is associated with LowerAlarmLimit attribute, which is available for Perseus A500, Atlan, Savina 300, Primus, Primus IE, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Fabius Family. This variable is available only with a precision to the tenth. For example: for a minute volume value of 7.43, the device displays 7.43 L/min but the DDI will output it as 7.4 L/min. For Zeus, Zeus IE, Primus, Primus IE, Apollo, Pallas: This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. Respiratory Minute Volume This variable is available only with a precision to the tenth. For example: for a minute volume value of 7.43, the device displays 7.43 L/min but the DDI will output it as 7.4 L/min. For Evita Infinity V500 and Babylog VN500: This variable is only output if the device version is less than 2.30, otherwise variable 2326 is exported. | Double | VitalSigns | ['Evita Infinity V500 [MEDIBUS.X]\xa0', 'Evita Infinity V800\xa0'] | [] | ['Evita V800', 'Evita V500']
1415 | Mean Airways Pressure | DragerMedibus | Mean Airway Pressure | ['mbar', '-'] | ['This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. Possible units: cmH2O, mbar, bar, hPa, kPa, mmHg\xa0', 'For Fabius GS, Fabius GS premium, Fabius Tiro, Fabius plus, Fabius plus XL, Fabius OS, Fabius MRI: This variable is associated with LowerAlarmLimit attribute. 1mbar = 1cmH2O. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300, Fabius Family. For Zeus, Zeus IE, Primus, Primus IE, Apollo, Pallas: Mean Breathing Pressure.\xa0'] | Double | VitalSigns | ['Monnal T60\xa0', 'Evita Infinity V500 [MEDIBUS.X]\xa0', 'Evita Infinity V800\xa0'] | [] | ['Monnal T60', 'Evita V800', 'Evita V500']
1426 | PEEP intrinseque | DragerMedibus | Intrinsic PEEP | mbar | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. | Double | VitalSigns | ['Evita Infinity V500 [MEDIBUS.X]\xa0', 'Evita Infinity V800\xa0'] | [] | ['Evita V800', 'Evita V500']
1569 | Minimum Airway Pressure | DragerMedibus | Minimum Airway Pressure | ['mbar', '-'] | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. | Double | VitalSigns | ['IntelliVue MX800\xa0', 'Evita Infinity V500 [MEDIBUS.X]\xa0', 'IntelliVue X3\xa0', 'Evita Infinity V800\xa0', 'IntelliVue MP70\xa0'] | AWPmin | ['X3', 'MX800', 'Evita V800', 'MP70', 'Evita V500']
1570 | Occlusion pressure | DragerMedibus | P0.1 Occlusion Pressure | ['mbar', '-'] | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. Occlusion pressure in P0.1 maneuver. For Evita Infinity V500, Babylog VN500: Only available for Evita Infinity V500 | Double | VitalSigns | ['IntelliVue MX800\xa0', 'Evita Infinity V500 [MEDIBUS.X]\xa0', 'IntelliVue X3\xa0', 'Evita Infinity V800\xa0', 'IntelliVue MP70\xa0'] | VENT: Poccl | ['X3', 'MX800', 'Evita V800', 'MP70', 'Evita V500']
1571 | Trapped volume | DragerMedibus | Lung Volume Trapped (TrpVol) | mL | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. Volume trapped in the lung by intrinsic PEEP, Vtrap. For Evita Infinity V500: Trapped Volume Vtrap. | Double | VitalSigns | ['Evita Infinity V500 [MEDIBUS.X]\xa0', 'Evita Infinity V800\xa0'] | [] | ['Evita V800', 'Evita V500']
2326 | Minute expired volume | DragerMedibus | Expired Minute Volume | ['L/min', '-'] | ['This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. Possible units: mL, L\xa0', 'For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. This variable is associated with UpperAlarmLimit attribute, which is available for Evita Infinity V500, Babylog VN500, Evita V300. This variable is associated with LowerAlarmLimit attribute, which is available for Evita Infinity V500, Babylog VN500, Evita V300. For Evita Infinity V500, Babylog VN500: This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes.\xa0'] | Double | VitalSigns | ['IntelliVue MX800\xa0', 'Evita Infinity V500 [MEDIBUS.X]\xa0', 'IntelliVue X3\xa0', 'Evita Infinity V800\xa0', 'Common\xa0', 'IntelliVue MP70\xa0'] | MINVOL | ['X3', 'MX800', 'Evita V800', 'Monnal T60', 'MP70', 'Evita V500']
2416 | Minute expired volume | DragerMedibus | Spontaneous Inspired Tidal Volume | mL | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. For Zeus, Zeus IE: This variable is associated with UpperAlarmLimit attribute. Insp. Tidal volume. | Double | VitalSigns | ['Evita Infinity V500 [MEDIBUS.X]\xa0', 'Evita Infinity V800\xa0'] | [] | ['Evita V800', 'Evita V500']
2432 | Expired Spontaneous Tidal Volume | DragerMedibus | Expired Spontaneous Tidal Volume | ['mL', '-'] | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. | Double | VitalSigns | ['IntelliVue MX800\xa0', 'Evita Infinity V500 [MEDIBUS.X]\xa0', 'IntelliVue X3\xa0', 'Evita Infinity V800\xa0', 'IntelliVue MP70\xa0'] | SpTVex | ['X3', 'MX800', 'Evita V800', 'MP70', 'Evita V500']
2564 | Peak Airways Pressure | DragerMedibus | Peak Airway Pressure | mbar | For Evita 4, Evita 2 dura, Evita XL: This variable is associated with UpperAlarmLimit attribute, which is also sent in separate variable 7040. For Evita 2, Evita, Savina: This variable is associated with UpperAlarmLimit attribute. For Fabius GS, Fabius GS premium, Fabius Tiro, Fabius plus, Fabius plus XL, Fabius OS, Fabius MRI: 1mbar = 1cmH2O. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300, Fabius Family. This variable is associated with UpperAlarmLimit attribute, which is available for Perseus A500, Atlan, Evita Infinity V500, Babylog VN500, Savina 300, Zeus IE, Oxylog 3000 plus, Carina, Evita V300, Fabius Family. This variable is associated with LowerAlarmLimit attribute, which is available for Zeus IE, Fabius Family. Peak inspiratory pressure, PIP. For Evita Infinity V500, Babylog VN500, Carina: This variable is associated with UpperAlarmLimit attribute. For Zeus, Zeus IE: This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. Peak Breathing Press. For Primus, Primus IE, Apollo, Pallas: Peak Breathing Pressure. | Double | VitalSigns | ['Evita Infinity V500 [MEDIBUS.X]\xa0', 'Evita Infinity V800\xa0'] | [] | ['Evita V800', 'Evita V500']
4003 | Breathing frequency | DragerMedibus | Generic Respiration Rate | 1/min | For Evita 4, Evita 2 dura, Evita XL: This variable is associated with UpperAlarmLimit attribute, which is also sent in separate variable 1339. Total Respiratory Rate (RRtotal). For Evita 2, Evita, Savina, Carina: This variable is associated with UpperAlarmLimit attribute. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300, Fabius Family. This variable is associated with UpperAlarmLimit attribute, which is available for Evita Infinity V500, Babylog VN500, Savina 300, Carina, Evita V300. Respiratory rate, RR. For Evita Infinity V500, Babylog VN500: This variable is associated with UpperAlarmLimit attribute. Respiratory Rate (RRtotal) RR. | Double | VitalSigns | ['Evita Infinity V500 [MEDIBUS.X]\xa0', 'Evita Infinity V800\xa0'] | [] | ['Evita V800', 'Evita V500']
5340 | Expired Mandatory Tidal Volume | DragerMedibus | Expired Mandatory Tidal Volume | mL | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. | Double | VitalSigns | ['Evita Infinity V500 [MEDIBUS.X]\xa0', 'Evita Infinity V800\xa0'] | [] | ['Evita V800', 'Evita V500']
171 | Device Date | RespironicsEsprit
172 | Device Time | RespironicsEsprit
1306 | Tidal Volume | RespironicsEsprit | Tidal Volume | ['mL', '-'] | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: Obsolete variable that may be removed in future release. This variable is marked as deprecated in order to synchronize with other models. Please use variable 60 instead. For Evita Infinity V500, Babylog VN500: Deprecated in order to synchronize with other models. Please use variable 60 instead. | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MX800\xa0', 'Evita Infinity V500 [MEDIBUS.X]\xa0', 'IntelliVue MP70\xa0'] | TV | ['MP70', 'Evita V500', 'MX800', 'X3']
1312 | Breathing Circuit Total Leakage | RespironicsEsprit
1364 | Low PEEP Alarm | RespironicsEsprit
2327 | Peak Inspiratory Pressure | RespironicsEsprit
2975 | Apnea Alarm | RespironicsEsprit
2978 | Low Minute Volume Alarm | RespironicsEsprit
4377 | Low Inspiratory Pressure Alarm | RespironicsEsprit
4378 | High Respiratory Rate Alarm | RespironicsEsprit
4390 | Device Alarm Silence Status | RespironicsEsprit
5945 | Total Respiration Rate | RespironicsEsprit
5973 | Circuit Occlusion Alarm | RespironicsEsprit
5975 | Low Battery Alarm | RespironicsEsprit
6216 | Current Ventilation Type | RespironicsEsprit
6266 | NPPV Mode Setting | RespironicsEsprit
6268 | NPPV Respiration Rate Setting | RespironicsEsprit
6269 | NPPV EPAP Setting | RespironicsEsprit
6270 | NPPV IPAP Setting | RespironicsEsprit
6271 | NPPV Inspiratory Time Setting | RespironicsEsprit
6272 | NPPV Rise Time Setting | RespironicsEsprit
6273 | NPPV I-Trigger Type | RespironicsEsprit
6277 | NPPV O2 Concentration Setting | RespironicsEsprit
6279 | NPPV Inspiratory Pressure Lower Alarm Limit | RespironicsEsprit
6281 | Tidal Volume Lower Alarm Limit | RespironicsEsprit | Tidal Volume Lower Alarm Limit | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | AlarmLimit: lowTV | ['MP70', 'MX800', 'X3']
6282 | NPPV Respiratory Rate Upper Alarm Limit | RespironicsEsprit
6283 | Minute Volume Lower Alarm Limit | RespironicsEsprit | Minute Volume Lower Alarm Limit | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | AlarmLimit: lowMV | ['MP70', 'MX800', 'X3']
6288 | Percent of Patient Triggered Breaths | RespironicsEsprit | Percent of Patient Triggered Breaths | % | [] | Double | OtherMeasuredData | Common | [] | V60
6289 | Safety Valve Alarm | RespironicsEsprit
6290 | Nonvolatile Memory Failure Alarm | RespironicsEsprit
6291 | Primary Alarm Failure | RespironicsEsprit
6292 | High Inspiratory Pressure Alarm | RespironicsEsprit
6293 | Air Valve Alarm | RespironicsEsprit
6294 | O2 Valve Alarm | RespironicsEsprit
6296 | O2 Supply Alarm | RespironicsEsprit | O2 Supply Alarm | unitless | [] | Boolean | Alarms | Monnal T60 | [] | Monnal T60
6299 | Low Tidal Volume Alarm | RespironicsEsprit
6300 | Low Expired Spontaneous Tidal Volume Alarm | RespironicsEsprit
6303 | High Enclosure Temperature Alarm | RespironicsEsprit
6305 | Low EPAP Alarm | RespironicsEsprit
6307 | Device Screen Lock Status | RespironicsEsprit
6376 | High Leak Alarm | RespironicsEsprit | High Leak Alarm | unitless | [] | Boolean | Alarms | Monnal T60 | [] | Monnal T60
6379 | Ratio of Inspiration Time to Total Breath Cycle | RespironicsEsprit | Ratio of Inspiration Time to Total Breath Cycle | - | This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. Possible units: % | Double | VitalSigns | Common | [] | Monnal T60
9223 | NPPV E-Cycle Type | RespironicsEsprit
2092 | Inspired Minute Volume | TaemaA | Inspired Minute Volume | ['L/min', '-'] | ['This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. Possible units: mL, L\xa0', 'For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Carina, Evita V300. This variable is associated with UpperAlarmLimit attribute, which is available for Carina. This variable is associated with LowerAlarmLimit attribute, which is available for Carina.\xa0'] | Double | VitalSigns | ['IntelliVue MX800\xa0', 'Evita Infinity V500 [MEDIBUS.X]\xa0', 'IntelliVue X3\xa0', 'Evita Infinity V800\xa0', 'Common\xa0', 'IntelliVue MP70\xa0'] | Available only in revF or earlier | ['X3', 'MX800', 'Evita V800', 'Monnal T60', 'MP70', 'Evita V500']
2434 | Spontaneous Expired Minute Volume | TaemaA | Spontaneous Expired Minute Volume | ['L/min', '-'] | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Primus, Primus Infinity Empowered, Apollo, Evita V300. | Double | VitalSigns | ['IntelliVue MX800\xa0', 'Evita Infinity V500 [MEDIBUS.X]\xa0', 'IntelliVue X3\xa0', 'Evita Infinity V800\xa0', 'IntelliVue MP70\xa0'] | %SpMV | ['X3', 'MX800', 'Evita V800', 'MP70', 'Evita V500']
26 |  |  | Device Inactivation Status | unitless | Possible values: "enabled" and "audio-paused". See "Multiset Alarms" on help file. | String | DeviceInfo | Common | Alarms | ['MP20', 'MP70', 'MX800', 'X3']
113 |  |  | Pleth Waveform | - | [] | Double | Waveforms | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Waves: Pleth | ['MP70', 'MX800', 'X3']
1314 |  |  | Spontaneous Respiration Rate | 1/min | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Zeus IE, Oxylog 3000 plus, Evita V300. This variable is associated with UpperAlarmLimit attribute, which is available for Oxylog 3000 plus. | Double | VitalSigns | ['Evita Infinity V500 [MEDIBUS.X]\xa0', 'Evita Infinity V800\xa0'] | [] | ['Evita V800', 'Evita V500']
1761 |  |  | Airway Resistance | mbar/L/s | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. | Double | VitalSigns | ['Evita Infinity V500 [MEDIBUS.X]\xa0', 'Evita Infinity V800\xa0'] | [] | ['Evita V800', 'Evita V500']
2433 |  |  | Mechanical Expired Minute Volume | L/min | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Primus, Primus Infinity Empowered, Apollo, Evita V300. Mandatory expiratory minute volume, MVemand | Double | VitalSigns | ['Evita Infinity V500 [MEDIBUS.X]\xa0', 'Evita Infinity V800\xa0'] | [] | ['Evita V800', 'Evita V500']
2570 |  |  | Spontaneous Tidal Volume | - | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Zeus IE, Evita V300. Possible units: L, mL. For Evita Infinity V500, Babylog VN500: Possible units: L, mL. | Double (Multiset) | VitalSigns | ['Evita Infinity V500 [MEDIBUS.X]\xa0', 'Evita Infinity V800\xa0'] | [] | ['Evita V800', 'Evita V500']
3326 |  |  | Device Average Gas Flow | L/min | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. Average device flow, FlowDev. | Double | OtherMeasuredData | ['Evita Infinity V500 [MEDIBUS.X]\xa0', 'Evita Infinity V800\xa0'] | [] | ['Evita V800', 'Evita V500']
3966 |  |  | Inspired Mandatory Tidal Volume | mL | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. | Double | VitalSigns | ['Evita Infinity V500 [MEDIBUS.X]\xa0', 'Evita Infinity V800\xa0'] | [] | ['Evita V800', 'Evita V500']
3968 |  |  | Spontaneous Fraction of Minute Volume | ['%', '-'] | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. Spontaneous breathing portion of minute volume, % Mvspon. For Evita Infinity V500, Babylog VN500: Spontaneous Fraction Min. Vol. %MVspon (MVspon : MVtotal). | Double | VitalSigns | ['IntelliVue MX800\xa0', 'Evita Infinity V500 [MEDIBUS.X]\xa0', 'IntelliVue X3\xa0', 'Evita Infinity V800\xa0', 'IntelliVue MP70\xa0'] | SpMVe | ['X3', 'MX800', 'Evita V800', 'MP70', 'Evita V500']
3971 |  |  | Negative Inspiratory Force | mbar | For Evita 4, Evita 2 dura, Evita XL: The value is not transmitted, if it is displayed on Evita 4 / Evita 2 dura as "[less than] –45 mbar". Firmware version 04.00 or higher For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. | Double | VitalSigns | ['Evita Infinity V500 [MEDIBUS.X]\xa0', 'Evita Infinity V800\xa0'] | [] | ['Evita V800', 'Evita V500']
4892 |  |  | Airway Correlation Coefficient | ['unitless', '-'] | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. Correlation factor, r2. For Evita Infinity V500, Babylog VN500: r². | Double | VitalSigns | ['IntelliVue MX800\xa0', 'Evita Infinity V500 [MEDIBUS.X]\xa0', 'IntelliVue X3\xa0', 'Evita Infinity V800\xa0', 'IntelliVue MP70\xa0'] | r | ['X3', 'MX800', 'Evita V800', 'MP70', 'Evita V500']
5141 |  |  | ECG Lead II Wave | - | [] | Double | Waveforms | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Waves: II | ['MP70', 'MX800', 'X3']
5142 |  |  | ECG Lead III Wave | - | [] | Double | Waveforms | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Waves: III | ['MP70', 'MX800', 'X3']
7640 |  |  | Mandatory Respiration Rate | ['1/min', '-'] | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Zeus IE, Evita V300. | Double | VitalSigns | ['IntelliVue MX800\xa0', 'Evita Infinity V500 [MEDIBUS.X]\xa0', 'IntelliVue X3\xa0', 'Evita Infinity V800\xa0', 'IntelliVue MP70\xa0'] | VENT: MRR | ['X3', 'MX800', 'Evita V800', 'MP70', 'Evita V500']
7642 |  |  | Mandatory Tidal Volume | - | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Babylog VN500, Babylog VN600, Babylog VN800, Zeus IE, Evita V300(added with SW 2.30 for Babylog VN500, added with SW 2.30 for Evita Infinity V500). Possible units: L, mL. Unit "L" available for Evita Infinity V500, Babylog VN500, Evita V300 (added with SW 2.30 for Babylog VN500, added with SW 2.30 for Evita Infinity V500). Unit "mL" available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Zeus IE, Evita V300. | ['Double\xa0', 'Double (Multiset)\xa0'] | VitalSigns | ['IntelliVue MX800\xa0', 'Evita Infinity V500 [MEDIBUS.X]\xa0', 'IntelliVue X3\xa0', 'Evita Infinity V800\xa0', 'IntelliVue MP70\xa0'] | VENT: MTV | ['X3', 'MX800', 'Evita V800', 'MP70', 'Evita V500']
8857 |  |  | Arterial Blood Pressure (ABP) Wave | - | [] | Double | Waveforms | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Waves: ABP | ['MP70', 'MX800', 'X3']
8870 |  |  | Impedance Respiration Wave | - | [] | Double | Waveforms | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Waves: Resp | ['MP70', 'MX800', 'X3']
9466 |  |  | Mandatory Triggered Respiration Rate | 1/min | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300(added with SW 2.42 for Evita Infinity V500, Evita V300 and Babylog VN500). Respiratory rate of triggered, mandatory breaths, RRtrig For Evita 4, Evita 2 dura, Evita XL: Available only with Evita XL firmware version 7.00 or higher. | Double | VitalSigns | ['Evita Infinity V500 [MEDIBUS.X]\xa0', 'Evita Infinity V800\xa0'] | [] | ['Evita V800', 'Evita V500']
9476 |  |  | Percent of Leakage | ['%', '-'] | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Evita V300. Leakage minute volume in % of inspiratory minute volume, % leak Leakage rel. % leak - Infinity V500 determines the difference between the delivered flow on the inspiratory side and the measured flow on the expiratory side. This difference provides a measure of the amount of leakage and is displayed by Infinity V500 as the leakage minute volume MVleak and relative as % leak (MVleak to MV). For Evita Infinity V500, Babylog VN500: Leakage rel. % leak. | Double | OtherMeasuredData | ['IntelliVue MX800\xa0', 'Evita Infinity V500 [MEDIBUS.X]\xa0', 'IntelliVue X3\xa0', 'Evita Infinity V800\xa0', 'IntelliVue MP70\xa0'] | %Leak | ['X3', 'MX800', 'Evita V800', 'MP70', 'Evita V500']
92001 |  |  | Ventilator Patient Disconnected Alarm | unitless | Consult detailed help for more information on this variable | Long | Alarms | Common | [] | V60
92011 |  |  | Ventilator Inoperative Alarm | unitless | Consult detailed help for more information on this variable | Long | Alarms | Common | [] | V60
2761 | Percent of Inspiration Time to Total Breath Cycle |  |  | % |  |  | VitalSigns
14 | Percent SpO2 | DragerMedibus | SpO2 | ['%', '-'] | [] | Double | VitalSigns | ['Common\xa0', 'PiCCO2\xa0'] | ['Continuous Parameter\xa0', 'SpO2\xa0'] | ['X3', 'MX800', 'MP20', 'PiCCO2', 'MP70', 'Expression IP5 et MR400']
15 | SpO2 pulse rate | DragerMedibus
20 | EtCO2 | DragerMedibus | End Tidal CO2 (EtCO2) | - | ['Possible units: mmHg or kPa. If PCMCIA protocol is used, the unit code of this variable is always 0 (unit unknown).\xa0', 'This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. Possible units: cmH2O, mbar, bar, hPa, kPa, mmHg, %\xa0', 'For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Evita V300, Fabius Family. This variable is associated with UpperAlarmLimit attribute, which is available for Perseus A500, Atlan, Evita Infinity V500, Babylog VN500, Savina 300, Primus, Primus IE, Apollo, Zeus IE, Oxylog 3000 plus, Evita V300. This variable is associated with LowerAlarmLimit attribute, which is available for Perseus A500, Atlan, Evita Infinity V500, Babylog VN500, Savina 300, Primus, Primus IE, Apollo, Zeus IE, Oxylog 3000 plus, Evita V300. Possible units: mmHg, %, kPa. This variable is in Multiset. See variables 2355 and 6551 for concentration and partial pressure respectively. For Evita Infinity V500, Babylog VN500: This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. end tidal CO2 etCO2 Note: etCO2 values and alarm limits are transmitted according to the configuration on the device screen (mmHg or Vol%). If kPa is configured be aware that the Medibus alarm limits will be transmitted with mmHg.\xa0'] | ['Double\xa0', 'Double (Multiset)\xa0'] | VitalSigns | ['Monnal T60\xa0', 'Common\xa0', 'Evita Infinity V500 [MEDIBUS.X]\xa0', 'Evita Infinity V800\xa0'] | Resp | ['Monnal T60', 'Evita V800', 'Expression IP5 et MR400', 'Evita V500']
62 | VM expiré mesuré | DragerMedibus
584 | Ventilation Mode | DragerMedibus | Ventilation Mode | unitless | ['The ventilation mode is reported in variable 584 and 2756. Variable 2756 contains mode ID numeric value, while variable 584 contains string representation which is the complete ventilation mode information concatenated from basic ventilation mode together with available options such as "Mode VC-CMV /AF". See variable 2756 for full list of supported ventilation mode.\xa0', 'Consult detailed help for more information on this variable\xa0'] | String | DeviceInfo | ['Common\xa0', 'Evita Infinity V500 [MEDIBUS.X]\xa0', 'Evita Infinity V800\xa0'] | [] | ['Monnal T60', 'Evita V800', 'Evita V500']
608 | EtCO2 mesurée | DragerMedibus | End Tidal CO2 | - | [] | Double | VitalSigns | Common | CO2 | ['MP20', 'MP70', 'MX800', 'X3']
616 | Airway Pressure | DragerMedibus | Airway Pressure | - | Measured in the inspiratory path. | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | AWPin | ['MP70', 'MX800', 'X3']
754 | Device Name | DragerMedibus | Device Name | unitless | ['This variable contains the device name defined by the manufacture. See variable 3591 for device model identification number.\xa0', 'This variable is obsolete for MultiFiltrate and might be removed in a future version.\xa0', 'The bed model name. \xa0', 'ASCII string up to 24 characters long.\xa0'] | String | DeviceInfo | ['Common\xa0', 'Bed Data Server (Common)\xa0'] | ['Device Information\xa0', 'Treatment Status/Data Management\xa0', 'BED\xa0'] | ['Total Scale', '5008', 'Evita V800', 'Monnal T60', 'Multifiltrate PRO', 'Evita V500']
1320 | Mandatory respiratory rate setting | DragerMedibus | Mandatory Respiration Rate Setting | 1/min | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300, Fabius Family. Respiratory rate, RR. For Evita Infinity V500, Babylog VN500: Respiratory Rate RR. | Double | Settings | ['Evita Infinity V500 [MEDIBUS.X]\xa0', 'Evita Infinity V800\xa0'] | [] | ['Evita V800', 'Evita V500']
5044 | Tidal Volume Setting | DragerMedibus | Inspired Tidal Volume Setting | ['mL', '-'] | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Zeus IE, Evita V300. Inspiratory tidal volume, Vti. For Evita Infinity V500, Babylog VN500: See also variable 1321 and 2771. For Zeus, Zeus IE: This variable contains same information as variable 1321, except that the value of different units (L or mL) are sent in separate variables. See also variable 2771. | Double | Settings | ['IntelliVue MX800\xa0', 'Evita Infinity V500 [MEDIBUS.X]\xa0', 'IntelliVue X3\xa0', 'Evita Infinity V800\xa0', 'IntelliVue MP70\xa0'] | sTVin | ['X3', 'MX800', 'Evita V800', 'MP70', 'Evita V500']
1332 | Pressure Support Level Aboce PEEP Setting | DragerMedibus | Pressure Support Level Above PEEP Setting | ['mbar', '-'] | For Evita 4, Evita 2 dura, Evita XL: This variable is calculated from variables 1392 and 2104. Available only with Evita XL. For Medibus.X compatible device, Perseus A500, Atlan, Evita V300, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300, Fabius Family. Pressure amplitude above PEEP in Pressure Support, DPsupp. For Evita Infinity V500, Babylog VN500: ΔPsupp. For Carina: Assisted spontaneous breathing. | Double | Settings | ['IntelliVue MX800\xa0', 'Evita Infinity V500 [MEDIBUS.X]\xa0', 'IntelliVue X3\xa0', 'Evita Infinity V800\xa0', 'IntelliVue MP70\xa0'] | sdPEEP | ['X3', 'MX800', 'Evita V800', 'MP70', 'Evita V500']
1333 | Vminute spontané | DragerMedibus | Spontaneous Minute Volume | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | VENT: SpMV | ['MP70', 'MX800', 'X3']
1349 | I/E ratio - I | DragerMedibus
1350 | I/E ratio - E | DragerMedibus
1381 | Inspiratory flow setting | DragerMedibus | Inspiratory Flow Setting | ['L/min', '-'] | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300, Fabius Family. | Double | Settings | ['IntelliVue MX800\xa0', 'Evita Infinity V500 [MEDIBUS.X]\xa0', 'IntelliVue X3\xa0', 'Evita Infinity V800\xa0', 'IntelliVue MP70\xa0'] | sInsFl | ['X3', 'MX800', 'Evita V800', 'MP70', 'Evita V500']
1383 | Flow trigger setting | DragerMedibus | Inspiratory Flow Trigger Setting | ['L/min', '-'] | ['Possible units: mL/min, mL/h, L/s, L/min, L/h\xa0', 'For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Evita V300, Fabius Family. Flow trigger threshold, Flow trigger. For Primus, Primus IE, Apollo, Pallas: Flow trigger.\xa0', 'Delivered by the Evita 2 VueLink Driver.\xa0'] | Double | Settings | ['IntelliVue MX800\xa0', 'Evita Infinity V500 [MEDIBUS.X]\xa0', 'IntelliVue X3\xa0', 'Evita Infinity V800\xa0', 'Common\xa0', 'IntelliVue MP70\xa0'] | VENT: sTrgFl | ['X3', 'MX800', 'Evita V800', 'Monnal T60', 'MP70', 'Evita V500']
1384 | Respiration Rate setting | DragerMedibus
1386 | Intermittent PEEP Setting | DragerMedibus | Intermittent PEEP Setting | mbar | For Fabius GS, Fabius GS premium, Fabius Tiro, Fabius plus, Fabius plus XL, Fabius OS, Fabius MRI: 1mbar = 1cmH2O. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Evita V300. Additional intermittent PEEP for sighs, PEEPsigh. For Evita Infinity V500, Babylog VN500: Interm. PEEP ΔintPEEP. For Zeus, Zeus IE: PEEP for all vent mode different from Man/Spont. | Double | Settings | ['Evita Infinity V500 [MEDIBUS.X]\xa0', 'Evita Infinity V800\xa0'] | [] | ['Evita V800', 'Evita V500']
1387 | Low pressure setting | DragerMedibus | APRV Low Pressure Setting | ['mbar', '-'] | For Evita: BIPAP low pressure For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Evita V300. Lower pressure level during APRV, Plow. For Evita Infinity V500, Babylog VN500: APRV Low Pressure Plow. | Double | Settings | ['IntelliVue MX800\xa0', 'Evita Infinity V500 [MEDIBUS.X]\xa0', 'IntelliVue X3\xa0', 'Evita Infinity V800\xa0', 'IntelliVue MP70\xa0'] | sPSVl | ['X3', 'MX800', 'Evita V800', 'MP70', 'Evita V500']
1388 | Low time setting | DragerMedibus | APRV Low Time Setting | s | For Evita: BIPAP low time For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Evita V300. Time of lower pressure level during APRV, Tlow. For Evita Infinity V500, Babylog VN500: APRV Low Time Tlow. | Double | Settings | ['Evita Infinity V500 [MEDIBUS.X]\xa0', 'Evita Infinity V800\xa0'] | [] | ['Evita V800', 'Evita V500']
1389 | High pressure setting | DragerMedibus | APRV High Pressure Setting | ['mbar', '-'] | For Evita: BIPAP high pressure For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Evita V300. Upper pressure level during APRV, Phigh. For Evita Infinity V500, Babylog VN500: APRV High Pressure Phigh. | Double | Settings | ['IntelliVue MX800\xa0', 'Evita Infinity V500 [MEDIBUS.X]\xa0', 'IntelliVue X3\xa0', 'Evita Infinity V800\xa0', 'IntelliVue MP70\xa0'] | sPSVh | ['X3', 'MX800', 'Evita V800', 'MP70', 'Evita V500']
1390 | High time setting | DragerMedibus | APRV High Time Setting | s | For Evita: BIPAP high time For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Evita V300. Time of upper pressure level during APRV, Thigh. For Evita Infinity V500, Babylog VN500: APRV High Time Thigh. | Double | Settings | ['Evita Infinity V500 [MEDIBUS.X]\xa0', 'Evita Infinity V800\xa0'] | [] | ['Evita V800', 'Evita V500']
1391 | Apnea time setting | DragerMedibus | Apnea Time Setting | ['s', '-'] | ['Possible units: ms, sec, min, h\xa0', 'For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Oxylog 3000 plus, Carina, Evita V300. For Evita Infinity V500, Babylog VN500: Apnea Time Tapn. For Carina: Apnea alarm time.\xa0'] | Double | Settings | ['IntelliVue MX800\xa0', 'Evita Infinity V500 [MEDIBUS.X]\xa0', 'IntelliVue X3\xa0', 'Evita Infinity V800\xa0', 'Common\xa0', 'IntelliVue MP70\xa0'] | sAlnTi | ['X3', 'MX800', 'Evita V800', 'Monnal T60', 'MP70', 'Evita V500']
1392 | Spontaneous breath setting | DragerMedibus
1393 | Max inspired pressure setting | DragerMedibus | Maximum Inspired Pressure Setting | mbar | For Fabius GS, Fabius GS premium, Fabius Tiro, Fabius plus, Fabius plus XL, Fabius OS, Fabius MRI: 1mbar = 1cmH2O. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300, Fabius Family. Pressure limitation, Pmax. For Evita Infinity V500, Babylog VN500: Max Pressure Pmax. For Carina: Pressure limitation. For Primus, Primus IE, Apollo, Pallas: Max. insp. Airway pressure. | Double | Settings | ['Evita Infinity V500 [MEDIBUS.X]\xa0', 'Evita Infinity V800\xa0'] | [] | ['Evita V800', 'Evita V500']
1396 | Tachypnea Respiratory Rate setting | DragerMedibus
1422 | Rapid Shallow Breathing Index | DragerMedibus | Rapid Shallow Breathing Index | - | For Evita 4, Evita 2 dura, Evita XL: Firmware version 04.00 or higher For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Atlan, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Evita V300. Possible units: 1/L * min, 1/mL*min. Unit "1/L * min" is available for Evita Infinity V500, Babylog VN500, Savina 300, Evita V300. Unit "1/mL*min" is available for Evita Infinity V500, Babylog VN500, Evita V300 (added with SW 2.30 for Babylog VN500, added with SW 2.30 for Evita Infinity V500). | ['Double\xa0', 'Double (Multiset)\xa0'] | VitalSigns | ['IntelliVue MX800\xa0', 'Evita Infinity V500 [MEDIBUS.X]\xa0', 'IntelliVue X3\xa0', 'Evita Infinity V800\xa0', 'IntelliVue MP70\xa0'] | RSBI | ['X3', 'MX800', 'Evita V800', 'MP70', 'Evita V500']
1596 | Minimum pressure | DragerMedibus
2102 | Temps insp réglé | DragerMedibus | Inspiratory Time Setting | ['s', '-'] | ['For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300, Fabius Family.\xa0', 'Possible units: ms, sec, min, h\xa0'] | Double | Settings | ['Common\xa0', 'Evita Infinity V500 [MEDIBUS.X]\xa0', 'Evita Infinity V800\xa0'] | [] | ['Monnal T60', 'Evita V800', 'Evita V500']
2104 | PEEP Setting | DragerMedibus | Positive End Expiratory Pressure (PEEP) Setting | ['mbar', '-'] | ['For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300, Fabius Family. For Evita Infinity V500, Babylog VN500: PEEP / CPAP PEEP. For Zeus, Zeus IE: PEEP (CPAP). PEEP for mode Man/Spont.\xa0', 'Possible units: cmH2O, mbar, bar, hPa, kPa, mmHg\xa0'] | Double | Settings | ['IntelliVue MX800\xa0', 'Evita Infinity V500 [MEDIBUS.X]\xa0', 'IntelliVue X3\xa0', 'Evita Infinity V800\xa0', 'Common\xa0', 'IntelliVue MP70\xa0'] | sEPAP | ['X3', 'MX800', 'Evita V800', 'Monnal T60', 'MP70', 'Evita V500']
2110 | FiO2 Setting | DragerMedibus | Inspired O2 (FiO2) Setting | ['%', '-'] | ['Possible units: %\xa0', 'For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Zeus IE, Oxylog 3000 plus, Carina, Evita V300. For Primus, Primus IE, Apollo, Pallas: Not in device Apollo and Pallas\xa0'] | Double | Settings | ['Monnal T60\xa0', 'Evita Infinity V500 [MEDIBUS.X]\xa0', 'Evita Infinity V800\xa0'] | [] | ['Monnal T60', 'Evita V800', 'Evita V500']
2324 | Inspired Tidal Volume | DragerMedibus | Inspired Tidal Volume | ['mL', '-'] | ['This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. Possible units: mL, L\xa0', 'For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Carina, Evita V300. This variable is associated with UpperAlarmLimit attribute, which is available for Evita Infinity V500, Zeus IE, Evita V300. This variable is associated with LowerAlarmLimit attribute, which is available for Evita Infinity V500, Evita V300. For Primus, Primus IE, Apollo, Pallas: Insp. Tidal Volume.\xa0'] | Double | VitalSigns | ['IntelliVue MX800\xa0', 'Evita Infinity V500 [MEDIBUS.X]\xa0', 'IntelliVue X3\xa0', 'Evita Infinity V800\xa0', 'Common\xa0', 'IntelliVue MP70\xa0'] | TVin | ['X3', 'MX800', 'Evita V800', 'Monnal T60', 'MP70', 'Evita V500']
2325 | Expired Tidal Volume | DragerMedibus | Expired Tidal Volume | ['mL', '-'] | ['This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. Possible units: mL, L\xa0', 'For Evita XL: Neo expired tidal volume. Available only with firmware version 7.01 or higher. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Evita V300.\xa0'] | Double | VitalSigns | ['IntelliVue MX800\xa0', 'Evita Infinity V500 [MEDIBUS.X]\xa0', 'IntelliVue X3\xa0', 'Evita Infinity V800\xa0', 'Common\xa0', 'IntelliVue MP70\xa0'] | TVexp | ['X3', 'MX800', 'Evita V800', 'Monnal T60', 'MP70', 'Evita V500']
2565 | Plateau Airway Pressure | DragerMedibus | Plateau Airway Pressure | mbar | For Fabius GS, Fabius GS premium, Fabius Tiro, Fabius plus, Fabius plus XL, Fabius OS, Fabius MRI: 1mbar = 1cmH2O. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Evita V300, Fabius Family. For Evita Infinity V500, Babylog VN500: Only available for Evita Infinity V500 | Double | VitalSigns | ['Evita Infinity V500 [MEDIBUS.X]\xa0', 'Evita Infinity V800\xa0'] | [] | ['Evita V800', 'Evita V500']
2583 | I/E Ratio Setting | DragerMedibus | I:E Ratio Setting | unitless | ['String format: 1:x if E > I, x:1 if I > E, or 1:1 if I = E. The useless leading or ending “0” or “.” is not transmitted, and the maximum digits after the floating point is 2. For Carina, Primus, Primus IE, Apollo, Pallas, Perseus A500: This variable is calculated by the DDI from respiration rate and inspiratory time settings. For Medibus.X compatible device: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300, Fabius Family. For Primus, Primus IE, Apollo, Perseus A500, Carina and Oxylog 3000 plus, it is calculated by the DDI from respiration rate and inspiratory time settings. For Atlan, Evita Infinity V500, Babylog VN500, Savina 300, Zeus IE, Evita V300, and Fabius Family, it is calculated by the DDI from inspiratory and expiratory components. For Evita, Evita 2, Evita 4, Evita 2 dura, and Evita XL, Fabius GS, Fabius GS premium, Fabius Tiro, Fabius plus, Fabius plus XL, Fabius OS, Fabius MRI, Evita Infinity V500, Babylog VN500, Zeus IE, Evita V300: This variable is calculated by the DDI from inspiratory and expiratory components of I:E ratio sent by the devices.\xa0', 'This variable is associated with ChannelID identifier. This variable is calculated from variable 2773, which is sent as a string in the format "1:xx.xx".\xa0'] | ['Double\xa0', 'String\xa0'] | Settings | ['IntelliVue MX800\xa0', 'Evita Infinity V500 [MEDIBUS.X]\xa0', 'IntelliVue X3\xa0', 'Evita Infinity V800\xa0', 'Common\xa0', 'IntelliVue MP70\xa0'] | Calculated Variables | ['X3', 'MX800', 'Evita V800', 'Monnal T60', 'MP70', 'Evita V500']
2771 | Tidal Volume Setting | DragerMedibus | Tidal Volume Setting | L | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300, Fabius Family. Inspiratory tidal volume, Vti. For Evita Infinity V500, Babylog VN500: See also variable 1321 and 5044. For Zeus, Zeus IE: This variable contains same information as variable 1321, except that the value of different units (L or mL) are sent in separate variables. See also variable 5044. | Double | Settings | ['Evita Infinity V500 [MEDIBUS.X]\xa0', 'Evita Infinity V800\xa0'] | [] | ['Evita V800', 'Evita V500']
2784 | ASB ramp time setting | DragerMedibus | ASB Ramp Time Setting | ['s', '-'] | ['Possible units: msec, sec, min\xa0', 'For Evita 2: Firmware version 2.00 or higher. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300. Pressure rise time, Slope. For Carina: ASB ramp. For Primus, Primus IE, Apollo, Pallas: Slope time (Ramp).\xa0'] | Double | Settings | ['Common\xa0', 'Evita Infinity V500 [MEDIBUS.X]\xa0', 'Evita Infinity V800\xa0'] | [] | ['V60', 'Evita V800', 'Evita V500']
2794 | VT inspiré réglé | DragerMedibus
2987 | Inspired Pressure Setting | DragerMedibus | Inspired Pressure Setting | mbar | For Evita 4, Evita 2 dura, Evita XL: Available only with EvitaXL firmware version 7.00 or higher. For Fabius GS, Fabius GS premium, Fabius Tiro, Fabius plus, Fabius plus XL, Fabius OS, Fabius MRI: 1mbar = 1cmH2O. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300, Fabius Family. | Double | Settings | ['Evita Infinity V500 [MEDIBUS.X]\xa0', 'Evita Infinity V800\xa0'] | [] | ['Evita V800', 'Evita V500']
3853 | Patient Type | DragerMedibus | Patient Type | unitless | ['Possible values: "unspecified", "adult", "pediatric", and "neonatal".\xa0', 'This variable contains patient type in string format, which is also send as a numerical value in variable 2467. See detailed description of variable 2467 for the list of supported values.\xa0', 'Value is ADULT, CHILD, INFANT, or BABY.\xa0'] | String | PatientInfo | ['Common\xa0', 'Evita Infinity V500 [MEDIBUS.X]\xa0', 'Evita Infinity V800\xa0'] | Patient Demographics | ['X3', 'MX800', 'MP20', 'Evita V800', 'Monnal T60', 'MP70', 'Evita V500']
3970 | Inspiratory spont support volume | DragerMedibus
4011 | Flow assist setting | DragerMedibus | Flow Assist Level Setting | 0.1 mbar/L/s | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. Flow assist during PPS ventilation, Flow Assist. For Evita Infinity V500, Babylog VN500: Flow Assist. | Double | Settings | ['Evita Infinity V500 [MEDIBUS.X]\xa0', 'Evita Infinity V800\xa0'] | [] | ['Evita V800', 'Evita V500']
4012 | Volume assist setting | DragerMedibus | Volume Assist Level Setting | ['0.1 mbar/L', '-'] | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. Volume assist during PPS ventilation, Vol. Assist. For Evita Infinity V500, Babylog VN500: Vol. Assist. | Double | Settings | ['IntelliVue MX800\xa0', 'Evita Infinity V500 [MEDIBUS.X]\xa0', 'IntelliVue X3\xa0', 'Evita Infinity V800\xa0', 'IntelliVue MP70\xa0'] | sWOBva | ['X3', 'MX800', 'Evita V800', 'MP70', 'Evita V500']
4013 | Backup Frequency Setting | DragerMedibus
4014 | Backup Tidal Volume Setting | DragerMedibus | Backup Tidal Volume Setting | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sTVbak | ['MP70', 'MX800', 'X3']
6389 | Endotracheal Tube Size | DragerMedibus | Endotracheal Tube Size | mm | For Evita 4, Evita 2 dura, Evita XL: Possible units: "0.1mm" in firmware versions 7.00 and 7.01, "mm" otherwise. Available only with EvitaXL firmware version 7.00 or higher. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Evita V300. Inner diameter of the endotracheal/tracheostoma tube. For Evita Infinity V500, Babylog VN500: ATC Tube Diameter. | Double | DeviceInfo | ['Evita Infinity V500 [MEDIBUS.X]\xa0', 'Evita Infinity V800\xa0'] | [] | ['Evita V800', 'Evita V500']
6702 | Inspiratory Component I:E ratio Setting | DragerMedibus | Inspiratory Component of I:E Ratio Setting | ['unitless', '-'] | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Zeus IE, Evita V300, Fabius Family, Atlan For Evita Infinity V500: Partial information. Please use variable 2583 instead. Obsolete variable that may be removed in future release. | Double | Settings | ['IntelliVue MX800\xa0', 'Evita Infinity V500 [MEDIBUS.X]\xa0', 'IntelliVue X3\xa0', 'Evita Infinity V800\xa0', 'IntelliVue MP70\xa0'] | sI | ['X3', 'MX800', 'Evita V800', 'MP70', 'Evita V500']
6703 | Expiratory Component I:E ratio Setting | DragerMedibus | Expiratory Component of I:E Ratio Setting | ['unitless', '-'] | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800,Savina 300, Zeus IE, Evita V300, Fabius Family, Atlan For Evita Infinity V500: Partial information. Please use variable 2583 instead. Obsolete variable that may be removed in future release. | Double | Settings | ['IntelliVue MX800\xa0', 'Evita Infinity V500 [MEDIBUS.X]\xa0', 'IntelliVue X3\xa0', 'Evita Infinity V800\xa0', 'IntelliVue MP70\xa0'] | sE | ['X3', 'MX800', 'Evita V800', 'MP70', 'Evita V500']
6735 | Lung Elastance | DragerMedibus | Lung Elastance | ['mbar/L', '-'] | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. | Double | VitalSigns | ['IntelliVue MX800\xa0', 'Evita Infinity V500 [MEDIBUS.X]\xa0', 'IntelliVue X3\xa0', 'Evita Infinity V800\xa0', 'IntelliVue MP70\xa0'] | E | ['X3', 'MX800', 'Evita V800', 'MP70', 'Evita V500']
7637 | Artificial Airway Compensation Status | DragerMedibus | Artificial Airway Compensation Status | % | For Evita 4, Evita 2 dura, Evita XL: Tube Compensation. Available only with EvitaXL firmware version 7.00 or higher. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. Compensation rate during ATC, % comp. For Evita Infinity V500, Babylog VN500: ATC Compensation Compens. | Double | DeviceInfo | ['Evita Infinity V500 [MEDIBUS.X]\xa0', 'Evita Infinity V800\xa0'] | [] | ['Evita V800', 'Evita V500']
7931 | End Tidal CO2 | DragerMedibus
9146 | End Inspiratory Pressure | DragerMedibus | End Inspiratory Pressure | mbar | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. End-inspiratory pressure, EIP. | Double | VitalSigns | ['Evita Infinity V500 [MEDIBUS.X]\xa0', 'Evita Infinity V800\xa0'] | [] | ['Evita V800', 'Evita V500']
9475 | Volume of Leakage | DragerMedibus | Volume of Leakage (MVLeak) | L/min | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Carina, Evita V300. Leakage minute volume, Mvleak. The leakage is determined by the difference between the delivered flow on the inspiratory side and the measured flow on the expiratory side, which is displayed as the leakage minute volume MVleak and relative as % leak (MVleak to MV). For Evita Infinity V500, Babylog VN500: Mvleak. The leakage is determined by the difference between the delivered flow on the inspiratory side and the measured flow on the expiratory side, which is displayed as the leakage minute volume MVleak and relative as % leak (MVleak to MV). For Carina: Leakage minute volume. For Evita XL: Leakage minute volume. Available since device SW version 7.02. | Double | OtherMeasuredData | ['Evita Infinity V500 [MEDIBUS.X]\xa0', 'Evita Infinity V800\xa0'] | [] | ['Evita V800', 'Evita V500']
9477 | Inspiratory Component Of Spontaneous I:E ratio | DragerMedibus | Inspiratory Component of Spontaneous I:E Ratio | unitless | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: Obsolete variable that may be removed in future release. Partial information of I:E ratio. For Evita Infinity V500: Partial information. Please use variable 5815 instead. Obsolete variable that may be removed in future release. | Double | VitalSigns | ['Evita Infinity V500 [MEDIBUS.X]\xa0', 'Evita Infinity V800\xa0'] | [] | ['Evita V800', 'Evita V500']
9478 | Expiratory Component Of Spontaneous I:E ratio | DragerMedibus | Expiratory Component of Spontaneous I:E Ratio | unitless | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: Obsolete variable that may be removed in future release. Partial information of I:E ratio. For Evita Infinity V500: Partial information. Please use variable 5815 instead. Obsolete variable that may be removed in future release. | Double | VitalSigns | ['Evita Infinity V500 [MEDIBUS.X]\xa0', 'Evita Infinity V800\xa0'] | [] | ['Evita V800', 'Evita V500']
997 | ModuleStatusStep | FreseniusVialC
1087 | BolusInfusedVolume | FreseniusVialC
1447 | PVolInf | FreseniusVialC
1448 | Rate | FreseniusVialC
1550 | TargetConcentration | FreseniusVialC
1880 | TotVol | FreseniusVialC
2055 | Drug | FreseniusVialC
2755 | Rate | FreseniusVialC
3609 | ModuleStatus | FreseniusVialC
3655 | InfusionPumpType | FreseniusVialC
5497 | ModeStatus | FreseniusVialC | Pump Infusion Mode | - | [] | Double | DeviceInfo | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sPuMod | ['MP70', 'MX800', 'X3']
6210 | SerialNo | FreseniusVialC
1 | Heart Rate | PhilipsIIC | Heart Rate | 1/min | ['This variable is associated with a SourceID identifier. For Expression, the SourceID is always ""MEAS_UNKNOWN" if using Expression Plus protocol.\xa0', 'The MeasurementTime of BP variables are only updated when a new measurement is performed. In between measurements, the last valid values are sent repeatedly.', 'This variable is associated with LowerAlarmLimit and UpperAlarmLimit attributes.\xa0'] | ['Double\xa0', 'Long\xa0', 'Short\xa0'] | VitalSigns | Common | ['Continuous Parameter\xa0', 'HR\xa0'] | ['CardioQ', 'Artis', '5008', 'PiCCO2', 'Expression IP5 et MR400']
922 | SimpleBolusVol | FreseniusE
1444 | SerialNumber | FreseniusE
1450 | InfusionPumpType | FreseniusE | Infusion Pump Type | unitless | 0: Smart or Go 2: Vision or Go+ | UByte | DeviceInfo | Common | [] | Applix Smart
1451 | PumpInfusionStatus | FreseniusE | Pump Infusion Status | ['unitless', '-'] | Consult detailed help for more information on this variable | ['UByte\xa0', 'Double\xa0'] | DeviceInfo | ['IntelliVue X3\xa0', 'Common\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | PuStat | ['Applix Smart', 'MP70', 'MX800', 'X3']
3638 | InfDoseRate | FreseniusE
4254 | InfusionRate | FreseniusE
5490 | TotVol | FreseniusE | Cumulated Infused Volume | mL | [] | UShort | VitalSigns | Common | [] | Applix Smart
3642 | PumpStatus | KendallA
7551 | FeedRate | KendallA
7550 | TotalVolume | KendallA
2201 | SerialNumber | KendallA | Serial Number | unitless | This variable is obsolete because it may contains incorrect information. Please use variables 764, 765 and 4201 must be used instead. | String | DeviceInfo | Common | Obsolete Variables | ['MP20', 'MP70', 'MX800', 'X3']
9009 | ErrorStatus | KendallA
3003 | Blood Flow Rate Setting | EdwardsA | Blood Flow Rate Setting | ['-', 'mL/min'] | [] | ['UShort\xa0', 'Double\xa0', 'Long\xa0'] | Settings | ['IntelliVue X3\xa0', 'Common\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | ['Dialysis Settings\xa0', 'sBFl\xa0'] | ['Multifiltrate PRO', 'X3', 'Artis', 'MX800', '5008', 'MP70']
4229 | Fluid Removal Cumulated Volume | EdwardsA | Fluid Removal Cumulated Volume | ['mL', '-'] | Effective total withdrawal, total | ['Double\xa0', 'ULong\xa0', 'Long\xa0'] | VitalSigns | ['IntelliVue X3\xa0', 'Common\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | ['Vufr\xa0', 'Dialysis Course\xa0'] | ['Multifiltrate PRO', 'X3', 'Artis', 'MX800', '5008', 'MP70']
3024 | Access Pod Pressure | EdwardsA | Access Pod Pressure | - | [] | Double | OtherMeasuredData | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | AccP | ['MP70', 'MX800', 'X3']
3025 | Return Pod Pressure | EdwardsA | Return Pod Pressure | - | [] | Double | OtherMeasuredData | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | RetP | ['MP70', 'MX800', 'X3']
3029 | Transmembrane Pressure | EdwardsA | Transmembrane Pressure | ['-', 'mmHg'] | This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. | ['Double\xa0', 'Long\xa0'] | VitalSigns | ['IntelliVue MX800\xa0', 'Multifiltrate PRO\xa0', 'IntelliVue X3\xa0', 'Common\xa0', 'IntelliVue MP70\xa0'] | ['TMP\xa0', 'Dialysis Course\xa0'] | ['Multifiltrate PRO', 'X3', 'Artis', 'MX800', '5008', 'MP70']
21 | Inspired CO2 (InCO2) Pressure | TaemaA | Inspired CO2 (InCO2) | - | Possible units: mmHg or kPa. If PCMCIA protocol is used, the unit code of this variable is always 0 (unit unknown). | Double | VitalSigns | Common | Resp | Expression IP5 et MR400
22 | Respiration Rate | TaemaA | Respiration Rate | ['1/min', '-'] | This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. Possible units: 1/min | Double | VitalSigns | Common | ['RR\xa0', 'Resp\xa0'] | ['X3', 'MX800', 'MP20', 'Monnal T60', 'MP70', 'Expression IP5 et MR400']
114 | CO2 Concentration Wave | TaemaA | CO2 Concentration Waveform | - | For Evita 4, Evita 2 dura, Evita XL: Possible units: mmHg, kPa, and % For Evita 2: Possible units: mmHg, kPa, and % Firmware version 2.00 or higher. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Evita V300. Possible units: mmHg, %, kPa. Unit "mmHg" is available for Perseus A500, Atlan, Evita Infinity V500, Babylog VN500, Savina 300, Primus, Primus IE, Apollo, Zeus IE, Oxylog 3000 plus, Evita V300. Unit "kPa" and "%" are available for Perseus A500, Atlan, Savina 300, Primus, Primus IE, Apollo, Zeus IE, Oxylog 3000 plus. See variables 8866 and 7872 for concentration and partial pressure respectively. For Evita Infinity V500, Babylog VN500: Possible units: mmHg, %. Expiratory CO2 concentration. For Zeus, Zeus IE, Primus, Primus IE, Apollo, Pallas, Vamos, Vamos plus: Possible units: mmHg, %, kPa. | Double | Waveforms | ['Evita Infinity V500 [MEDIBUS.X]\xa0', 'Evita Infinity V800\xa0'] | [] | ['Evita V800', 'Evita V500']
115 | O2 Concentration Wave | TaemaA
118 | Airway Flow Wave | TaemaA | Airway Flow Wave | L/min | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300, Fabius Family. For Evita Infinity V500, Babylog VN500, Carina, Zeus, Zeus IE, Primus, Primus IE, Apollo, Pallas: Flow (insp./exp.) | Double | Waveforms | ['Evita Infinity V500 [MEDIBUS.X]\xa0', 'Evita Infinity V800\xa0'] | [] | ['Evita V800', 'Evita V500']
169 | Patient Age | TaemaA | Patient Age | - | Patient age. The Monitor device calculates the age of the patient from the birth data manually entered on the Monitor. | Double | PatientInfo | Common | Patient Demographics | ['MP20', 'MP70', 'MX800', 'X3']
541 | Minimum Tidal Volume Threshold | TaemaA
544 | Minimum Respiration Rate Threshold Setting | TaemaA | Minimum Respiration Rate Threshold Setting | - | Possible units: 1/min | Double | Settings | Common | [] | Monnal T60
618 | N2O Concentration Wave | TaemaA | Airway N2O Wave | - | [] | Double | Waveforms | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Waves: N2O | ['MP70', 'MX800', 'X3']
638 | Apnea Index | TaemaA
756 | Patient Height | TaemaA | Patient Height | cm | [] | Short | PatientInfo | Common | Patient Setting | PiCCO2
757 | Patient Weight | TaemaA | Patient Weight | kg | The last patient weight captured by a caregiver from the bedside displayed in kilograms. -9999.9 -- 9999.9 | Double | PatientInfo | ['Common\xa0', 'Bed Data Server (Common)\xa0'] | ['Patient Setting\xa0', 'BED\xa0'] | ['5008', 'Artis', 'PiCCO2', 'Total Scale']
784 | Inspiration Time Percent Setting | TaemaA | Inspiration Time Percent Setting | - | Possible units: % | Double | Settings | Common | [] | Monnal T60
1177 | Identified Agent Name | TaemaA
1187 | Low Inspired Tidal Volume Alarm | TaemaA | Low Inspired Tidal Volume Alarm | unitless | [] | Boolean | Alarms | Common | [] | Monnal T60
1188 | Low Expired Tidal Volume Alarm | TaemaA | Low Expired Tidal Volume Alarm | unitless | [] | Boolean | Alarms | Common | [] | Monnal T60
1321 | Tidal Volume Setting | TaemaA | Tidal Volume Setting | - | ['For Evita 4, Evita 2 dura, Evita XL, Evita 2, Evita, Savina, Fabius GS, Fabius GS premium, Fabius Tiro, Fabius plus, Fabius plus XL, Fabius OS, Fabius MRI, Carina, Zeus, Zeus IE, Primus, Primus IE, Apollo, Pallas: Inspired tidal volume setting. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: Obsolete variable that may be removed in future release. Variable is in Multiset format. Please use separate variable 2771 and 5044 instead. For Evita Infinity V500, Babylog VN500: Variable 2771 and 5044 contains same information but in different units. This variable combines these two variables and sends whichever is significant. For Zeus, Zeus IE: Obsolete variable that may be removed in future release. Multiset variable. Please use variables 2771 and 5044 instead.\xa0', 'Possible units: mL, L\xa0'] | ['Double\xa0', 'Double (Multiset)\xa0'] | Settings | ['IntelliVue MX800\xa0', 'Evita Infinity V500 [MEDIBUS.X]\xa0', 'IntelliVue X3\xa0', 'Common\xa0', 'IntelliVue MP70\xa0'] | sTV | ['X3', 'MX800', 'Monnal T60', 'MP70', 'Evita V500']
1324 | Sigh Time Setting | TaemaA | Sigh Interval Setting | ['s', '-'] | ['For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. Time interval between sighs, Interval sigh.\xa0', 'Possible units: ms, sec, min, h\xa0'] | Double | Settings | ['IntelliVue MX800\xa0', 'Evita Infinity V500 [MEDIBUS.X]\xa0', 'IntelliVue X3\xa0', 'Evita Infinity V800\xa0', 'Common\xa0', 'IntelliVue MP70\xa0'] | sTsigh | ['X3', 'MX800', 'Evita V800', 'Monnal T60', 'MP70', 'Evita V500']
1326 | Sigh Number Setting | TaemaA | Sigh Number Setting | unitless | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. Number of sigh cycles, Cycles sigh. | Double | Settings | ['Evita Infinity V500 [MEDIBUS.X]\xa0', 'Evita Infinity V800\xa0'] | [] | ['Evita V800', 'Evita V500']
1327 | Sigh Pressure Limit Setting | TaemaA | Sigh Pressure Limit Setting | - | Possible units: cmH2O, mbar, bar, hPa, kPa, mmHg | Double | Settings | Common | [] | Monnal T60
1328 | Apnea Tidal Volume Setting | TaemaA | Apnea Tidal Volume Setting | - | ['For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Oxylog 3000 plus, Carina, Evita V300. Tidal volume for apnea ventilation, Vtapn. Possible units: L, mL. Unit "L" is available for Evita Infinity V500, Babylog VN500, Savina 300, Oxylog 3000 plus, Carina, Evita V300. Unit "mL" is available for Evita Infinity V500, Babylog VN500, Evita V300. For Evita Infinity V500, Babylog VN500: Possible units: L, mL. VT Apnea VTapn.\xa0', 'Possible units: mL, L\xa0'] | ['Double\xa0', 'Double (Multiset)\xa0'] | Settings | ['Common\xa0', 'Evita Infinity V500 [MEDIBUS.X]\xa0', 'Evita Infinity V800\xa0'] | [] | ['Monnal T60', 'Evita V800', 'Evita V500']
1329 | Apnea Respiration Rate Setting | TaemaA | Apnea Respiration Rate Setting | ['1/min', '-'] | ['Possible units: 1/min\xa0', 'For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Oxylog 3000 plus, Carina, Evita V300. Respiratory rate for apnea ventilation, Rrapn.\xa0'] | Double | Settings | ['Common\xa0', 'Evita Infinity V500 [MEDIBUS.X]\xa0', 'Evita Infinity V800\xa0'] | [] | ['Monnal T60', 'Evita V800', 'Evita V500']
1360 | Nebulizer Activation Status | TaemaA
1370 | Device Low Battery Alarm | TaemaA | Battery Low Alarm | unitless | [] | Boolean | Alarms | Common | [] | Monnal T60
1371 | Apnea Alarm | TaemaA
1395 | Inspiratory Pressure Trigger Setting | TaemaA
1416 | Peak Inspiratory Flow | TaemaA
1417 | Peak Expiratory Flow | TaemaA
1520 | Inspired N2O | TaemaA | Inspired N2O | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | N2O | ['MP70', 'MX800', 'X3']
1538 | O2 Sensor Alarm | TaemaA | O2 Sensor Alarm | unitless | [] | Boolean | Alarms | Common | [] | Monnal T60
1647 | Expired Xenon | TaemaA
1648 | Airway Xenon Concentration | TaemaA
1649 | Maximum Support Pressure Setting | TaemaA
1657 | Expiratory Trigger Level Setting | TaemaA | Expiratory Trigger Level Setting | - | Possible units: % | Double | Settings | Common | [] | Monnal T60
1658 | Ventilation Boost Duration | TaemaA
1659 | Ventilation Boost Coefficient | TaemaA
1663 | Low PEEP Cycle Setting | TaemaA
1664 | Nebulizer Flow Setting | TaemaA
1666 | Inspired Xenon Setting | TaemaA
1667 | Sigh Coefficient Setting | TaemaA
1668 | Air Concentration Wave | TaemaA
1669 | Xenon Concentration Wave | TaemaA
1673 | Leakage of Pressure (Pleak) | TaemaA
1674 | Cumulated Inspiratory Time | TaemaA
1675 | Cumulated Expiratory Time | TaemaA
1677 | Cumulated Respiration Time With Low FiO2 | TaemaA
1678 | Cumulated Respiration Time With Low SaO2 | TaemaA
1679 | Cumulated Respiration Time With Leakage | TaemaA
1891 | Device Power Supply Status | TaemaA | ['Power Supply Status', 'Device Power Supply Status'] | unitless | ['Value is Sector, Internal Battery, or External Battery\xa0', 'The current status of the AC power system. \xa0', 'This variable is 24 bits long.'] | ['UShort\xa0', 'String\xa0'] | DeviceInfo | ['Common\xa0', 'Bed Data Server (Common)\xa0'] | BED | ['Monnal T60', 'Total Scale', 'Xenios Console']
1913 | Device Serial Number | TaemaA | Device Serial Number | unitless | ['Format: 8 characters string. This variable is buffered and sent repeatedly by the DDI with each data message.\xa0', 'The serial number of the bed communication controller.\xa00x000000000000 -\xa00xFFFFFFFFFFFF', 'ASCII string up to 24 characters long.\xa0'] | String | DeviceInfo | ['Common\xa0', 'Bed Data Server (Common)\xa0'] | ['BED\xa0', 'Treatment Status\xa0'] | ['Total Scale', 'Applix Smart', '5008', 'Monnal T60', 'Multifiltrate PRO']
2063 | Cumulated Time in Apnea | TaemaA
2064 | Cumulated Time in Hypopnea | TaemaA
2065 | Cumulated Time in Boost Ventilation | TaemaA
2066 | Low Inflation Airway FlowSetting | TaemaA
2067 | Low Inflation Tidal Volume | TaemaA
2068 | Low Inflation Inspiratory Pressure | TaemaA
2077 | Airway Gas Flow Rate | TaemaA | Airway Gas Flow Rate | - | Measured in the inspiratory path. | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | VENT: AWFin | ['MP70', 'MX800', 'X3']
2078 | Inspired Airway Pressure | TaemaA
2088 | SIMV Respiration Rate Setting | TaemaA | SIMV Respiration Rate Setting | - | Possible units: 1/min | Double | Settings | Common | [] | Monnal T60
2117 | High Airway Pressure Alarm | TaemaA
2120 | Battery Alarm | TaemaA | Battery Alarm | unitless | This alarm is raised if the internal battery is discharged. | Boolean | Alarms | Common | [] | Monnal T60
2351 | End Tidal O2 | TaemaA | End Tidal O2 | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | O2 | ['MP70', 'MX800', 'X3']
2359 | Inspired Agent Concentration | TaemaA
2360 | End Tidal Agent Concentration | TaemaA | End Tidal Agent Concentration | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | AGT | ['MP70', 'MX800', 'X3']
2381 | Breathing Circuit Type | TaemaA
2382 | Ventilator Operation Mode | TaemaA | Ventilator Operation Mode | unitless | This variable contains ventilation operation mode in string format, which is also send as a numerical value in variable 2758. See detailed description of variable 2758 for the list of supported values. For Evita 4, Evita 2 dura, Evita XL: Firmware version 04.00 or higher For Evita 2: Variable label is not correct. Please use variable 3853 instead. Obsolete variable that may be removed in future release. | String | DeviceInfo | ['Evita Infinity V500 [MEDIBUS.X]\xa0', 'Evita Infinity V800\xa0'] | [] | ['Evita V800', 'Evita V500']
2542 | Airway Pressure Wave | TaemaA | Airway Pressure Wave | ['mbar', '-'] | For Fabius GS, Fabius GS premium, Fabius Tiro, Fabius plus, Fabius plus XL, Fabius OS, Fabius MRI: 1mbar = 1cmH2O. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300, Fabius Family. For Evita Infinity V500, Babylog VN500: Airway pressure Paw. | Double | Waveforms | ['IntelliVue MX800\xa0', 'Evita Infinity V500 [MEDIBUS.X]\xa0', 'IntelliVue X3\xa0', 'Evita Infinity V800\xa0', 'IntelliVue MP70\xa0'] | Waves: AWP | ['X3', 'MX800', 'Evita V800', 'MP70', 'Evita V500']
2546 | Low Expired Minute Volume Alarm | TaemaA | Low Expired Minute Volume Alarm | unitless | [] | Boolean | Alarms | Common | [] | Monnal T60
2551 | End Tidal N2O | TaemaA
2573 | Ventilator Control Mode | TaemaA | Ventilator Control Mode | unitless | The variable indicates if the ventilator is in Non-Invasive Ventilation. ON means non-invasive ventilation and OFF means invasive ventilation. | String | DeviceInfo | Common | [] | Monnal T60
2582 | Target Respiration Rate | TaemaA | Target Respiration Rate | - | Possible units: 1/min | Double | Settings | Monnal T60 | [] | Monnal T60
2590 | Generic Device Alarm | TaemaA | Generic Device Alarm | unitless | ['Format: list of maximum 16 alarm IDs separated with a "¦".', 'This is a multiset variable containing the text messages for all active alarms. Associated with variable 6164 for the matching alarm codes.\xa0', 'This variable is 40 bits long.'] | ['ULong\xa0', 'String (Multiset)\xa0', 'String\xa0'] | Alarms | Common | [] | ['Monnal T60', '5008', 'Xenios Console']
2759 | Support Pressure Setting | TaemaA | Support Pressure Setting | - | Possible units: cmH2O, mbar, bar, hPa, kPa, mmHg | Double | Settings | Common | [] | Monnal T60
2776 | Total PEEP | TaemaA | Total PEEP | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | VENT: tPEEP | ['MP70', 'MX800', 'X3']
2777 | Auto PEEP | TaemaA
2879 | Device Software Revision | TaemaA | Device Software Revision | unitless | ['Format: vv.r.c vv = software version r = software revision c = software correction\xa0', 'Format: XY.ZZ\xa0', 'Release PiCCO software version, composed by one letter (type of protocol: V for usual protocol, D for demo version and X for unreleased version) and from one to\xa0three digits seperated by dots (version number). For example: D5.1 \xa0', 'ASCII string up to 24 characters long.\xa0', 'This variable is obsolete because it may contains incorrect information. Please use variables 764, 765 and 4201 must be used instead.\xa0', 'The software revision of the bed communication controller.\xa00.00 -- 9.99'] | String | DeviceInfo | ['Common\xa0', 'Bed Data Server (Common)\xa0'] | ['BED\xa0', 'Obsolete Variables\xa0', 'Device Information\xa0', 'Treatment Status\xa0'] | ['MP70', 'Total Scale', 'X3', 'Applix Smart', 'MX800', 'MP20', '5008', 'PiCCO2', 'Monnal T60', 'Multifiltrate PRO', 'Criticool']
2960 | Tidal Volume Target | TaemaA | Tidal Volume Target | - | Possible units: mL, L | Double | Settings | Common | [] | Monnal T60
2964 | I:E Ratio | TaemaA | I:E Ratio | unitless | This variable is associated with ChannelID identifier This variable is calculated from variable 1190 and sent as a string in the format "1:xx.xx". | ['Double\xa0', 'String\xa0'] | VitalSigns | ['IntelliVue X3\xa0', 'Common\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Calculated Variables | ['Monnal T60', 'MP70', 'MX800', 'X3']
2968 | Plateau Airway Pressure | TaemaA | Plateau Airway Pressure | - | This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. Possible units: cmH2O, mbar, bar, hPa, kPa, mmHg | Double | VitalSigns | ['IntelliVue X3\xa0', 'Common\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Pplat | ['Monnal T60', 'MP70', 'MX800', 'X3']
2976 | Patient Disconnect Alarm | TaemaA | Patient Disconnect Alarm | unitless | [] | Boolean | Alarms | Common | [] | Monnal T60
3470 | Pressure Trigger Actual Value | TaemaA
3474 | Expiratory Resistance | TaemaA | Expiratory Resistance | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Rexp | ['MP70', 'MX800', 'X3']
3476 | Ventilation Limit Pressure Measurement | TaemaA
3480 | Low Inflation Activation Status | TaemaA
3481 | Low Suflation Activation Status | TaemaA
3482 | Aspiration Activation Status | TaemaA
3483 | Automatic Ventilation Enable Status | TaemaA
3485 | Number of Inspiratory Cycle Trigged by Ventilator | TaemaA
3486 | Number of Expiratory Cycle Trigged by Ventilator | TaemaA
3487 | Number of Boost Ventilation | TaemaA
3488 | Expiratory Pause Activation Status | TaemaA
3489 | Inspiratory Pause Activation Status | TaemaA
3789 | Arterial O2 Saturation (SaO2) | TaemaA | Arterial O2 Saturation (SaO2) | ['%', '-'] | ['Obsolete variable that may be removed in a future release. The label of this variable is not correct. Please use variable 14 instead.\xa0', 'Arterial oxygenation saturation (input value for display of parameters DO2 and VO2). \xa0'] | Double | VitalSigns | ['Common\xa0', 'PiCCO2\xa0'] | ['ScvO2 Calibration\xa0', 'Obsolete Variables\xa0'] | ['X3', 'MX800', 'MP20', 'PiCCO2', 'MP70']
3901 | Airway Peak Flow | TaemaA
3922 | Expiratory Hold Status | TaemaA
3961 | Plateau Time Setting | TaemaA | Plateau Time Setting | - | ['This variable is obsolete that may be removed in a future release. The label of this variable is incorrect, please use variable 9585 instead.\xa0', 'Possible units: ms, sec, min, h\xa0'] | Double | Settings | ['IntelliVue X3\xa0', 'Monnal T60\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Obsolete Variables | ['Monnal T60', 'MP70', 'MX800', 'X3']
4004 | Breathing System Leakage | TaemaA | Breathing System Leakage | - | This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. Possible units: % | Double | OtherMeasuredData | ['IntelliVue X3\xa0', 'Common\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | VENT: Leak | ['Monnal T60', 'MP70', 'MX800', 'X3']
4022 | Inspired Xenon | TaemaA
4029 | Inspired Agent Setting | TaemaA | Inspired Agent Setting | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sInAGT | ['MP70', 'MX800', 'X3']
4367 | Respiration Rate Setting | TaemaA | Respiration Rate Setting | - | Possible units: 1/min | Double | Settings | Common | [] | Monnal T60
4371 | Sigh Tidal Volume Setting | TaemaA | Sigh Tidal Volume Setting | - | Possible units: mL, L | Double | Settings | ['IntelliVue X3\xa0', 'Common\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | VENT: sSghTV | ['Monnal T60', 'MP70', 'MX800', 'X3']
4383 | Static Lung Compliance | TaemaA | Static Lung Compliance | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Cstat | ['MP70', 'MX800', 'X3']
4384 | Static Lung Resistance | TaemaA | Static Lung Resistance | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Raw | ['MP70', 'MX800', 'X3']
4385 | Dynamic Lung Compliance | TaemaA | Dynamic Lung Compliance | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Cdyn | ['MP70', 'MX800', 'X3']
4386 | Dynamic Lung Resistance | TaemaA | Dynamic Lung Resistance | - | NLS_NOM_RES_AWAY_DYN | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Rdyn | ['MP70', 'MX800', 'X3']
4391 | Apnea Ventilation Alarm | TaemaA | Apnea Ventilation Alarm | unitless | [] | Boolean | Alarms | Common | [] | Monnal T60
4992 | Maximum Inspired Pressure Setting | TaemaA | Maximum Inspired Pressure Setting | - | Possible units: cmH2O, mbar, bar, hPa, kPa, mmHg | Double | Settings | Common | [] | Monnal T60
4997 | Continuous Positive Airway Pressure (CPAP) Setting | TaemaA | Continuous Positive Airway Pressure (CPAP) Setting | - | This is the device's measurement with label="NLS_NOM_SETT_CONC_AWAY_O2" and physio_id="NOM_CONC_AWAY_O2". | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sCPAP | ['MP70', 'MX800', 'X3']
5010 | Inspiratory Pressure Setting | TaemaA | Inspiratory Pressure Setting | - | Possible units: cmH2O, mbar, bar, hPa, kPa, mmHg | Double | Settings | Common | [] | Monnal T60
5029 | Pressure Ramp Setting (ASB Ramp) | TaemaA | Pressure Ramp Setting (ASB Ramp) | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
5178 | Number of Inspiratory Cycle Trigged by Patient | TaemaA
5179 | Number of Expiratory Cycle Trigged by Patient | TaemaA
5359 | Airway Peak Flow Setting | TaemaA | Airway Peak Flow Setting | - | Possible units: mL/min, mL/h, L/s, L/min, L/h | Double | Settings | Monnal T60 | [] | Monnal T60
5386 | Work of Breathing | TaemaA
5523 | Low Expiration PEEP Setting | TaemaA
5577 | Hypopnea Index | TaemaA
5618 | Number of SaO2 Limitation Reached | TaemaA
5619 | SaO2 Low Threshold | TaemaA
5712 | Low EtCO2 Alarm | TaemaA | Low EtCO2 Alarm | unitless | [] | Boolean | Alarms | Monnal T60 | [] | Monnal T60
5713 | High EtCO2 Alarm | TaemaA | High EtCO2 Alarm | unitless | [] | Boolean | Alarms | Monnal T60 | [] | Monnal T60
5738 | Inspired N2O Setting | TaemaA
5746 | Inspired 2nd Agent Concentration Setting | TaemaA
5946 | Waveform Shape Setting | TaemaA | Waveform Shape Setting | unitless | Consult detailed help for more information on this variable | String | Settings | Common | [] | Monnal T60
5947 | Airway Flow Rate Setting | TaemaA
5985 | Nurse Call Alarm | TaemaA
6164 | General Alarm Code | TaemaA | General Alarm Code | unitless | ['This is a multiset variable containing the alarm codes for all active alarms. Associated with variable 2590 for the matching alarm text messages.', 'Consult detailed help for more information on this variable\xa0', 'This variable is in multiset format. This variable contains alarm codes, variable 3878, alarm messages, variable 2596 contains alarm priority level. They need to be interpreted together. One alarm code may have one or more priority values and one or more alarm messages associated with it. The possible alarms are depended on device firmware version. Please see device user manuals for details. See variables 177 and 588 and 4131 for string format.'] | ['Byte (Multiset)\xa0', 'UShort (Multiset)\xa0', 'Long\xa0'] | Alarms | ['Common\xa0', 'Evita Infinity V500 [MEDIBUS.X]\xa0', 'Evita Infinity V800\xa0', 'Multifiltrate PRO\xa0'] | ['Alarm And Text Message\xa0', 'Dialysis Course\xa0'] | ['Monnal T60', 'Multifiltrate PRO', 'Evita V800', 'Evita V500']
6276 | NPPV Expiratory Trigger Level Setting | TaemaA
6397 | Expired 2nd Agent | TaemaA
6701 | Expiratory Time Setting | TaemaA | Expiratory Time Setting | s | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Zeus IE, Evita V300. | Double | Settings | ['Evita Infinity V500 [MEDIBUS.X]\xa0', 'Evita Infinity V800\xa0'] | [] | ['Evita V800', 'Evita V500']
6746 | Total Respiration Duration Setting | TaemaA | Total Respiration Duration Setting | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sBreaT | ['MP70', 'MX800', 'X3']
7927 | Inspired 2nd Agent Concentration | TaemaA
8379 | Peak Pressure | TaemaA | Peak Pressure | - | This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. Possible units: cmH2O, mbar, bar, hPa, kPa, mmHg | Double | VitalSigns | ['IntelliVue X3\xa0', 'Monnal T60\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['Monnal T60', 'MP70', 'MX800', 'X3']
8405 | Airway Volume Wave | TaemaA | Airway Volume Wave | - | [] | Double | Waveforms | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Waves: AWV | ['MP70', 'MX800', 'X3']
8548 | Identified 2nd Agent Name | TaemaA
8579 | Plimit Reached Alarm | TaemaA | Plimit Reached Alarm | unitless | [] | Boolean | Alarms | Monnal T60 | [] | Monnal T60
8589 | High Expired Tidal Volume Alarm | TaemaA | High Expired Tidal Volume Alarm | unitless | [] | Boolean | Alarms | Common | [] | Monnal T60
8596 | Low FiO2 Alarm | TaemaA | Low FiO2 Alarm | unitless | [] | Boolean | Alarms | Common | [] | Monnal T60
8597 | High FiO2 Alarm | TaemaA | High FiO2 Alarm | unitless | [] | Boolean | Alarms | Common | [] | Monnal T60
8598 | Low Respiration Rate Alarm | TaemaA | Low Respiration Rate Alarm | unitless | [] | Boolean | Alarms | Common | [] | Monnal T60
8599 | High Respiration Rate Alarm | TaemaA | High Respiration Rate Alarm | unitless | [] | Boolean | Alarms | Common | [] | Monnal T60
8600 | CO2 Apnea Alarm | TaemaA | CO2 Apnea Alarm | unitless | [] | Boolean | Alarms | Monnal T60 | [] | Monnal T60
8601 | High FiCO2 Alarm | TaemaA
8606 | Battery Failure Alarm | TaemaA | Battery Failure Alarm | unitless | [] | Boolean | Alarms | Common | [] | Monnal T60
9154 | Humidifier Activation Status | TaemaA | Humidifier Activation Status | - | [] | Double | DeviceInfo | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sHmCtl | ['MP70', 'MX800', 'X3']
9155 | Nebulizer Time Setting | TaemaA | Nebulizer Time Setting | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sNebTi | ['MP70', 'MX800', 'X3']
9293 | Pressure Rise Time Setting | TaemaA | Pressure Rise Time Setting | cmH2O/s | [] | Double | Settings | Common | [] | Monnal T60
9307 | Nebulizer Cycle Setting | TaemaA
9310 | Breathing Circuit Compliance | TaemaA
9353 | Low PEEP Setting | TaemaA
9355 | Low PEEP Time Setting | TaemaA
9358 | Number of Breath Cycles | TaemaA
9485 | Maximum Inspiratory Time Setting | TaemaA | Maximum Inspiratory Time Setting | ['s', '-'] | ['For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Carina, Evita V300. Maximal inspiratory time for flow cycled strokes, Timax. For Carina: Maximal inspiration time for CPAP, ASB.\xa0', 'Possible units: ms, sec, min, h\xa0'] | Double | Settings | ['IntelliVue MX800\xa0', 'Evita Infinity V500 [MEDIBUS.X]\xa0', 'IntelliVue X3\xa0', 'Evita Infinity V800\xa0', 'Common\xa0', 'IntelliVue MP70\xa0'] | sTimax | ['X3', 'MX800', 'Evita V800', 'Monnal T60', 'MP70', 'Evita V500']
9584 | High Expired Minute Volume Alarm | TaemaA | High Expired Minute Volume Alarm | unitless | [] | Boolean | Alarms | Common | [] | Monnal T60
4350 | FeedRate | FreseniusD | Current Delivery Rate | mL/h | If the pump is OFF, the return rate is 0. | UShort | Settings | Common | [] | Applix Smart
539 | Machine ID | BaxterC
850 | Disposable Brand Name | BaxterC
854 | Total Number of Used Disposable | BaxterC | Total Number of Used Disposable | - | [] | Double | DeviceInfo | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | nFilt | ['MP70', 'MX800', 'X3']
946 | Cumulated Syringe Infused Volume | BaxterC | Cumulated Syringe Infused Volume | - | [] | Double | OtherMeasuredData | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | SyrC | ['MP70', 'MX800', 'X3']
966 | Current Therapy Time | BaxterC | Current Therapy Time | ['s', '-'] | [] | ['Double\xa0', 'Long\xa0'] | OtherMeasuredData | ['IntelliVue X3\xa0', 'Common\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | FiltTi | ['MP70', '5008', 'MX800', 'X3']
1011 | Therapy Running Status | BaxterC | Therapy Running Status | unitless | ['Possible values:', 'Consult detailed help for more information on this variable\xa0'] | ['Long\xa0', 'String\xa0'] | DeviceInfo | Common | [] | ['5008', 'Artis']
1043 | Manually Entered Hematocrit | BaxterC | Manually Entered Hematocrit | ['%', '-'] | [] | ['UShort\xa0', 'Double\xa0'] | Settings | ['IntelliVue X3\xa0', 'Common\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | scHCTi | ['MP70', '5008', 'MX800', 'X3']
1246 | Hematocrit | BaxterC | Hematocrit | ['%', '-'] | [] | Double | VitalSigns | ['IntelliVue MX800\xa0', 'PiCCO2\xa0', 'IntelliVue X3\xa0', 'Common\xa0', 'IntelliVue MP70\xa0'] | ['ScvO2 Calibration\xa0', 'Hct\xa0'] | ['X3', 'MX800', '5008', 'PiCCO2', 'MP70']
1929 | Patient ID | BaxterC | Patient ID | unitless | ['The globally unique identifier for the Patient entity associated with this Bed entity or an empty string if no association currently exists. \xa0', '1–16 characters length alphanumeric string.\xa0', '9 digits maximum, only contain numbers from 0 to 9. \xa0'] | String | PatientInfo | ['Common\xa0', 'Bed Data Server (Common)\xa0', 'Expression IP5\xa0'] | ['Patient Info (EXPRESS Plus, Express Connect Only)\xa0', 'BED\xa0', 'Administrative Data\xa0', 'Patient Setting\xa0', 'Patient Demographics\xa0'] | ['MP70', 'Total Scale', 'X3', 'MX800', 'MP20', '5008', 'PiCCO2', 'Multifiltrate PRO', 'Expression IP5 et MR400']
2595 | General Alarm Status | BaxterC | General Alarm Status | unitless | Contains every alarm code that is currently active | Long (Multiset) | Alarms | Common | [] | V60
2898 | Patient Weight | BaxterC | Patient Weight | - | [] | Double | PatientInfo | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Weight | ['MP70', 'MX800', 'X3']
2997 | Blood Leak Detector Signal 1 | BaxterC
2998 | Blood Leak Detector Signal 2 | BaxterC
2999 | Blood Leak Detector Signal Difference | BaxterC
3000 | Blood Leak Detector Signal Average | BaxterC
3001 | Blood Leak Detector Normalization Value | BaxterC
3005 | Dialysate Flow Rate Setting | BaxterC | Dialysate Flow Rate Setting | ['mL/h', '-', 'mL/min'] | [] | ['Double\xa0', 'ULong\xa0', 'Long\xa0'] | Settings | ['IntelliVue X3\xa0', 'Common\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | ['Dialysis Settings\xa0', 'sDialF\xa0'] | ['Multifiltrate PRO', 'X3', 'MX800', '5008', 'MP70']
3006 | Fluid Removal Setting |  | Fluid Removal Volume Setting | mL | UF goal setting | Long | Settings | Common | Dialysis Settings | ['Multifiltrate PRO', '5008', 'Artis']
3007 | Effluent Flow Rate Setting | BaxterC | Effluent Flow Rate Setting | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sEfflF | ['MP70', 'MX800', 'X3']
3008 | Anticoagulant Delivery Method | BaxterC
3026 | Filter Pod Pressure | BaxterC | Filter Pod Pressure | ['-', 'mmHg'] | Filtrate pressure | ['Double\xa0', 'Long\xa0'] | OtherMeasuredData | ['IntelliVue X3\xa0', 'IntelliVue MX800\xa0', 'IntelliVue MP70\xa0', 'Multifiltrate PRO\xa0'] | ['FiltP\xa0', 'Dialysis Course\xa0'] | ['MP70', 'Multifiltrate PRO', 'MX800', 'X3']
3027 | Effluent Pod Pressure | BaxterC | Effluent Pod Pressure | - | [] | Double | OtherMeasuredData | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | EfflP | ['MP70', 'MX800', 'X3']
3047 | Next Intervention Time (hours) | BaxterC
3048 | Next Intervention Time (minutes) | BaxterC
3049 | Next Intervention Type | BaxterC
3050 | Therapy Type | BaxterC | Therapy Type | unitless | Consult detailed help for more information on this variable | String | DeviceInfo | Multifiltrate PRO | Treatment Status | Multifiltrate PRO
3168 | Fluid Removal Actual Rate | BaxterC | Fluid Removal Actual Rate | ['mL/h', '-'] | [] | ['Double\xa0', 'Long\xa0'] | VitalSigns | ['IntelliVue X3\xa0', 'Common\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Vrem/h | ['X3', 'Artis', 'MX800', '5008', 'MP70']
3248 | Generic Venous Line Pressure | BaxterC | Generic Venous Line Pressure | mmHg | This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. | Long | VitalSigns | ['Common\xa0', 'Multifiltrate PRO\xa0'] | Dialysis Course | ['Multifiltrate PRO', '5008', 'Artis']
3252 | Total Replacement Cumulated Volume | BaxterC | Total Replacement Cumulated Volume | L | Expected total substitution volume (pre + post) at the end of the treatment. | Double | VitalSigns | Common | [] | 5008
3270 | Total Replacement Dosage Cumulated Value | BaxterC
3274 | Filtration Fraction | BaxterC
3275 | Plasma Cumulated Volume | BaxterC
3317 | Effluent Dosage Cumulated Value | BaxterC
3323 | Filtration Dosage Cumulated Value | BaxterC
3325 | Pre-delivered Total Liquid Cumulated Volume | BaxterC
3337 | Pre-delivered Total Liquid Dosage Cumulated Value | BaxterC
3339 | Replacement Dosage Post-delivered Cumulated Value | BaxterC
3436 | Replacement Pre-delivered Percentage | BaxterC
3437 | Replacement Post-delivered Percentage | BaxterC | Replacement Post-delivered Percentage | % | [] | UShort | VitalSigns | Common | [] | 5008
4226 | Replacement Cumulated Volume | BaxterC | Replacement Cumulated Volume | ['L', 'mL', '-'] | Substituate-Total, total | ['Double\xa0', 'Long\xa0'] | VitalSigns | ['IntelliVue MX800\xa0', 'Multifiltrate PRO\xa0', 'IntelliVue X3\xa0', 'Common\xa0', 'IntelliVue MP70\xa0'] | ['Vrep/h\xa0', 'Dialysis Course\xa0'] | ['Multifiltrate PRO', 'X3', 'Artis', 'MX800', 'MP70']
4227 | Dialysate Cumulated Volume | BaxterC | Dialysate Cumulated Volume | - | This variable is associated with ChannelID identifier. | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | DialC/Vdly/h | ['MP70', 'MX800', 'X3']
4228 | Effluent Cumulated Volume | BaxterC | Effluent Cumulated Volume | - | This variable is associated with ChannelID identifier. | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Veff/h / EfflC | ['MP70', 'MX800', 'X3']
4315 | Force Applied to Syringe Pump | BaxterC
4347 | Therapy Remaning Time | BaxterC | Therapy Remaining Time | ['s', 'min'] | Remaining UF time including ISO UF time. | ['ULong\xa0', 'Long\xa0'] | OtherMeasuredData | Common | [] | ['5008', 'Artis']
4348 | Therapy Treatment Time | BaxterC
5855 | Blood Flow Rate | BaxterC | Blood Flow Rate | mL/min | This variable is associated with LowerAlarmLimit and UpperAlarmLimit attributes. | ['Long\xa0', 'Short\xa0'] | VitalSigns | Common | [] | ['Artis', 'Xenios Console']
5856 | Dialysate Flow Rate | BaxterC | Dialysate Flow Rate | mL/min | [] | ['ULong\xa0', 'Long\xa0'] | VitalSigns | Common | [] | ['5008', 'Artis']
5863 | Dialysate Temperature | BaxterC | Dialysate Temperature | °C | [] | Double | VitalSigns | Common | [] | ['5008', 'Artis']
5871 | Citrate Dose Setting | BaxterC | Citrate Dose Setting | 0.01 mmol/L | [] | Long | Settings | Multifiltrate PRO | Dialysis Settings | Multifiltrate PRO
5873 | Citrate Solution Concentration | BaxterC | Citrate Solution Concentration | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sCitCo | ['MP70', 'MX800', 'X3']
5874 | Calcium Solution Concentration | BaxterC | Calcium Solution Concentration | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sCConc | ['MP70', 'MX800', 'X3']
5885 | Replacement Solution Calcium Concentration | BaxterC
5886 | Citrate Acid Concentration | BaxterC | Citrate Acid Concentration | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sCSCAC | ['MP70', 'MX800', 'X3']
5887 | Citrate Load Setting | BaxterC | Citrate Load Setting | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sCitLo | ['MP70', 'MX800', 'X3']
845892 | Bicarbonate Conductivity | BaxterC
7438 | Generic Arterial Line Pressure | BaxterC | Generic Arterial Line Pressure | mmHg | This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. | Long | VitalSigns | ['Common\xa0', 'Multifiltrate PRO\xa0'] | Dialysis Course | ['Multifiltrate PRO', '5008', 'Artis']
7979 | Dialysate Conductivity | BaxterC | Dialysate Conductivity | mS/cm | [] | Double | VitalSigns | Common | [] | ['5008', 'Artis']
7986 | Calcium Flow Rate Setting | BaxterC | Calcium Flow Rate Setting | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sCSR | ['MP70', 'MX800', 'X3']
7987 | Calcium Compensation Setting | BaxterC | Calcium Compensation Setting | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sCComp | ['MP70', 'MX800', 'X3']
8252 | Syringe Force ADC Value | BaxterC
8792 | Fifth Pod Pressure | BaxterC
8793 | Pressure Drop | BaxterC | Pressure Drop | - | [] | Double | OtherMeasuredData | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | DP | ['MP70', 'MX800', 'X3']
8794 | Pre-blood Pump Rotation Rate | BaxterC
8795 | Replacement Pump Rotation Rate | BaxterC
8796 | Dialysate Pump Rotation Rate | BaxterC
8797 | Effluent Pump Rotation Rate | BaxterC
8798 | Blood Pump Rotation Rate | BaxterC
8799 | Weight on Pre-blood Scale | BaxterC
8800 | Weight on Replacement Scale | BaxterC
8801 | Weight on Dialysate Scale | BaxterC
8802 | Weight on Effluent Scale | BaxterC
8803 | Replacement Pre-delivered Rate Setting | BaxterC | Replacement Pre-delivered Rate Setting | ['-', 'mL/min'] | [] | ['UShort\xa0', 'Double\xa0'] | Settings | ['IntelliVue X3\xa0', 'Common\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sRPrFl | ['MP70', '5008', 'MX800', 'X3']
8804 | Replacement Post-delivered Rate Setting | BaxterC | Replacement Post-delivered Rate Setting | ['-', 'mL/min'] | [] | ['UShort\xa0', 'Double\xa0'] | Settings | ['IntelliVue X3\xa0', 'Common\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sRPoFl | ['MP70', '5008', 'MX800', 'X3']
8805 | Pre-blood Flow Rate Setting | BaxterC | Pre-blood Flow Rate Setting | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sPBPFl | ['MP70', 'MX800', 'X3']
8806 | Patient Weight Balance Setting | BaxterC | Patient Weight Balance Setting | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sWTbal | ['MP70', 'MX800', 'X3']
8807 | Anticoagulant Bolus Volume Setting | BaxterC | Anticoagulant Bolus Volume Setting | mL | [] | Double | Settings | Common | [] | 5008
8808 | Anticoagulant Bolus Interval Setting | BaxterC | Anticoagulant Bolus Interval Setting | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sBolIn | ['MP70', 'MX800', 'X3']
8809 | Anticoagulant Infusion Rate Setting | BaxterC | Anticoagulant Infusion Rate Setting | ['mL/h', '-', '0.1 mL/h'] | Heparin continuous rate | ['Double\xa0', 'Long\xa0'] | Settings | ['IntelliVue X3\xa0', 'Common\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | ['sHepRt\xa0', 'Dialysis Settings\xa0'] | ['Multifiltrate PRO', 'X3', 'MX800', '5008', 'MP70']
8810 | Ultrafiltration Ratio Setting | BaxterC | Ultrafiltration Ratio Setting | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sUFR | ['MP70', 'MX800', 'X3']
8811 | Therapy Elapsed Time | BaxterC | Therapy Elapsed Time | ['-', 'min'] | [] | ['Double\xa0', 'Long\xa0'] | OtherMeasuredData | ['IntelliVue X3\xa0', 'Common\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | RunTi | ['MP70', 'Artis', 'MX800', 'X3']
8813 | Processed Blood Volume | BaxterC | Processed Cumulated Blood Volume | ['L', '-'] | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'Common\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | BVolPr | ['X3', 'Artis', 'MX800', '5008', 'MP70']
8814 | Replacement Pre-delivered Cumulated Volume | BaxterC | Replacement Pre-delivered Cumulated Volume | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | RpreC | ['MP70', 'MX800', 'X3']
8815 | Replacement Post-delivered Cumulated Volume | BaxterC | Replacement Post-delivered Cumulated Volume | ['mL', '-'] | This variable is associated with ChannelID identifier. | ['Double\xa0', 'Long\xa0'] | VitalSigns | ['IntelliVue X3\xa0', 'Common\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Vdpo/h / RpostC | ['MP70', '5008', 'MX800', 'X3']
8816 | Pre-blood Cumulated Volume | BaxterC | Pre-blood Cumulated Volume | - | This variable is associated with ChannelID identifier. | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | VPBP/h / PreBC | ['MP70', 'MX800', 'X3']
8817 | Patient Weight Cumulated Value | BaxterC | Patient Weight Cumulated Value | g | Absolute balancing error | Long | VitalSigns | Multifiltrate PRO | Dialysis Course | Multifiltrate PRO
8818 | Pre-blood Weight ADC Value | BaxterC
8819 | Replacement Weight ADC Value | BaxterC
8820 | Dialysate Weight ADC Value | BaxterC
8821 | Effluent Weight ADC Value | BaxterC
8822 | Access Pressure ADC Value | BaxterC
8823 | Filter Pressure ADC Value | BaxterC
8824 | Return Pressure ADC Value | BaxterC
8825 | Effluent Pressure ADC Value | BaxterC
8826 | Fifth Pod Pressure ADC Value | BaxterC
8827 | Blood Pump Counts | BaxterC
8828 | PBP Pump Counts | BaxterC
8829 | Replacement Pump Counts | BaxterC
8830 | Dialysate Pump Counts | BaxterC
8831 | Effluent Pump Counts | BaxterC
8832 | PWM Normalization Value | BaxterC
8833 | PWM Actual Value | BaxterC
8834 | Encoder Pulses of Syringe Pump | BaxterC
8835 | Syringe Speed | BaxterC | Syringe Speed | - | [] | Double | OtherMeasuredData | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Vsyr/h | ['MP70', 'MX800', 'X3']
8836 | Syringe Pump Overload Threshold | BaxterC
8837 | Arps Reference 1 | BaxterC
8838 | Arps Reference 2 | BaxterC
8839 | Arps Reference 3 | BaxterC
8840 | Patient Sensor Value 1 | BaxterC | Patient Sensor Value 1 | g | Weight of scale 1 | Long | OtherMeasuredData | Multifiltrate PRO | Dialysis Course | Multifiltrate PRO
8841 | Patient Sensor Value 2 | BaxterC | Patient Sensor Value 2 | g | Weight of scale 2 | Long | OtherMeasuredData | Multifiltrate PRO | Dialysis Course | Multifiltrate PRO
8842 | Patient Sensor Value 3 | BaxterC | Patient Sensor Value 3 | g | Weight of scale 3 | Long | OtherMeasuredData | Multifiltrate PRO | Dialysis Course | Multifiltrate PRO
8843 | Patient Sensor Value 4 | BaxterC | Patient Sensor Value 4 | g | Weight of scale 4 | Long | OtherMeasuredData | Multifiltrate PRO | Dialysis Course | Multifiltrate PRO
8844 | Voltage AD Value 1 | BaxterC
8845 | Voltage AD Value 2 | BaxterC
8846 | Voltage AD Value 3 | BaxterC
8847 | Voltage AD Value 4 | BaxterC
8848 | Voltage AD Value 5 | BaxterC
8849 | Syringe Pump Slope | BaxterC
8850 | Syringe Pump Load Threshold | BaxterC
8851 | Maximum Volume Gain Setting | BaxterC | Maximum Volume Gain Setting | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sMaxG | ['MP70', 'MX800', 'X3']
8852 | Excess Patient Fluid Limit | BaxterC
8853 | Patient Fluid Balance | BaxterC | Patient Fluid Balance | ['mL', '-'] | Balance, total | ['Double\xa0', 'Long\xa0'] | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MX800\xa0', 'IntelliVue MP70\xa0', 'Multifiltrate PRO\xa0'] | ['Wtcum\xa0', 'Dialysis Course\xa0'] | ['MP70', 'Multifiltrate PRO', 'MX800', 'X3']
8854 | PBP Method | BaxterC
8855 | Syringe Pump Overload Type | BaxterC
8856 | Therapy Operation Status | BaxterC | Therapy Operation Status | unitless | Consult detailed help for more information on this variable | ['Short\xa0', 'Long\xa0', 'String\xa0'] | DeviceInfo | Common | Treatment Status | ['Multifiltrate PRO', '5008', 'Artis']
9638 | Filtration Cumulated Volume | BaxterC
3636 | Alarmstatus | FreseniusD | Device Alarm Status | unitless | ['The alarms code are output as decimal values.\xa0Consult detailed help for more information on this variable\xa0', 'This variable is 8 bits long.'] | ['UShort\xa0', 'UByte\xa0'] | Alarms | Common | [] | ['Applix Smart', 'Xenios Console']
414 | Cardiac Output | PulsionA | Cardiac Output | ['L/min', '-'] | This variable is available only if patient category 'adult' or 'paediatric' is selected. | Double | VitalSigns | ['IntelliVue X3\xa0', 'Common\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | ['Discontinuous Mean Parameter\xa0', 'C.O.\xa0'] | ['X3', 'CardioQ', 'MX800', 'PiCCO2', 'MP70']
519 | Central Venous Pressure (CVP) | PulsionA | Central Venous Pressure (CVP) | ['mmHg', '-'] | ['Available on IntelliVue devices with firmware revision G or later.\xa0', 'This variable transmits either a continuous measurement or a manual input value. \xa0'] | ['Double\xa0', 'Short\xa0'] | VitalSigns | Common | ['Continuous Parameter\xa0', 'IBP: CVP\xa0'] | ['X3', 'MX800', 'MP20', 'PiCCO2', 'MP70']
585 | Patient Category | PulsionA | Patient Category | unitless | ['A: adult C: paediatric L: linear\xa0', 'Consult detailed help for more information on this variable\xa0'] | ['Short\xa0', 'String\xa0'] | PatientInfo | ['Common\xa0', 'PiCCO2\xa0'] | Patient Setting | ['CardioQ', 'PiCCO2']
649 | Cardiac Index | PulsionA | ['Cardiac Index', 'Cardiac Output Index'] | ['L/min/m2', '-'] | This variable is available only if patient category 'adult' or 'paediatric' is selected. | Double | VitalSigns | ['IntelliVue MX800\xa0', 'PiCCO2\xa0', 'IntelliVue X3\xa0', 'Common\xa0', 'IntelliVue MP70\xa0'] | ['Discontinuous Mean Parameter\xa0', 'C.I.\xa0'] | ['X3', 'CardioQ', 'MX800', 'PiCCO2', 'MP70']
650 | Stroke Volume | PulsionA | Stroke Volume | ['mL', '-'] | ['Beat to beat stroke volume. \xa0', "This variable is available only if patient category 'adult' or 'paediatric' is selected.\xa0"] | ['Double\xa0', 'Long\xa0'] | VitalSigns | ['IntelliVue X3\xa0', 'Common\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | ['Continuous Parameter\xa0', 'SV\xa0'] | ['X3', 'CardioQ', 'Artis', 'MX800', 'PiCCO2', 'MP70']
654 | Systemic Vascular Resistance | PulsionA | Systemic Vascular Resistance | ['dyn·s/cm5', '-'] | [] | ['Double\xa0', 'Short\xa0'] | VitalSigns | ['IntelliVue X3\xa0', 'Common\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | ['Continuous Parameter\xa0', 'SVR\xa0'] | ['MP70', 'PiCCO2', 'MX800', 'X3']
768 | Generic Error Code | PulsionA | Generic Error Code | unitless | TD measurement error, "0" is transmitted if no error occurs. | Short | Alarms | Common | Error | PiCCO2
1434 | Patient Body Surface Area | PulsionA | Patient Body Surface Area | m2 | [] | Double | PatientInfo | PiCCO2 | Patient Calculation | PiCCO2
1435 | Systemic Vascular Resistance Index | PulsionA | Systemic Vascular Resistance Index | ['dyn·s·m2/cm5', '-'] | [] | ['Double\xa0', 'Short\xa0'] | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'PiCCO2\xa0', 'IntelliVue MX800\xa0'] | ['SVRI\xa0', 'Continuous Parameter\xa0'] | ['MP70', 'PiCCO2', 'MX800', 'X3']
1468 | O2 Consumption (VO2) | PulsionA | O2 Consumption (VO2) | mL/min | [] | Short | VitalSigns | PiCCO2 | Continuous Parameter | PiCCO2
1472 | O2 Consumption Index | PulsionA | ['O2 Consumption Index', 'O2 Consumption Index (VO2I)'] | ['mL/min/m2', '-'] | [] | ['Double\xa0', 'Short\xa0'] | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'PiCCO2\xa0', 'IntelliVue MX800\xa0'] | ['Continuous Parameter\xa0', 'VO2I\xa0'] | ['MP70', 'PiCCO2', 'MX800', 'X3']
1710 | Cardiac Power Index (CCO) | PulsionA | Cardiac Power Index (CCO) | ['Not Available', '-'] | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'PiCCO2\xa0', 'IntelliVue MX800\xa0'] | ['Continuous Parameter\xa0', 'CPI\xa0'] | ['MP70', 'PiCCO2', 'MX800', 'X3']
1711 | Cardiac Power Output (CCO) | PulsionA | Cardiac Power Output (CCO) | ['W', '-'] | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'PiCCO2\xa0', 'IntelliVue MX800\xa0'] | ['Continuous Parameter\xa0', 'CPO\xa0'] | ['MP70', 'PiCCO2', 'MX800', 'X3']
1720 | ICG Measurement Counter | PulsionA | ICG Measurement Counter | unitless | Counter for signalizing update of ICG measurement. | Short | OtherMeasuredData | PiCCO2 | Discontinuous ICG | PiCCO2
1721 | Plasma Disappearance Rate | PulsionA | Plasma Disappearance Rate | Not Available | [] | Double | VitalSigns | PiCCO2 | Discontinuous ICG | PiCCO2
1722 | ICG Retention Rate | PulsionA | ICG Retention Rate | % | [] | Double | VitalSigns | PiCCO2 | Discontinuous ICG | PiCCO2
1730 | Cardiac Index ((Thermodilution) | PulsionA | Cardiac Index ((Thermodilution) | L/min/m2 | [] | Double | VitalSigns | PiCCO2 | Discontinuous TD | PiCCO2
1731 | Cardiac Output (Thermodilution) | PulsionA | Cardiac Output (Thermodilution) | L/min | [] | Double | VitalSigns | PiCCO2 | Discontinuous TD | PiCCO2
1732 | Global Ejection Fraction (Thermodilution) | PulsionA | Global Ejection Fraction (Thermodilution) | % | [] | Short | VitalSigns | PiCCO2 | Discontinuous TD | PiCCO2
1733 | Global End-Diastolic Volum (Thermodilution) | PulsionA | ['Global End-Diastolic Volum (Thermodilution)', 'Global End-Diastolic Volume (Thermodilution)'] | ['mL', '-'] | [] | ['Double\xa0', 'Short\xa0'] | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'PiCCO2\xa0', 'IntelliVue MX800\xa0'] | ['iGEDV\xa0', 'Discontinuous TD\xa0'] | ['MP70', 'PiCCO2', 'MX800', 'X3']
1734 | Global End-Diastolic Volum Index (Thermodilution) | PulsionA | Global End-Diastolic Volum Index (Thermodilution) | ['mL/m2', '-'] | [] | ['Double\xa0', 'Short\xa0'] | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'PiCCO2\xa0', 'IntelliVue MX800\xa0'] | ['Discontinuous TD\xa0', 'iGEDVI\xa0'] | ['MP70', 'PiCCO2', 'MX800', 'X3']
1735 | Pulmonary Vascular Permeability Index (Thermodilution) | PulsionA | Pulmonary Vascular Permeability Index (Thermodilution) | ['unitless', '-'] | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'PiCCO2\xa0', 'IntelliVue MX800\xa0'] | ['iPVPI\xa0', 'Discontinuous TD\xa0'] | ['MP70', 'PiCCO2', 'MX800', 'X3']
1821 | Generic ICG Error | PulsionA | Generic ICG Error | unitless | ICG measurement error, "0" is transmitted if no error occurs. | Short | Alarms | PiCCO2 | Error | PiCCO2
1930 | Patient Name | PulsionA | Patient Name | unitless | ["Allowed characters include A..Z, 0..9, ' ', '.', and '-'. The device sends a single character each 400 ms, so it takes time to transmit the full patient name. \xa0", 'Easily recognizable patient name that can be used by authorized personnel to locate the specific physical patient (e.g. j_doe…). \xa0'] | String | PatientInfo | ['Common\xa0', 'Bed Data Server (Common)\xa0', 'PiCCO2\xa0'] | ['Patient Setting\xa0', 'PATIENT\xa0', 'Patient Demographics\xa0'] | ['Total Scale', 'X3', 'MX800', 'MP20', 'PiCCO2', 'MP70']
1943 | Predicted Patient Body Surface Area | PulsionA | Predicted Patient Body Surface Area | m2 | [] | Double | PatientInfo | PiCCO2 | Patient Calculation | PiCCO2
1944 | Predicted Patient Weight | PulsionA | Predicted Patient Weight | kg | [] | Double | PatientInfo | PiCCO2 | Patient Calculation | PiCCO2
2255 | Device Date/Time | PulsionA | Device Date/Time | unitless | ['Format: YYYYMMDDhhmmss, where ss is always present but might be 00.\xa0', 'Format: 16 ascii digits string of "YYYYMMDDhhmmssff", where the "ff" is second_fractions that is not used by IntelliVue monitors.\xa0', 'Format: YYYY-MM-DD HH:MM:SS\xa0', 'Format: YYYYMMDDHHMMSS\xa0', 'This variable contains the current time\xa0of the device\xa0in the form of\xa0"MM/dd/yyyy - hh:mm:ss". It is also sent in the "DeviceTime" attribute of other variables. \xa0'] | String | DeviceInfo | ['Common\xa0', 'Evita Infinity V500 [MEDIBUS.X]\xa0', 'Evita Infinity V800\xa0', 'PiCCO2\xa0'] | ['System Information\xa0', 'Device Information\xa0', 'Device Info\xa0'] | ['X3', 'CardioQ', 'MX800', 'MP20', 'Evita V800', 'PiCCO2', 'MP70', 'Expression IP5 et MR400', 'Evita V500', 'Criticool']
2259 | Arterial BP Systolic | PulsionA | Arterial BP Systolic | ['mmHg', '-'] | [] | Double | VitalSigns | Common | ['Continuous Parameter\xa0', 'IBP: ABP\xa0'] | ['X3', 'MX800', 'MP20', 'PiCCO2', 'MP70']
2260 | Arterial BP Diastolic | PulsionA | Arterial BP Diastolic | ['mmHg', '-'] | [] | Double | VitalSigns | Common | ['Continuous Parameter\xa0', 'IBP: ABP\xa0'] | ['X3', 'MX800', 'MP20', 'PiCCO2', 'MP70']
2261 | Arterial BP Mean | PulsionA | Arterial BP Mean | ['mmHg', '-'] | [] | Double | VitalSigns | Common | ['Continuous Parameter\xa0', 'IBP: ABP\xa0'] | ['X3', 'MX800', 'MP20', 'PiCCO2', 'MP70']
2341 | Patient Blood Temperature | PulsionA | Patient Blood Temperature | ['°C', '-'] | A offset of 30°C is already taken into account by the DDI and consequently, the user receives the blood temperature real value, without having to interpret it. | Double | VitalSigns | ['IntelliVue X3\xa0', 'Common\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | ['Discontinuous TD\xa0', 'Tblood\xa0'] | ['MP70', 'PiCCO2', 'MX800', 'X3']
2342 | Injectate Temperature | PulsionA | Injectate Temperature | ['°C', '-'] | A offset of 15°C is already taken into account by the DDI and consequently, the user receives the injectate temperature real value, without having to interpret it. | Double | OtherMeasuredData | ['IntelliVue X3\xa0', 'Common\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | ['Tinj\xa0', 'Discontinuous TD\xa0'] | ['MP70', 'PiCCO2', 'MX800', 'X3']
2452 | O2 Delivery (DO2) | PulsionA | O2 Delivery (DO2) | mL/min | [] | Short | VitalSigns | PiCCO2 | Continuous Parameter | PiCCO2
2462 | SvO2 Alarm | PulsionA | SvO2 Alarm | unitless | High/Low Alarm of ScvO2 Consult detailed help for more information on this variable | Short | Alarms | PiCCO2 | Alarm | PiCCO2
2819 | SvO2 Calibration Status | PulsionA | SvO2 Calibration Status | unitless | ScvO2Cal: Point of time of calibration. "1" once during “Bloodsample drawn”, otherwise "0". | Boolean | DeviceInfo | PiCCO2 | ScvO2 Calibration | PiCCO2
2915 | Ejection Fraction | PulsionA | Ejection Fraction | ['%', '-'] | Global Ejection Fraction | ['Double\xa0', 'Short\xa0'] | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'PiCCO2\xa0', 'IntelliVue MX800\xa0'] | ['REF\xa0', 'Discontinuous Mean Parameter\xa0'] | ['MP70', 'PiCCO2', 'MX800', 'X3']
3184 | Alternate Patient ID | PulsionA | Alternate Patient ID | unitless | Alternate display of Patient ID. Allowed characters include A..Z, 0..9, ' ', '.', and '-'. The device sends a single character each 400 ms, so it takes time to transmit the full patient ID. | String | PatientInfo | PiCCO2 | Patient Setting | PiCCO2
3877 | Device Protocol Revision | PulsionA | Device Protocol Revision | unitless | Release PiCCO2 software version, composed by one letter (type of protocol: V for usual protocol, D for demo version and X for unreleased version) and from one to three digits seperated by dots (version number). For example: D5.1 | String | DeviceInfo | PiCCO2 | Device Information | PiCCO2
4532 | Stroke Volume Variation | PulsionA | Stroke Volume Variation | ['%', '-'] | [] | ['Double\xa0', 'Short\xa0'] | VitalSigns | ['IntelliVue X3\xa0', 'Common\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | ['Continuous Parameter\xa0', 'SVV\xa0'] | ['MP70', 'PiCCO2', 'MX800', 'X3']
4575 | O2 Delivery Index (DO2I) | PulsionA | O2 Delivery Index (DO2I) | mL/min/m2 | [] | Short | VitalSigns | PiCCO2 | Continuous Parameter | PiCCO2
4865 | Cardiac Function Index (Thermodilution) | PulsionA | Cardiac Function Index (Thermodilution) | ['1/min', '-'] | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'PiCCO2\xa0', 'IntelliVue MX800\xa0'] | ['Discontinuous TD\xa0', 'iCFI\xa0'] | ['MP70', 'PiCCO2', 'MX800', 'X3']
4866 | Intrathoracic Blood Volume (Thermodilution) | PulsionA | Intrathoracic Blood Volume (Thermodilution) | ['mL', '-'] | [] | ['Double\xa0', 'Short\xa0'] | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'PiCCO2\xa0', 'IntelliVue MX800\xa0'] | ['iITVB\xa0', 'Discontinuous TD\xa0'] | ['MP70', 'PiCCO2', 'MX800', 'X3']
4867 | Intrathoracic Blood Volume Index (Thermodilution) | PulsionA | Intrathoracic Blood Volume Index (Thermodilution) | ['mL/m2', '-'] | [] | ['Double\xa0', 'Short\xa0'] | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'PiCCO2\xa0', 'IntelliVue MX800\xa0'] | ['Discontinuous TD\xa0', 'iITBVI\xa0'] | ['MP70', 'PiCCO2', 'MX800', 'X3']
4868 | Extravascular Lung Water (Thermodilution) | PulsionA | Extravascular Lung Water (Thermodilution) | ['mL', '-'] | [] | ['Double\xa0', 'Short\xa0'] | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'PiCCO2\xa0', 'IntelliVue MX800\xa0'] | ['Discontinuous TD\xa0', 'iEVLW\xa0'] | ['MP70', 'PiCCO2', 'MX800', 'X3']
4869 | Extravascular Lung Water Index (Thermodilution) | PulsionA | Extravascular Lung Water Index (Thermodilution) | ['mL/kg', '-'] | [] | ['Double\xa0', 'Short\xa0'] | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'PiCCO2\xa0', 'IntelliVue MX800\xa0'] | ['iEVLWI\xa0', 'Discontinuous TD\xa0'] | ['MP70', 'PiCCO2', 'MX800', 'X3']
5072 | Intrathoracic Blood Volume (ITBV) | PulsionA | Intrathoracic Blood Volume (ITBV) | ['mL', '-'] | [] | ['Double\xa0', 'Short\xa0'] | VitalSigns | ['IntelliVue X3\xa0', 'Common\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | ['Discontinuous Mean Parameter\xa0', 'ITBV\xa0'] | ['MP70', 'PiCCO2', 'MX800', 'X3']
5073 | Intrathoracic Blood Volume Index (ITBI) | PulsionA | Intrathoracic Blood Volume Index (ITBI) | ['mL/m2', '-'] | [] | ['Double\xa0', 'Short\xa0'] | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'PiCCO2\xa0', 'IntelliVue MX800\xa0'] | ['ITBVI\xa0', 'Discontinuous Mean Parameter\xa0'] | ['MP70', 'PiCCO2', 'MX800', 'X3']
5074 | Extravascular Lung Water | PulsionA | Extravascular Lung Water | ['mL', '-'] | [] | ['Double\xa0', 'Short\xa0'] | VitalSigns | ['IntelliVue X3\xa0', 'Common\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | ['Discontinuous Mean Parameter\xa0', 'EVLW\xa0'] | ['MP70', 'PiCCO2', 'MX800', 'X3']
5075 | Extravascular Lung Water Index | PulsionA | Extravascular Lung Water Index | ['mL/kg', '-'] | [] | ['Double\xa0', 'Short\xa0'] | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'PiCCO2\xa0', 'IntelliVue MX800\xa0'] | ['Discontinuous Mean Parameter\xa0', 'EVLWI\xa0'] | ['MP70', 'PiCCO2', 'MX800', 'X3']
5076 | Global End Diastolic Volume | PulsionA | Global End Diastolic Volume | ['mL', '-'] | [] | ['Double\xa0', 'Short\xa0'] | VitalSigns | ['IntelliVue X3\xa0', 'Common\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | ['Discontinuous Mean Parameter\xa0', 'GEDV\xa0'] | ['MP70', 'PiCCO2', 'MX800', 'X3']
5077 | Global End Diastolic Volume Index | PulsionA | Global End Diastolic Volume Index | ['mL/m2', '-'] | [] | ['Double\xa0', 'Short\xa0'] | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'PiCCO2\xa0', 'IntelliVue MX800\xa0'] | ['Discontinuous Mean Parameter\xa0', 'GEDVI\xa0'] | ['MP70', 'PiCCO2', 'MX800', 'X3']
5078 | Cardiac Function Index | PulsionA | Cardiac Function Index | ['1/min', '-'] | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'PiCCO2\xa0', 'IntelliVue MX800\xa0'] | ['Discontinuous Mean Parameter\xa0', 'CFI\xa0'] | ['MP70', 'PiCCO2', 'MX800', 'X3']
5642 | SvO2 Calibration Data | PulsionA | SvO2 Calibration Data | unitless | ScvO2Cal: ScvO2 Calibration input. | Double | DeviceInfo | PiCCO2 | ScvO2 Calibration | PiCCO2
5684 | Continuous Cardiac Output (CCO) Alarm | PulsionA | Continuous Cardiac Output (CCO) Alarm | unitless | High/Low Alarm of PCCO Consult detailed help for more information on this variable | Short | Alarms | PiCCO2 | Alarm | PiCCO2
6403 | Pulse Contour Cardiac Output | PulsionA | Pulse Contour Cardiac Output | L/min | [] | Double | VitalSigns | Common | Continuous Parameter | PiCCO2
6404 | Thermodilution Measurement Status | PulsionA | Thermodilution Measurement Status | unitless | Consult detailed help for more information on this variable | Short | OtherMeasuredData | Common | Discontinuous TD | PiCCO2
6407 | Index Of Left Ventricular Contractility (dPmx) | PulsionA | Index Of Left Ventricular Contractility (dPmx) | mmHg/s | [] | Short | VitalSigns | Common | Continuous Parameter | PiCCO2
6410 | Mean Transit Time (MTt) | PulsionA | Mean Transit Time (MTt) | ['s', '-'] | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'Common\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | ['Discontinuous TD\xa0', 'Available only in revF or earlier\xa0'] | ['MP70', 'PiCCO2', 'MX800', 'X3']
6411 | Down Slope Time | PulsionA | Down Slope Time | s | [] | Double | VitalSigns | Common | Discontinuous TD | PiCCO2
6422 | Temperature Difference (Thermodilution) | PulsionA | Temperature Difference (Thermodilution) | °C | [] | Double | VitalSigns | PiCCO2 | Discontinuous TD | PiCCO2
6424 | Injectate Temperature (Thermodilution) | PulsionA | Injectate Temperature (Thermodilution) | °C | [] | Double | OtherMeasuredData | PiCCO2 | Discontinuous TD | PiCCO2
6428 | PICCO Catheter | PulsionA | PICCO Catheter | unitless | Range 0-14 Consult detailed help for more information on this variable | Short | DeviceInfo | Common | Device Information | PiCCO2
6434 | Pulse Pressure Variation | PulsionA | Pulse Pressure Variation | % | [] | Short | VitalSigns | Common | Continuous Parameter | PiCCO2
6493 | Pulmonary Vascular Permeability Index | PulsionA | Pulmonary Vascular Permeability Index | ['unitless', '-'] | This is the device's measurement with label="NLS_NOM_PERM_VASC_PULM_INDEX" and physio_id="NOM_PERM_VASC_PULM_INDEX". See also variable 8791. | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'PiCCO2\xa0', 'IntelliVue MX800\xa0'] | ['Discontinuous Mean Parameter\xa0', 'PVPI\xa0'] | ['MP70', 'PiCCO2', 'MX800', 'X3']
6533 | Patient Gender | PulsionA | Patient Gender | unitless | Consult detailed help for more information on this variable | Short | PatientInfo | PiCCO2 | Patient Setting | PiCCO2
6689 | Pulse Contour Cardiac Index | PulsionA | Pulse Contour Cardiac Index | L/min/m2 | [] | Double | VitalSigns | PiCCO2 | Continuous Parameter | PiCCO2
6690 | Thermodilution Measurement Counter | PulsionA | Thermodilution Measurement Counter | unitless | Counter for signalizing when all discontinuous thermodilution parameters have been updated. | Short | OtherMeasuredData | PiCCO2 | Discontinuous TD | PiCCO2
6692 | Right to Left Heart Shunt Fraction | PulsionA | Right to Left Heart Shunt Fraction | % | [] | Short | VitalSigns | PiCCO2 | Discontinuous TD | PiCCO2
6694 | Arterial Blood Pressure Wave | PulsionA | Arterial Blood Pressure Wave | mmHg | [] | Double | Waveforms | Common | Continuous Parameter | PiCCO2
6695 | Injectate Volume Setting | PulsionA | Injectate Volume Setting | mL | [] | Double | Settings | Common | Discontinuous TD | PiCCO2
6696 | Stroke Volume Min | PulsionA | Stroke Volume Min | mL | [] | Double | VitalSigns | Common | Continuous Parameter | PiCCO2
6697 | Stroke Volume Max | PulsionA | Stroke Volume Max | mL | [] | Double | VitalSigns | Common | Continuous Parameter | PiCCO2
6698 | Stroke Volume Index | PulsionA | Stroke Volume Index | mL/m2 | [] | Double | VitalSigns | PiCCO2 | Continuous Parameter | PiCCO2
6699 | Stroke Volume Mean | PulsionA | Stroke Volume Mean | mL | [] | Double | VitalSigns | Common | Continuous Parameter | PiCCO2
7823 | Central Venous O2 Saturation (ScvO2) | PulsionA | Central Venous O2 Saturation (ScvO2) | ['%', '-'] | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'PiCCO2\xa0', 'IntelliVue MX800\xa0'] | ['ScvO2\xa0', 'Continuous Parameter\xa0'] | ['MP70', 'PiCCO2', 'MX800', 'X3']
7941 | SpO2 Alarm | PulsionA | SpO2 Alarm | unitless | High/Low Alarm of SpO2 Consult detailed help for more information on this variable | Short | Alarms | PiCCO2 | Alarm | PiCCO2
2907 | Patient Skin Temperature | MTREA | Patient Temperature | ['°C', '-'] | ['This variable is associated with LowerAlarmLimit and UpperAlarmLimit attributes. This variable is associated with ChannelID identifier.\xa0', 'This variable is associated with SourceID identifier.\xa0', 'Possible units: °C or °F.\xa0'] | Double | VitalSigns | Common | Temp | ['X3', 'MX800', 'MP20', 'Xenios Console', 'MP70', 'Expression IP5 et MR400', 'Criticool']
3291 | Patient Temperature Setting | MTREA | Patient Temperature Setting | ['°C', '-'] | [] | Double | Settings | ['IntelliVue X3\xa0', 'Common\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sTskin | ['MP70', 'X3', 'MX800', 'Criticool']
651 | Stroke Volume Index | DeltexA | Stroke Volume Index | ['mL/m2', '-'] | This variable is available only if patient category 'adult' or 'paediatric' is selected. | Double | VitalSigns | ['IntelliVue X3\xa0', 'Common\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | SI | ['CardioQ', 'MP70', 'MX800', 'X3']
6412 | Blood Minute Distance | DeltexA | Blood Minute Distance | cm | [] | Double | VitalSigns | Common | [] | CardioQ
6413 | Blood Stroke Distance | DeltexA | Blood Stroke Distance | cm | [] | Double | VitalSigns | Common | [] | CardioQ
6414 | Flow Time To Peak | DeltexA | Flow Time To Peak | ms | [] | Double | VitalSigns | Common | [] | CardioQ
6415 | Blood Mean Acceleration | DeltexA | Blood Mean Acceleration | m/s2 | [] | Double | VitalSigns | Common | [] | CardioQ
8117 | Corrected Flow Time | DeltexA | Corrected Flow Time | ms | [] | Double | VitalSigns | Common | [] | CardioQ
8118 | Peak Velocity | DeltexA | Peak Velocity | cm/s | [] | Double | VitalSigns | Common | [] | CardioQ
1033 | Air Bubble Detection State | XeniosA | Air Bubble Detection State | unitless | Consult detailed help for more information on this variable | Byte | DeviceInfo | Common | [] | Xenios Console
1869 | Module State | XeniosA | Module State | unitless | This variable is 16 bits long. | UShort | DeviceInfo | Common | [] | Xenios Console
2132 | AC Power Supply Failure Alarm | XeniosA | AC Power Supply Failure Alarm | unitless | This variable is 32 bits long. | ULong | Alarms | Common | [] | Xenios Console
2899 | Flow Sensor Type | XeniosA | Flow Sensor Type | unitless | Consult detailed help for more information on this variable | UByte | DeviceInfo | Common | [] | Xenios Console
3052 | Device Screen State | XeniosA | Device Screen State | unitless | Possible values : | UByte | DeviceInfo | Common | [] | Xenios Console
3267 | Pulsatile Pulse Percentage Setting | XeniosA | Pulsatile Pulse Percentage Setting | % | This variable is for pulsatile mode. | UByte | Settings | Common | [] | Xenios Console
3288 | Generic Pump Rotation Speed | XeniosA | Generic Pump Rotation Speed | rpm | This variable is associated with LowerAlarmLimit and UpperAlarmLimit attributes. | Short | VitalSigns | Common | [] | Xenios Console
3335 | Pulsatile Simulated Frequency Setting | XeniosA | Pulsatile Simulated Frequency Setting | 1/min | This variable is for pulsatile mode. | UByte | Settings | Common | [] | Xenios Console
3691 | HLM System Control Module Technical Alarm | XeniosA | HLM System Control Module Technical Alarm | unitless | This variable is 24 bits long. | ULong | Alarms | Common | [] | Xenios Console
3696 | Generic Pump Running Status | XeniosA | Generic Pump Running Status | unitless | This variable is 32 bits long. | ULong | DeviceInfo | Common | [] | Xenios Console
3701 | Generic Pump Flow Rate Setting | XeniosA | Generic Pump Flow Rate Setting | mL/min | [] | Short | Settings | Common | [] | Xenios Console
3702 | Generic Pump Rotation Speed Setting | XeniosA | Generic Pump Rotation Speed Setting | rpm | [] | Short | Settings | Common | [] | Xenios Console
3705 | Generic Pump Flow Alarm | XeniosA | Generic Pump Flow Alarm | unitless | This variable is 16 bits long. | UShort | Alarms | Common | [] | Xenios Console
5233 | Blood Level Measurement Status | XeniosA | Blood Level Measurement Status | unitless | Consult detailed help for more information on this variable | Byte | OtherMeasuredData | Common | [] | Xenios Console
7244 | Generic Pressure Line Value | XeniosA | Generic Pressure Line Value | mmHg | This variable is associated with LowerAlarmLimit and UpperAlarmLimit attributes. This variable is associated with ChannelID identifier. | Short | VitalSigns | Common | [] | Xenios Console
7945 | Pulsatile Pump Speed Amplitude Setting | XeniosA | Pulsatile Pump Speed Amplitude Setting | rpm | This variable is for pulsatile mode. | Short | Settings | Common | [] | Xenios Console
9266 | Generic ECG Alarm | XeniosA | Generic ECG Alarm | unitless | This variable is 8 bits long. | UByte | Alarms | Common | [] | Xenios Console
965 | Cumulated Therapy Time | FreseniusG | Cumulated Therapy Time | ['s', 'min'] | Treatment time, running, total | Long | OtherMeasuredData | ['Common\xa0', 'Multifiltrate PRO\xa0'] | Dialysis Course | ['Multifiltrate PRO', '5008']
973 | Device Communication Module Software Revision | FreseniusG
1012 | Device Communication Module Name | FreseniusG
1013 | Device Communication Module Serial Number | FreseniusG
1014 | Device Communication Module Hardware Revision | FreseniusG
1059 | Anticoagulant Current Continuous Infused Volume | FreseniusG
1060 | Anticoagulant Cumulated Continuous Infused Volume | FreseniusG | Anticoagulant Cumulated Continuous Infused Volume | 0.1 mL | Anticoagulation, continuous, total | Long | VitalSigns | Multifiltrate PRO | Dialysis Course | Multifiltrate PRO
1089 | Anticoagulant Current Bolus Infused Volume | FreseniusG
1090 | Anticoagulant Cumulated Bolus Infused Volume | FreseniusG | Anticoagulant Cumulated Bolus Infused Volume | ['mL', '0.1 mL'] | Anticoagulation bolus, total | ['Double\xa0', 'Long\xa0'] | VitalSigns | ['Common\xa0', 'Multifiltrate PRO\xa0'] | Dialysis Course | ['Multifiltrate PRO', '5008']
2003 | Therapy Options | FreseniusG | Therapy Options | unitless | Consult detailed help for more information on this variable | Long | DeviceInfo | Common | [] | 5008
2022 | Device System Code | FreseniusG | Device System Code | unitless | [] | String | DeviceInfo | Multifiltrate PRO | Treatment Status | Multifiltrate PRO
2200 | Device Hardware Revision | FreseniusG | Device Hardware Revision | unitless | This variable is obsolete because it may contains incorrect information. Please use variables 764, 765 and 4201 must be used instead. | String | DeviceInfo | Common | Obsolete Variables | ['MP20', 'MP70', 'MX800', 'X3']
3004 | Replacement Flow Rate Setting | FreseniusG | Replacement Flow Rate Setting | ['mL/h', 'mL/min'] | Substitution rate setting | Long | Settings | ['Common\xa0', 'Multifiltrate PRO\xa0'] | Dialysis Settings | ['Multifiltrate PRO', '5008']
3440 | Recall of Stored Data Set | FreseniusG
3593 | Device Communications Status | FreseniusG | Device Communications Status | unitless | ['Value "stop" sent when a request to stop the communication is received from the device\xa0', 'The current connection state between the Bed entity and the Bed Data Server application.'] | String | DeviceInfo | ['Common\xa0', 'Bed Data Server (Common)\xa0'] | ['Device Information\xa0', 'BED\xa0'] | ['Evita V800', 'Evita V500', 'Total Scale']
3628 | Therapy Starting Time | FreseniusG | Therapy Starting Time | unitless | ['Date/time in the format of "YYYYMMDDhhmmss"\xa0', 'Format: YYYYMMDDhhmm\xa0'] | String | OtherMeasuredData | Common | Treatment Status | ['Multifiltrate PRO', '5008']
3654 | Module Software Revision | FreseniusG
3829 | Device Total Operation Time | FreseniusG | Device Total Operation Time | unitless | Format: [hours]:[minutes] | String | DeviceInfo | Multifiltrate PRO | Treatment Status | Multifiltrate PRO
3844 | Device Screen Display Mode | FreseniusG
4365 | Device Communication Mode | FreseniusG | Device Communication Mode | unitless | Consult detailed help for more information on this variable | UByte | DeviceInfo | Common | [] | 5008
4485 | Device Communication Software Revision | FreseniusG
5849 | Therapy Duration Setting | FreseniusG | Therapy Duration Setting | s | [] | Long | Settings | Common | [] | 5008
5860 | Blood Sensed Status | FreseniusG | Blood Sensed Status | unitless | 0=no blood sensed 1=blood sensed | Long | DeviceInfo | Multifiltrate PRO | Dialysis Course | Multifiltrate PRO
5867 | Filtration Volume Setting | FreseniusG
5868 | Fluid Removal Rate Setting | FreseniusG | Fluid Removal Rate Setting | mL/h | UF rate setting | Long | Settings | Common | Dialysis Settings | ['Multifiltrate PRO', '5008']
5869 | Dialysate Temperature Setting | FreseniusG | Dialysate Temperature Setting | ['°C', '0.1 °C'] | [] | ['Double\xa0', 'Long\xa0'] | Settings | Common | Dialysis Settings | ['Multifiltrate PRO', '5008']
5870 | Plasma Volume Setting | FreseniusG | Plasma Volume Setting | mL | [] | Long | Settings | Multifiltrate PRO | Dialysis Settings | Multifiltrate PRO
5872 | Calcium Dose Setting | FreseniusG | Calcium Dose Setting | 0.01 mmol/L | [] | Long | Settings | Multifiltrate PRO | Dialysis Settings | Multifiltrate PRO
5875 | Citrate Flow Rate | FreseniusG | Citrate Flow Rate | mL/h | [] | Long | VitalSigns | Multifiltrate PRO | Dialysis Course | Multifiltrate PRO
5876 | Calcium Flow Rate | FreseniusG | Calcium Flow Rate | mL/h | [] | Long | VitalSigns | Multifiltrate PRO | Dialysis Course | Multifiltrate PRO
5877 | Citrate Current Volume | FreseniusG
5878 | Citrate Cumulated Volume | FreseniusG | Citrate Cumulated Volume | 0.1 mL | [] | Long | VitalSigns | Multifiltrate PRO | Dialysis Course | Multifiltrate PRO
5879 | Calcium Current Volume | FreseniusG
5880 | Calcium Cumulated Volume | FreseniusG | Calcium Cumulated Volume | 0.1 mL | [] | Long | VitalSigns | Multifiltrate PRO | Dialysis Course | Multifiltrate PRO
5881 | Filtration Fluid Current Balance | FreseniusG
5882 | Filtration Fluid Cumulated Balance | FreseniusG | Filtration Fluid Cumulated Balance | mL | DF-Total, total balance | Long | VitalSigns | Multifiltrate PRO | Dialysis Course | Multifiltrate PRO
5883 | Replacement Current Bolus Volume | FreseniusG
5884 | Replacement Cumulated Bolus Volume | FreseniusG | Replacement Cumulated Bolus Volume | mL | Bolus-Substituate-Total, total | Long | VitalSigns | ['Common\xa0', 'Multifiltrate PRO\xa0'] | Dialysis Course | ['Multifiltrate PRO', '5008', 'Artis']
6768 | Device Quick Test Status | FreseniusG | Device Quick Test Status | unitless | T0 Test Status. Possible values: Start, End, E1, E2, E3. | String | DeviceInfo | Multifiltrate PRO | Treatment Status | Multifiltrate PRO
7982 | Patient Fluid Current Balance | FreseniusG
7983 | Filtration Current Volume | FreseniusG | Filtration Current Volume | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Vufr/h | ['MP70', 'MX800', 'X3']
7984 | Fluid Removal Current Volume | FreseniusG
7985 | Replacement Current Volume | FreseniusG
8696 | Generic Message Text | FreseniusG | Generic Message Text | unitless | This variable is in multiset format. This variable, along with variable 8366, contains text message code and alarm text respectively. They need to be interpreted together. | String (Multiset) | Messages | ['Evita Infinity V500 [MEDIBUS.X]\xa0', 'Evita Infinity V800\xa0'] | Alarm And Text Message | ['Evita V800', 'Evita V500']
8967 | Case ID | FreseniusG | Case ID | unitless | [] | String | PatientInfo | Common | Administrative Data | Multifiltrate PRO
9522 | Device Manufacturer | FreseniusG | Device Manufacturer | unitless | [] | String | DeviceInfo | Common | [] | 5008
9644 | Prefilter Pod Pressure | FreseniusG | Prefilter Pod Pressure | mmHg | Pre-Hemofilter pressure | Long | OtherMeasuredData | Multifiltrate PRO | Dialysis Course | Multifiltrate PRO
9662 | Device Location | FreseniusG | Device Location | unitless | [] | String | DeviceInfo | Common | [] | 5008
2 |  |  | NIBP Systolic | ['-', 'mmHg'] | ['The MeasurementTime of BP variables are only updated when a new measurement is performed. In between measurements, the last valid values are sent repeatedly.\xa0', 'This variable is associated with LowerAlarmLimit and UpperAlarmLimit attributes.\xa0', 'Possible units: mmHg or kPa.'] | ['Double\xa0', 'Long\xa0'] | VitalSigns | Common | ['NBP\xa0', 'NIBP\xa0'] | ['X3', 'Artis', 'MX800', 'MP20', '5008', 'MP70', 'Expression IP5 et MR400']
3 |  |  | NIBP Diastolic | ['-', 'mmHg'] | ['The MeasurementTime of BP variables are only updated when a new measurement is performed. In between measurements, the last valid values are sent repeatedly.\xa0', 'This variable is associated with LowerAlarmLimit and UpperAlarmLimit attributes.\xa0', 'Possible units: mmHg or kPa.'] | ['Double\xa0', 'Long\xa0'] | VitalSigns | Common | ['NBP\xa0', 'NIBP\xa0'] | ['X3', 'Artis', 'MX800', 'MP20', '5008', 'MP70', 'Expression IP5 et MR400']
4 |  |  | NIBP Mean | ['-', 'mmHg'] | ['This variable is associated with LowerAlarmLimit and UpperAlarmLimit attributes.\xa0', 'Possible units: mmHg or kPa.'] | ['Double\xa0', 'Long\xa0'] | VitalSigns | Common | ['NBP\xa0', 'NIBP\xa0'] | ['X3', 'MX800', 'MP20', '5008', 'MP70', 'Expression IP5 et MR400']
7 |  |  | Patient Delta Temperature | ['°C', '-'] | [] | Double | VitalSigns | Common | DeltaTemp | ['X3', 'MX800', 'MP20', '5008', 'MP70']
8 |  |  | INVP1 Systolic | - | Possible units: mmHg or kPa. If PCMCIA protocol is used, the unit code of this variable is always 0 (unit unknown). | Double | VitalSigns | Common | IBP | Expression IP5 et MR400
9 |  |  | INVP1 Diastolic | - | Possible units: mmHg or kPa. If PCMCIA protocol is used, the unit code of this variable is always 0 (unit unknown). | Double | VitalSigns | Common | IBP | Expression IP5 et MR400
10 |  |  | INVP1 Mean | - | Possible units: mmHg or kPa. If PCMCIA protocol is used, the unit code of this variable is always 0 (unit unknown). | Double | VitalSigns | Common | IBP | Expression IP5 et MR400
11 |  |  | INVP2 Systolic | - | Possible units: mmHg or kPa. If PCMCIA protocol is used, the unit code of this variable is always 0 (unit unknown). | Double | VitalSigns | Common | IBP | Expression IP5 et MR400
12 |  |  | INVP2 Diastolic | - | Possible units: mmHg or kPa. If PCMCIA protocol is used, the unit code of this variable is always 0 (unit unknown). | Double | VitalSigns | Common | IBP | Expression IP5 et MR400
13 |  |  | INVP2 Mean | - | Possible units: mmHg or kPa. If PCMCIA protocol is used, the unit code of this variable is always 0 (unit unknown). | Double | VitalSigns | Common | IBP | Expression IP5 et MR400
16 |  |  | NIBP Pulse Rate | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Pulse | ['MP70', 'MX800', 'X3']
19 |  |  | ECG Heart Rate | - | [] | Double | VitalSigns | Common | ECG: HR | ['MP20', 'MP70', 'MX800', 'X3']
28 |  |  | Active Alarms | unitless | See "Multiset Alarms" on help file. | Double (Multiset) | Alarms | Common | Alarms | ['MP20', 'MP70', 'MX800', 'X3']
29 |  |  | Latched Alarms | unitless | See "Multiset Alarms" on help file. | Double (Multiset) | Alarms | Common | Alarms | ['MP20', 'MP70', 'MX800', 'X3']
33 |  |  | Inactive Alarms | unitless | See "Multiset Alarms" on help file. | Double (Multiset) | Alarms | Common | Alarms | ['MP20', 'MP70', 'MX800', 'X3']
34 |  |  | General ST Level 1 | - | Physiological identifier: NOM_ECG_AMPL_ST_I | Double | VitalSigns | Common | ECG: ST | ['MP20', 'MP70', 'MX800', 'X3']
35 |  |  | General ST Level 2 | - | Physiological identifier: NOM_ECG_AMPL_ST_II | Double | VitalSigns | Common | ECG: ST | ['MP20', 'MP70', 'MX800', 'X3']
36 |  |  | General ST Level 3 | - | Physiological identifier: NOM_ECG_AMPL_ST_III | Double | VitalSigns | Common | ECG: ST | ['MP20', 'MP70', 'MX800', 'X3']
45 |  |  | Temperature 3 | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
46 |  |  | Temperature 4 | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
48 |  |  | O2 Saturation (SO2) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | SO2 | ['MP70', 'MX800', 'X3']
127 |  |  | ECG Measurement Status | - | ECG Rhythm Status | Double | OtherMeasuredData | Common | ECG Status (on IntelliVue) | ['MP20', 'MP70', 'MX800', 'X3']
168 |  |  | Device ID | unitless | ['The globally unique identifier for a Bed entity as defined by Hill-Rom (e.g. “VersaCare-0x0001”). \xa0', '1–8 characters length alphanumeric string. For MR Patient Care Portal 5000, the length can be up to 31 characters.\xa0'] | String | DeviceInfo | ['Expression IP5\xa0', 'Bed Data Server (Common)\xa0'] | ['Monitor Info (EXPRESS Plus, Express Connect Only)\xa0', 'BED\xa0'] | ['Expression IP5 et MR400', 'Total Scale']
170 |  |  | Patient Sex | unitless | Possible values: "unknown", "male", "female", and "unspecified". | String | PatientInfo | Common | Patient Demographics | ['MP20', 'MP70', 'MX800', 'X3']
177 |  |  | Alarm Data | unitless | This variable is same as 3878 but sent as a string, which is an array containing all alarm messages and separated by delimiter "¦¦". This variable, along with variables 588 and 4131, are marked as obsolete because they use either"¦¦" or "," in the array string as delimiter, which are often not compatible with customers' applications. Variables 6164, 3878, and 2596 should be used instead. The HL7 output can contain an '&' character to indicate a LOW alarm. In order to avoid issues, the '&' character will be replaced by a '?' character in the DDI output. | String | Alarms | Evita Infinity V500 [MEDIBUS.X] | Alarm And Text Message | Evita V500
356 |  |  | Generic Dialysate Alarm | unitless | Value of "1" when alarm is active. | Boolean | Alarms | Common | [] | 5008
357 |  |  | Urea Clearance Alarm | unitless | Value of "1" when alarm is active. | Boolean | Alarms | Common | [] | 5008
358 |  |  | Blood Volume Alarm | unitless | Value of "1" when alarm is active. | Boolean | Alarms | Common | [] | 5008
359 |  |  | Mixed Dilution Alarm | unitless | Value of "1" when alarm is active. | Boolean | Alarms | Common | [] | 5008
393 |  |  | Single Needle Ratio Setting | % | [] | Long | Settings | Common | [] | 5008
405 |  |  | Generic O2 Saturation (1) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | SO2 1 | ['MP70', 'MX800', 'X3']
406 |  |  | Generic O2 Saturation (2) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | SO2 2 | ['MP70', 'MX800', 'X3']
419 |  |  | Inspired O2 (FiO2) | % | [] | Double | VitalSigns | Common | Resp | Expression IP5 et MR400
426 |  |  | Inspired N2O | % | [] | Double | VitalSigns | Expression IP5 | Agent (EXPRESS Plus, Express Connect Only) | Expression IP5 et MR400
427 |  |  | End Tidal N2O | % | [] | Double | VitalSigns | Common | Agent | Expression IP5 et MR400
444 |  |  | Target Treatement Effectiveness (Kt/V) Remaining Time | s | [] | Long | VitalSigns | Common | [] | 5008
448 |  |  | Shunt Blood Flow Rate | mL/min | [] | Short | VitalSigns | Common | [] | 5008
449 |  |  | Recirculation Measurement Active Status | unitless | Possible values: 0=off, 1=on | Boolean | DeviceInfo | Common | [] | 5008
450 |  |  | Isolated Ultrafiltration Therapy Time | s | [] | Long | VitalSigns | Common | [] | 5008
451 |  |  | Isolated Ultrafiltration Rate | mL/h | [] | Long | VitalSigns | Common | [] | 5008
452 |  |  | Initial Hemoglobin | g/dL | [] | Double | VitalSigns | Common | [] | 5008
453 |  |  | Initial Protein Concentration | g/dL | [] | Double | VitalSigns | Common | [] | 5008
454 |  |  | Single Needle Time | min | [] | Long | VitalSigns | Common | [] | 5008
455 |  |  | Blood Outlet Pressure | mmHg | [] | Short | VitalSigns | Common | [] | 5008
456 |  |  | Dialysate Inlet Pressure | mmHg | [] | ['Long\xa0', 'Short\xa0'] | VitalSigns | Common | [] | ['5008', 'Artis']
457 |  |  | Dialysate Outlet Pressure | mmHg | [] | ['Long\xa0', 'Short\xa0'] | VitalSigns | Common | [] | ['5008', 'Artis']
458 |  |  | Maximum Transmembrane Pressure | mmHg | [] | Short | VitalSigns | Common | [] | 5008
459 |  |  | Auto Dialysate Flow Enabled Status | unitless | Possible values: 0=disabled, 1=enabled | Boolean | DeviceInfo | Common | [] | 5008
460 |  |  | Auto Dialysate Flow Factor Setting | unitless | [] | Double | Settings | Common | [] | 5008
461 |  |  | Auto AntiCoagulant Enabled Status | unitless | Possible values: 0=disabled, 1=enabled | Boolean | DeviceInfo | Common | [] | 5008
462 |  |  | Auto Anticoagulant Bolus Enabled Status | unitless | Possible values: 0=off, 1=on | Boolean | DeviceInfo | Common | [] | 5008
463 |  |  | Auto Anticoagulant Bolus Rate Enabled Status | unitless | Possible values: 0=disabled, 1=enabled | Boolean | DeviceInfo | Common | [] | 5008
464 |  |  | Single Needle Auto Rate Enabled Status | unitless | Possible values: 0=disabled, 1=enabled | Boolean | DeviceInfo | Common | [] | 5008
465 |  |  | Fetal Heart Rate (FHR) | - | This variable is associated with ChannelID identifier. | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | DFHR/DFHR1/DFHR2/DFHR3 | ['MP70', 'MX800', 'X3']
467 |  |  | Fetal Beat-to-Beat Rate (FHR) | - | This variable is associated with ChannelID identifier. | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | bDFHR/bDFHR1/bDFHR2/bDFHR3 | ['MP70', 'MX800', 'X3']
468 |  |  | Replacement Pump Running Status | unitless | Possible values: 0=off, 1=on | Boolean | DeviceInfo | Common | [] | 5008
469 |  |  | TOCO | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Toco | ['MP70', 'MX800', 'X3']
497 |  |  | ECG Lead Name | - | Physiological identifier: NOM_ECG_AMPL_ST | Double | OtherMeasuredData | Common | Available only in revF or earlier | ['MP20', 'MP70', 'MX800', 'X3']
501 |  |  | Arterial Pressure (ART) Systolic | - | [] | Double | VitalSigns | Common | IBP: ART | ['MP20', 'MP70', 'MX800', 'X3']
502 |  |  | Arterial Pressure (ART) Diastolic | - | [] | Double | VitalSigns | Common | IBP: ART | ['MP20', 'MP70', 'MX800', 'X3']
503 |  |  | Arterial Pressure (ART) Mean | - | [] | Double | VitalSigns | Common | IBP: ART | ['MP20', 'MP70', 'MX800', 'X3']
508 |  |  | Clearance Measurement Interval Setting | s | [] | UShort | Settings | Common | [] | 5008
512 |  |  | Left Atrial Mean Pressure (LAP) | - | [] | Double | VitalSigns | Common | IBP: LAP | ['MP20', 'MP70', 'MX800', 'X3']
538 |  |  | Device Language | unitless | Consult detailed help for more information on this variable | String | DeviceInfo | Common | System Information | ['MP20', 'MP70', 'MX800', 'X3']
540 |  |  | Low Venous Line Pressure Alarm | unitless | Value of "1" when alarm is active. | Boolean | Alarms | Common | [] | 5008
542 |  |  | Maximum Tidal Volume Limit | - | Possible units: mL, L | Double | Settings | Common | [] | V60
543 |  |  | High Venous Line Pressure Alarm | unitless | Value of "1" when alarm is active. | Boolean | Alarms | Common | [] | 5008
546 |  |  | High Arterial Line Pressure Alarm | unitless | Value of "1" when alarm is active. | Boolean | Alarms | Common | [] | 5008
547 |  |  | Low Arterial Line Pressure Alarm | unitless | Value of "1" when alarm is active. | Boolean | Alarms | Common | [] | 5008
558 |  |  | Heparin Pump Alarm | unitless | Value of "1" when alarm is active. | Boolean | Alarms | Common | [] | 5008
569 |  |  | Patient Specific Filtration Ratio Setting | % | [] | UShort | Settings | Common | [] | 5008
576 |  |  | Average Dialysate Temperature | °C | [] | Double | VitalSigns | Common | [] | 5008
588 |  |  | Alarm Message | unitless | This variable is same as 6164 but sent as a string, which is an array containing all alarm codes and separated by delimiter ",". This variable, along with variables 177 and 4131, are marked as obsolete because they use either"¦¦" or "," in the array string as delimiter, which are often not compatible with customers' applications. Variables 6164, 3878, and 2596 should be used instead. | String | Alarms | Evita Infinity V500 [MEDIBUS.X] | Alarm And Text Message | Evita V500
591 |  |  | ECG Lead II | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
592 |  |  | Airway N2O Concentration | - | Measurement from VueLink | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | N2O | ['MP70', 'MX800', 'X3']
593 |  |  | Inspired Desflurane | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | DES | ['MP70', 'MX800', 'X3']
595 |  |  | Inspired Sevoflurane | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | SEV | ['MP70', 'MX800', 'X3']
597 |  |  | Fresh Gas Agent Value | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | fgAGT | ['MP70', 'MX800', 'X3']
598 |  |  | Fresh Gas Halothane | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | fgHAL | ['MP70', 'MX800', 'X3']
599 |  |  | Fresh Gas Enflurane | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | fgENF | ['MP70', 'MX800', 'X3']
600 |  |  | Fresh Gas Isoflurane | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | fgISO | ['MP70', 'MX800', 'X3']
601 |  |  | Inspired Halothane | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | HAL | ['MP70', 'MX800', 'X3']
603 |  |  | Inspired Enflurane | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | ENF | ['MP70', 'MX800', 'X3']
605 |  |  | Inspired Isoflurane | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | ISO | ['MP70', 'MX800', 'X3']
610 |  |  | Cuff Inflation Pressure | mmHg | [] | UShort | OtherMeasuredData | Common | [] | 5008
620 |  |  | Halothane Concentration Wave | - | [] | Double | Waveforms | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Waves: HAL | ['MP70', 'MX800', 'X3']
621 |  |  | Enflurane Concentration Wave | - | [] | Double | Waveforms | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Waves: ENF | ['MP70', 'MX800', 'X3']
622 |  |  | Isoflurane Concentration Wave | - | [] | Double | Waveforms | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Waves: ISO | ['MP70', 'MX800', 'X3']
623 |  |  | Desflurane Concentration Wave | - | [] | Double | Waveforms | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Waves: DES | ['MP70', 'MX800', 'X3']
624 |  |  | Sevoflurane Concentration Wave | - | [] | Double | Waveforms | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Waves: SEV | ['MP70', 'MX800', 'X3']
625 |  |  | Apnea Duration | - | This is the device's measurement with label="NLS_NOM_TIME_PD_APNEA" and physio_id="NOM_TIME_PD_APNEA". | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | ApneaD | ['MP70', 'MX800', 'X3']
626 |  |  | End Tidal CO2 | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | CRT | ['MP70', 'MX800', 'X3']
630 |  |  | Cardiac Beat-to-Beat Rate | - | Cardiac Beat-to-Beat Rate (NOM_ECG_CARD_BEAT_RATE_BTB). | Double | VitalSigns | Common | ECG: btbHR | ['MP20', 'MP70', 'MX800', 'X3']
633 |  |  | Airway Delta O2 Concentration | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | VENT: DeltaO2 | ['MP70', 'MX800', 'X3']
669 |  |  | Normalized Ionized Calcium Concentration (Ca) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | iCa-N | ['MP70', 'MX800', 'X3']
676 |  |  | Bicarbonate Concentration (HCO3) | mmol/L | [] | Double | VitalSigns | Common | [] | 5008
678 |  |  | Potassium Ion Concentration (K+) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
696 |  |  | QT Interval (QT) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | ECG: QT | ['MP70', 'MX800', 'X3']
762 |  |  | Predilution Enable Status | unitless | Internal status for the mixed HDF control | UShort | DeviceInfo | Common | [] | 5008
764 |  |  | Device Specification ID | unitless | Variables 764, 765 and 4201 must be interpreted conjointly. | String (Multiset) | DeviceInfo | Common | System Information | ['MP20', 'MP70', 'MX800', 'X3']
765 |  |  | Device Specification Type | unitless | Variables 764, 765 and 4201 must be interpreted conjointly. | String (Multiset) | DeviceInfo | Common | System Information | ['MP20', 'MP70', 'MX800', 'X3']
794 |  |  | Pleth Wave (2) | - | [] | Double | Waveforms | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Waves: PLETH2 | ['MP70', 'MX800', 'X3']
845 |  |  | Clearance Measurement Enabled Status | unitless | Possible values: 0=off, 1=on | Boolean | DeviceInfo | Common | [] | 5008
846 |  |  | Maximum Ultrafiltration Rate Setting | mL/h | [] | Long | Settings | Common | [] | 5008
847 |  |  | Ultrafiltration Goal Positive Deviation Setting | mL | [] | Long | Settings | Common | [] | 5008
848 |  |  | Ultrafiltration Goal Negative Deviation Setting | mL | [] | Long | Settings | Common | [] | 5008
851 |  |  | Residual Blood Volume Active Status | unitless | Possible values: 0=off, 1=on | Boolean | DeviceInfo | Common | [] | 5008
852 |  |  | Temperature Change Rate Setting | °C/h | [] | Double | Settings | Common | [] | 5008
859 |  |  | Reinfusion Type Setting | unitless | Consult detailed help for more information on this variable | UShort | Settings | Common | [] | 5008
860 |  |  | Auto Reinfusion Active Status | unitless | Possible values: 0=off, 1=on | Boolean | DeviceInfo | Common | [] | 5008
914 |  |  | Low Volume Treatment Active Status | unitless | Possible values: 0=off, 1=on | Boolean | DeviceInfo | Common | [] | 5008
955 |  |  | Battery Running Alarm | unitless | ['Value of "1" when alarm is active.\xa0', 'Consult detailed help for more information on this variable\xa0'] | ['Boolean\xa0', 'Long\xa0'] | Alarms | Common | [] | ['V60', '5008', 'Monnal T60']
967 |  |  | Current Therapy ID | unitless | [] | String | PatientInfo | Common | [] | 5008
980 |  |  | Base Device ID | unitless | The globally unique identifier for the BedConnector entity associated with this Bed entity or an empty string if no association currently exists. | String | DeviceInfo | Bed Data Server (Common) | BED | Total Scale
982 |  |  | Device IP Address | unitless | [] | String | DeviceInfo | Common | [] | 5008
1055 |  |  | Infusion Drug Name | - | [] | Double | PatientInfo | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sDrug | ['MP70', 'MX800', 'X3']
1072 |  |  | Infusion Time Remaining | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | InfTim | ['MP70', 'MX800', 'X3']
1095 |  |  | Infusion Dose Rate Setting | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sDosRt | ['MP70', 'MX800', 'X3']
1096 |  |  | Volume To Be Infused Original Setting | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sVtot | ['MP70', 'MX800', 'X3']
1100 |  |  | Average Anticoagulant Rate | mL/h | [] | Double | VitalSigns | Common | [] | 5008
1101 |  |  | Single Needle Stroke Volume | mL | [] | Long | Settings | Common | [] | 5008
1102 |  |  | Average Urea Clearance | mL/min | [] | Long | VitalSigns | Common | [] | 5008
1103 |  |  | Online Bolus Volume | mL | [] | Long | VitalSigns | Common | [] | 5008
1104 |  |  | Online Bolus Rate Setting | mL/min | [] | Long | Settings | Common | [] | 5008
1110 |  |  | KVO Rate Setting | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sKVO | ['MP70', 'MX800', 'X3']
1186 |  |  | Inspired Minute Volume | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | MINVOL | ['MP70', 'MX800', 'X3']
1197 |  |  | Message Text | unitless | This variable, along with variable 3909 are marked as obsolete because they use either"¦¦" or "," in the array string as delimiter, which are often not compatible with customers' applications. Variables 8366 and 8696 should be used instead. | String | Messages | Evita Infinity V500 [MEDIBUS.X] | Alarm And Text Message | Evita V500
1208 |  |  | SpO2 | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
1220 |  |  | Patient Birth Date | unitless | ['Format: YYYYMMDD\xa0', 'Format: 14 ascii digits string of "YYYYMMDDhhmmss"\xa0'] | String | PatientInfo | Common | Patient Demographics | ['X3', 'MX800', 'MP20', '5008', 'MP70']
1228 |  |  | High Inspired Tidal Volume Alarm | unitless | [] | Boolean | Alarms | Common | [] | Monnal T60
1234 |  |  | SpO2 (Pre-Ductal) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | SpO2pr | ['MP70', 'MX800', 'X3']
1236 |  |  | SpO2 (Post-Ductal) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | SpO2po | ['MP70', 'MX800', 'X3']
1244 |  |  | ECG Lead III | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
1250 |  |  | Pulmonary Vascular Resistance | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | PVR | ['MP70', 'MX800', 'X3']
1251 |  |  | Body Surface Area | - | [] | Double | PatientInfo | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | BSA | ['MP70', 'MX800', 'X3']
1255 |  |  | ECG Lead MCL1 | - | Physiological identifier: NOM_ECG_AMPL_ST_MCL | Double | VitalSigns | Common | ECG: ST | ['MP20', 'MP70', 'MX800', 'X3']
1257 |  |  | ECG Lead V1 | - | Physiological identifier: NOM_ECG_AMPL_ST_V1 | Double | VitalSigns | Common | ECG: ST | ['MP20', 'MP70', 'MX800', 'X3']
1258 |  |  | ECG Lead V6 | - | Physiological identifier: NOM_ECG_AMPL_ST_V6 | Double | VitalSigns | Common | ECG: ST | ['MP20', 'MP70', 'MX800', 'X3']
1259 |  |  | ECG Lead aVL | - | Physiological identifier: NOM_ECG_AMPL_ST_AVL | Double | VitalSigns | Common | ECG: ST | ['MP20', 'MP70', 'MX800', 'X3']
1260 |  |  | ECG Lead aVF | - | Physiological identifier: NOM_ECG_AMPL_ST_AVF | Double | VitalSigns | Common | ECG: ST | ['MP20', 'MP70', 'MX800', 'X3']
1261 |  |  | ECG Lead aVR | - | Physiological identifier: NOM_ECG_AMPL_ST_AVR | Double | VitalSigns | Common | ECG: ST | ['MP20', 'MP70', 'MX800', 'X3']
1262 |  |  | ECG Lead V2 | - | Physiological identifier: NOM_ECG_AMPL_ST_V2 | Double | VitalSigns | Common | ECG: ST | ['MP20', 'MP70', 'MX800', 'X3']
1263 |  |  | ECG Lead V3 | - | Physiological identifier: NOM_ECG_AMPL_ST_V3 | Double | VitalSigns | Common | ECG: ST | ['MP20', 'MP70', 'MX800', 'X3']
1264 |  |  | ECG Lead V4 | - | Physiological identifier: NOM_ECG_AMPL_ST_V4 | Double | VitalSigns | Common | ECG: ST | ['MP20', 'MP70', 'MX800', 'X3']
1265 |  |  | ECG Lead V5 | - | Physiological identifier: NOM_ECG_AMPL_ST_V5 | Double | VitalSigns | Common | ECG: ST | ['MP20', 'MP70', 'MX800', 'X3']
1283 |  |  | Patient Temperature | - | This variable is associated with a SourceID Identifier. | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Infrared Temp | ['MP70', 'MX800', 'X3']
1286 |  |  | Micro Environment Air Temperature | - | [] | Double | OtherMeasuredData | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | INCUBATOR: Air T | ['MP70', 'MX800', 'X3']
1318 |  |  | Static Compliance | mL/mbar | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500 Evita Infinity V600, Evita Infinity V800. For Evita Infinity V500: Cstat (Low Flow) Cstat. | Double | VitalSigns | ['Evita Infinity V500 [MEDIBUS.X]\xa0', 'Evita Infinity V800\xa0'] | [] | ['Evita V800', 'Evita V500']
1319 |  |  | Static Resistance | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Rstat | ['MP70', 'MX800', 'X3']
1323 |  |  | Pressure Sensitivity Setting | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sSenPr | ['MP70', 'MX800', 'X3']
1325 |  |  | Sigh Volume Setting | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sVsigh | ['MP70', 'MX800', 'X3']
1340 |  |  | Negative Inspiratory Pressure | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | NgInsP | ['MP70', 'MX800', 'X3']
1342 |  |  | Peak Spontaneous Flow | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | SpPkFl | ['MP70', 'MX800', 'X3']
1343 |  |  | Flow-by Base Flow Setting | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | VENT: sBasFl | ['MP70', 'MX800', 'X3']
1344 |  |  | Flow Sensitivity Setting | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | VENT: sSenFl | ['MP70', 'MX800', 'X3']
1348 |  |  | Expiratory Time | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | ExpTi | ['MP70', 'MX800', 'X3']
1378 |  |  | Device Configuration | unitless | For CritiCool: This variable contains information on the fixed step rate at which the patient gradually warms up. For Allon: This variable contains information on if the temperature alarms are on or off.Consult detailed help for more information on this variable | String | DeviceInfo | Common | [] | Criticool
1382 |  |  | Expiratory Flow Setting | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sExpFl | ['MP70', 'MX800', 'X3']
1398 |  |  | Minute Volume Setting | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sMV | ['MP70', 'MX800', 'X3']
1418 |  |  | Positive End Expiratory Pressure (PEEP) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | VENT: PEEP | ['MP70', 'MX800', 'X3']
1419 |  |  | Ventilator Work of Breathing | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | WOBv | ['MP70', 'MX800', 'X3']
1424 |  |  | Patient Work of Breathing | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | WOBp | ['MP70', 'MX800', 'X3']
1436 |  |  | Pulmonary Vascular Resistance Index | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | PVRI | ['MP70', 'MX800', 'X3']
1442 |  |  | Continuous Cardiac Output | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | CCO | ['MP70', 'MX800', 'X3']
1443 |  |  | Continuous Cardiac Output Index | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | CCI | ['MP70', 'MX800', 'X3']
1463 |  |  | Hemoglobin (Hb) | ['-', 'g/dL'] | This variable is associated with SourceID identifier. | Double | VitalSigns | ['IntelliVue X3\xa0', 'Common\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sHb | ['MP70', '5008', 'MX800', 'X3']
1464 |  |  | Pleth Wave (Pre-Ductal) | - | [] | Double | Waveforms | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Waves: PLTHpr | ['MP70', 'MX800', 'X3']
1465 |  |  | Pleth Wave (Post-Ductal) | - | [] | Double | Waveforms | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Waves: PLTHpo | ['MP70', 'MX800', 'X3']
1466 |  |  | Arterial Venous O2 Content Difference | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Ca-vO2 | ['MP70', 'MX800', 'X3']
1467 |  |  | O2 Delivery (DO2) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | DO2 | ['MP70', 'MX800', 'X3']
1469 |  |  | O2 Extraction Ratio (O2ER) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | O2ER | ['MP70', 'MX800', 'X3']
1471 |  |  | O2 Delivery Index (DO2I) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | DO2I | ['MP70', 'MX800', 'X3']
1473 |  |  | Oxygenation Ratio (PaO2/FiO2) | - | PaO2 to FIO2 ratio expressed in mmHg to % ratio. | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | PaFIO2 | ['MP70', 'MX800', 'X3']
1475 |  |  | Physiologic Dead Space | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | VENT: Vd | ['MP70', 'MX800', 'X3']
1477 |  |  | Alveolar Ventilation Volume (VA) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | VENT: ALVENT | ['MP70', 'MX800', 'X3']
1478 |  |  | Urea Nitrogen Concentration (BUN) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | BUN | ['MP70', 'MX800', 'X3']
1479 |  |  | Creatinine Concentration (Crea) | - | Measured by the i-Stat module. | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Crea | ['MP70', 'MX800', 'X3']
1480 |  |  | Plasma Osmolarity (Osm) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | PlOsm | ['MP70', 'MX800', 'X3']
1482 |  |  | Urine Creatinine | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | UrCrea | ['MP70', 'MX800', 'X3']
1484 |  |  | Urine Potassium (K+) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | UrK | ['MP70', 'MX800', 'X3']
1486 |  |  | Urine Osmolarity | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | UrOsm | ['MP70', 'MX800', 'X3']
1487 |  |  | Blood Urea Nitrogen Creatinine Ratio (BUN/cr) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | BUN/cr | ['MP70', 'MX800', 'X3']
1488 |  |  | Osmolar Clearance (COSM) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | COsm | ['MP70', 'MX800', 'X3']
1489 |  |  | Creatinine Clearance (CreaCl) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | CreaCl | ['MP70', 'MX800', 'X3']
1490 |  |  | Fractional Excretion of Sodium (FeNa) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | FeNa | ['MP70', 'MX800', 'X3']
1493 |  |  | Urine Serum Creatinine Ratio | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | U/SCr | ['MP70', 'MX800', 'X3']
1495 |  |  | Urine Potassium Excretion | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | UrKEx | ['MP70', 'MX800', 'X3']
1496 |  |  | Urine Sodium Excretion | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | UrNaEx | ['MP70', 'MX800', 'X3']
1497 |  |  | Free Water Clearance (CH2O) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | CH2O | ['MP70', 'MX800', 'X3']
1510 |  |  | Concentrate Source | unitless | Consult detailed help for more information on this variable | Long | DeviceInfo | Common | [] | 5008
1523 |  |  | Premature Ventricular Contractions | - | [] | Double | VitalSigns | Common | ECG: PVC | ['MP20', 'MP70', 'MX800', 'X3']
1525 |  |  | Maximum Fluid Removal Rate Setting | mL/h | [] | Long | Settings | Common | [] | 5008
1533 |  |  | LAP Systolic | - | [] | Double | VitalSigns | Common | IBP: LAP | ['MP20', 'MP70', 'MX800', 'X3']
1534 |  |  | LAP Diastolic | - | [] | Double | VitalSigns | Common | IBP: LAP | ['MP20', 'MP70', 'MX800', 'X3']
1540 |  |  | TcGas Heating Power | - | [] | Double | DeviceInfo | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | TCUT: HeatPw | ['MP70', 'MX800', 'X3']
1556 |  |  | Abdominal Fetal Heart Rate (FHR) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | aFHR | ['MP70', 'MX800', 'X3']
1557 |  |  | Abdominal Fetal Beat-to-Beat Rate (FHR) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | baFHR | ['MP70', 'MX800', 'X3']
1558 |  |  | Patient Activity Level (Movement) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | mAct | ['MP70', 'MX800', 'X3']
1559 |  |  | Patient Activity Level (Wearable) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | wAct | ['MP70', 'MX800', 'X3']
1573 |  |  | Patient CO2 Production | - | Volume of CO2 produced by respiration. | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | VCO2 | ['MP70', 'MX800', 'X3']
1574 |  |  | Relative Dead Space | - | Relative Dead Space Volume. | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | VENT: relVd | ['MP70', 'MX800', 'X3']
1575 |  |  | Respired Volume Wave | mL | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Evita V300. Respiratory volume since start of inspiration, Volume. For Evita Infinity V500, Babylog VN500: Volume since inspiration began. | Double | Waveforms | ['Evita Infinity V500 [MEDIBUS.X]\xa0', 'Evita Infinity V800\xa0'] | [] | ['Evita V800', 'Evita V500']
1590 |  |  | Procalcitonin (PCT) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | PCT | ['MP70', 'MX800', 'X3']
1591 |  |  | Average Blood Flow Rate | mL/min | [] | Long | VitalSigns | Common | [] | 5008
1602 |  |  | Reinfusion Volume | mL | [] | Long | VitalSigns | Common | [] | 5008
1607 |  |  | Fetal Heart Rate (USound) | - | This variable is associated with ChannelID identifier. | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | FHR/FHR1/FHR2/FHR3 | ['MP70', 'MX800', 'X3']
1609 |  |  | Fetal Beat-to-Beat Rate (USound) | - | This variable is associated with ChannelID identifier. | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | btbFHR/bFHR1/bFHR2/bFHR3 | ['MP70', 'MX800', 'X3']
1630 |  |  | HFV Mean Pressure Setting | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sHFMAP | ['MP70', 'MX800', 'X3']
1631 |  |  | HFV Tidal Volume Setting | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sHFVTV | ['MP70', 'MX800', 'X3']
1637 |  |  | Ventilation Pressure Variability Setting | % | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800. For Evita Infinity V500, Babylog VN500: Supported through Medibus.X protocol. | Double | Settings | ['Evita Infinity V500 [MEDIBUS.X]\xa0', 'Evita Infinity V800\xa0'] | [] | ['Evita V800', 'Evita V500']
1638 |  |  | APRV Low Time Limitation | s | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. Maximum time of low pressure level during APRV/AutoRelease, Tlow max. For Evita Infinity V500, Babylog VN500: Tlow max. | Double | Settings | ['Evita Infinity V500 [MEDIBUS.X]\xa0', 'Evita Infinity V800\xa0'] | [] | ['Evita V800', 'Evita V500']
1640 |  |  | Spontaneous Mean Tidal Volume | mL | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500 Evita Infinity V600, Evita Infinity V800. Spontaneous mean tidal volume, VTspon mean. For Evita Infinity V500, Babylog VN500: VTspon mean (Not available on Babylog VN500). Only available for Evita Infinity V500 | Double | VitalSigns | ['Evita Infinity V500 [MEDIBUS.X]\xa0', 'Evita Infinity V800\xa0'] | [] | ['Evita V800', 'Evita V500']
1641 |  |  | Spontaneous Inspired Mean Tidal Volume | mL | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500 Evita Infinity V600, Evita Infinity V800. For Evita Infinity V500, Babylog VN500: VTispon mean (Not available on Babylog VN500). Only available for Evita Infinity V500 | Double | VitalSigns | ['Evita Infinity V500 [MEDIBUS.X]\xa0', 'Evita Infinity V800\xa0'] | [] | ['Evita V800', 'Evita V500']
1642 |  |  | Spontaneous Expired Mean Tidal Volume | mL | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800. For Evita Infinity V500, Babylog VN500: VTespon mean (Not available on Babylog VN500). Only available for Evita Infinity V500 | Double | VitalSigns | ['Evita Infinity V500 [MEDIBUS.X]\xa0', 'Evita Infinity V800\xa0'] | [] | ['Evita V800', 'Evita V500']
1645 |  |  | PAV-based Resistance Setting | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sRppv | ['MP70', 'MX800', 'X3']
1646 |  |  | PAV-based Elastance Setting | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sEppv | ['MP70', 'MX800', 'X3']
1650 |  |  | SmartCare FiO2 Limitation | % | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. FiO2 in SmartCare - upper limit, SC-FiO2high For Evita Infinity V500, Babylog VN500: SC-Upper limit inspiratory O2 FiO2 max. | Double | Settings | ['Evita Infinity V500 [MEDIBUS.X]\xa0', 'Evita Infinity V800\xa0'] | [] | ['Evita V800', 'Evita V500']
1651 |  |  | SmartCare PEEP Limitation | mbar | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. PEEP in SmartCare - upper limit, SC-PEEPmax. For Evita Infinity V500, Babylog VN500: SC-Upper limit PEEP, PEEP max. | Double | Settings | ['Evita Infinity V500 [MEDIBUS.X]\xa0', 'Evita Infinity V800\xa0'] | [] | ['Evita V800', 'Evita V500']
1652 |  |  | SmartCare Spontaneous Respiration Rate Lower Limitation | 1/min | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. Spontaneous respiratory rate in SmartCare - lower limit, SC-RRspon low. For Evita Infinity V500, Babylog VN500: SC-Lower limit respiratory rate in the guideline, RRspon low. | Double | Settings | ['Evita Infinity V500 [MEDIBUS.X]\xa0', 'Evita Infinity V800\xa0'] | [] | ['Evita V800', 'Evita V500']
1653 |  |  | SmartCare Spontaneous Respiration Rate Upper Limitation | 1/min | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. Spontaneous respiratory rate in SmartCare - upper limit, SC-RRspon high. For Evita Infinity V500, Babylog VN500: SC-Upper limit respiratory rate in the guideline, RRspon high. | Double | Settings | ['Evita Infinity V500 [MEDIBUS.X]\xa0', 'Evita Infinity V800\xa0'] | [] | ['Evita V800', 'Evita V500']
1654 |  |  | SmartCare Tidal Volume Limitation | mL/kg | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. Tidal volume in SmartCare - lower limit, SC-VTlow. For Evita Infinity V500, Babylog VN500: SC-Lower limit tidal volume, VT low. | Double | Settings | ['Evita Infinity V500 [MEDIBUS.X]\xa0', 'Evita Infinity V800\xa0'] | [] | ['Evita V800', 'Evita V500']
1655 |  |  | SmartCare EtCO2 Limitation | mmHg | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. End-tidal CO2 in SmartCare - upper limit, SC-etCO2high. For Evita Infinity V500, Babylog VN500: SC-Upper limit endtidal CO2 in the guideline etCO2 high. | Double | Settings | ['Evita Infinity V500 [MEDIBUS.X]\xa0', 'Evita Infinity V800\xa0'] | [] | ['Evita V800', 'Evita V500']
1656 |  |  | SmartCare Patient Height | cm | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. Patient height in SmartCare, SC-Ht. For Evita Infinity V500: SC-Body height. | Double | Settings | ['Evita Infinity V500 [MEDIBUS.X]\xa0', 'Evita Infinity V800\xa0'] | [] | ['Evita V800', 'Evita V500']
1703 |  |  | Abdominal TOCO | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | aToco | ['MP70', 'MX800', 'X3']
1704 |  |  | Chest Movement Respiration Rate | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | cmRR | ['MP70', 'MX800', 'X3']
1705 |  |  | Patient Movement Respiration Rate | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | mRR | ['MP70', 'MX800', 'X3']
1706 |  |  | Patient Movement Pulse Rate | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | mPulse | ['MP70', 'MX800', 'X3']
1707 |  |  | Respiration Signal Quality Index (Movement) | - | [] | Double | OtherMeasuredData | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | mSQI | ['MP70', 'MX800', 'X3']
1715 |  |  | Central Venous Pressure (Thermodilution) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | iCVP | ['MP70', 'MX800', 'X3']
1716 |  |  | Mean Arterial Pressure (Thermodilution) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | iARTm | ['MP70', 'MX800', 'X3']
1718 |  |  | Average Blood Temperature (Thermodilution) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | iAvgBT | ['MP70', 'MX800', 'X3']
1723 |  |  | O2 Delivery DO2 (Thermodilution) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | iDO2 | ['MP70', 'MX800', 'X3']
1724 |  |  | O2 Consumption (Thermodilution) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | iVO2 | ['MP70', 'MX800', 'X3']
1725 |  |  | O2 Consumption Index (Thermodilution) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | iVO2I | ['MP70', 'MX800', 'X3']
1728 |  |  | Ultrafiltration/Blood Flow Ratio | % | [] | Long | VitalSigns | Multifiltrate PRO | Dialysis Course | Multifiltrate PRO
1737 |  |  | Stroke Volume (Thermodilution) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | iSV | ['MP70', 'MX800', 'X3']
1738 |  |  | Stroke Volume Index (Thermodilution) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | iSI | ['MP70', 'MX800', 'X3']
1739 |  |  | O2 Consumption Estimate | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | iVO2e | ['MP70', 'MX800', 'X3']
1745 |  |  | EEG1 Wave | - | [] | Double | Waveforms | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Waves: EEG1 | ['MP70', 'MX800', 'X3']
1746 |  |  | EEG2 Wave | - | [] | Double | Waveforms | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Waves: EEG2 | ['MP70', 'MX800', 'X3']
1747 |  |  | EEG3 Wave | - | [] | Double | Waveforms | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Waves: EEG3 | ['MP70', 'MX800', 'X3']
1748 |  |  | EEG4 Wave | - | [] | Double | Waveforms | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Waves: EEG4 | ['MP70', 'MX800', 'X3']
1751 |  |  | Spectral Edge Frequency | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | EEG: SEF | ['MP70', 'MX800', 'X3']
1757 |  |  | Burst Suppression Ratio | - | This variable is associated with ChannelID identifier. | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | BSR1/BSR2/BSR3/BSR4 | ['MP70', 'MX800', 'X3']
1759 |  |  | Patient Energy Expenditure | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | EE | ['MP70', 'MX800', 'X3']
1760 |  |  | Patient Respiratory Quotient | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | RQ | ['MP70', 'MX800', 'X3']
1782 |  |  | Spectral Edge Frequency (CH3) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | EEG: Frequ3 | ['MP70', 'MX800', 'X3']
1784 |  |  | Relative Power In Delta Band (CH3) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Delta3 | ['MP70', 'MX800', 'X3']
1785 |  |  | Relative Power In Theta Band (CH3) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Theta3 | ['MP70', 'MX800', 'X3']
1786 |  |  | Relative Power In Alpha Band (CH3) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Alpha3 | ['MP70', 'MX800', 'X3']
1787 |  |  | Relative Power In Beta Band (CH3) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Beta3 | ['MP70', 'MX800', 'X3']
1790 |  |  | Spectral Edge Frequency (CH4) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | SEF4 | ['MP70', 'MX800', 'X3']
1792 |  |  | Relative Power In Delta Band (CH4) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Delta4 | ['MP70', 'MX800', 'X3']
1793 |  |  | Relative Power In Theta Band (CH4) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Theta4 | ['MP70', 'MX800', 'X3']
1794 |  |  | Relative Power In Alpha (CH4) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Alpha4 | ['MP70', 'MX800', 'X3']
1795 |  |  | Relative Power In Beta Band (CH4) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Beta4 | ['MP70', 'MX800', 'X3']
1803 |  |  | Device Alarm Global Status | unitless | Format: Bit 2: 0 --> Alarm | String | Alarms | Common | [] | 5008
1822 |  |  | Blood Pump Alarm | unitless | Value of "1" when alarm is active. | Boolean | Alarms | Common | [] | 5008
1831 |  |  | Bed Head Angle Alarm | unitless | The current status for the head angle alarm system. | Byte | Alarms | Bed Data Server (Common) | BED | Total Scale
1842 |  |  | Arterial Pressure Upper Alarm Limit | mmHg | This variable is also sent as alarm limit of variableID 7438. | Long | Settings | Multifiltrate PRO | Dialysis Settings | Multifiltrate PRO
1843 |  |  | Arterial Pressure Lower Alarm Limit | mmHg | This variable is also sent as alarm limit of variableID 7438. | Long | Settings | Multifiltrate PRO | Dialysis Settings | Multifiltrate PRO
1911 |  |  | Device Unique ID | unitless | [] | String | DeviceInfo | Common | System Information | ['MP20', 'MP70', 'MX800', 'X3']
1914 |  |  | Patient Bed ID | unitless | This variable contains a printable string identifying the system location. | String | PatientInfo | Common | System Information | ['MP20', 'MP70', 'MX800', 'X3']
1927 |  |  | Instrument Code | unitless | Format: 6 characters string | String | DeviceInfo | Common | [] | Applix Smart
1931 |  |  | Patient Location ID | unitless | The globally unique identifier for the Location entity associated with this Bed entity or an empty string if no association currently exists. | String | PatientInfo | Bed Data Server (Common) | BED | Total Scale
1932 |  |  | Patient Location Text | unitless | Easily recognizable location name that can be used by authorized personnel to locate the specific physical location (e.g. Tower1_WingC_Floor2_Room103). | String | PatientInfo | Bed Data Server (Common) | LOCATION | Total Scale
1942 |  |  | Patient Comment Text | unitless | [] | String (Multiset) | PatientInfo | Common | Patient Demographics | ['MP20', 'MP70', 'MX800', 'X3']
1946 |  |  | Remote Alarm Application Mode | unitless | The current configuration mode of the SafeView alarm system. | String | DeviceInfo | Bed Data Server (Common) | BED | Total Scale
1947 |  |  | Remote Alarm Application Indicator Status | unitless | The current status of the SafeView alarm system indicators. | String | DeviceInfo | Bed Data Server (Common) | BED | Total Scale
1948 |  |  | Patient Weight Scale Activity | unitless | The last command to be successfully processed by the weigh scale system. | String | DeviceInfo | Bed Data Server (Common) | BED | Total Scale
1955 |  |  | Patient Stress Index | - | [] | Double | PatientInfo | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | StrssI | ['MP70', 'MX800', 'X3']
1957 |  |  | Remote CIS Application Status | unitless | The current status for the NaviCare alerts switch. | String | DeviceInfo | Bed Data Server (Common) | BED | Total Scale
1958 |  |  | Remote CIS Application Indicator Status | unitless | The current status for the NaviCare alerts indicator. | String | DeviceInfo | Bed Data Server (Common) | BED | Total Scale
1963 |  |  | Nurse Call Indicator Status | unitless | The current status for the nurse call indicator. | String | DeviceInfo | Bed Data Server (Common) | BED | Total Scale
1964 |  |  | Nurse Answer Indicator Status | unitless | The current status for the nurse answer indicator. | String | DeviceInfo | Bed Data Server (Common) | BED | Total Scale
1982 |  |  | Generic Patient Temperature | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Tbody | ['MP70', 'MX800', 'X3']
1986 |  |  | Bed Height Position | unitless | The bed’s physical articulated position. | String | DeviceInfo | Bed Data Server (Common) | BED | Total Scale
1987 |  |  | Bed Lockout Status | unitless | [] | String | DeviceInfo | Bed Data Server (Common) | BED | Total Scale
1988 |  |  | Bed Head Rails Position | unitless | The physical positions of the head-end siderails. | String | DeviceInfo | Bed Data Server (Common) | BED | Total Scale
1989 |  |  | Bed Foot RailsPosition | unitless | The physical positions of the foot-end siderails. | String | DeviceInfo | Bed Data Server (Common) | BED | Total Scale
1990 |  |  | Bed Surface Support Mode | unitless | The current configuration mode of the patient support surface system. | String | DeviceInfo | Bed Data Server (Common) | BED | Total Scale
1991 |  |  | Bed Emergency CPR Status | unitless | The current status for the CPR switch. | String | DeviceInfo | Bed Data Server (Common) | BED | Total Scale
1992 |  |  | Bed Clearance Status | unitless | The current status for the Bed Cleaned switch. | String | DeviceInfo | Bed Data Server (Common) | BED | Total Scale
1993 |  |  | Bed Rotation Therapy Status | unitless | The current status for the rotation therapy system. | String | DeviceInfo | Bed Data Server (Common) | BED | Total Scale
1994 |  |  | Bed Percussion Therapy Status | unitless | The current status for the percussion therapy system. | String | DeviceInfo | Bed Data Server (Common) | BED | Total Scale
1995 |  |  | Bed Vibration Therapy Status | unitless | The current status for the vibration therapy system. | String | DeviceInfo | Bed Data Server (Common) | BED | Total Scale
1996 |  |  | Patient Environment Control Activity | unitless | The last command completed by the patient environment control system. | String | DeviceInfo | Bed Data Server (Common) | BED | Total Scale
1997 |  |  | Patient Environment Activity Cleared Status | unitless | The last command completed by the patient environment control system. | Byte | DeviceInfo | Bed Data Server (Common) | BED | Total Scale
1998 |  |  | Bed Head Angle Alarm Mode | unitless | The current mode configuration for the head angle alarm system. | String | DeviceInfo | Bed Data Server (Common) | BED | Total Scale
1999 |  |  | Bed Head Angle Alarm Silence Status | unitless | The current mode configuration for the audible annunciation portion of the head angle alarm system. | String | DeviceInfo | Bed Data Server (Common) | BED | Total Scale
2000 |  |  | Bed Clearance Indicator Status | unitless | The current status for the bed cleaned indicator. | String | DeviceInfo | Bed Data Server (Common) | BED | Total Scale
2001 |  |  | SmartBed Patient Movement State | unitless | The current direction of patient activity as detected by the patient positioning monitor (PPM) alarm system. | String | OtherMeasuredData | Bed Data Server (Common) | BED | Total Scale
2032 |  |  | Volume Not Delivered Alarm | unitless | Consult detailed help for more information on this variable | Long | Alarms | Common | [] | V60
2044 |  |  | P01 Repetition Interval Setting | min | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. P0.1 repetition interval. | Double | Settings | ['Evita Infinity V500 [MEDIBUS.X]\xa0', 'Evita Infinity V800\xa0'] | [] | ['Evita V800', 'Evita V500']
2049 |  |  | Breath Cycle Type | - | [] | Double | DeviceInfo | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sCycTy | ['MP70', 'MX800', 'X3']
2163 |  |  | Spectral Edge Frequency - 1st EEG | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | EEG: Frequ1 | ['MP70', 'MX800', 'X3']
2167 |  |  | Bispectral Index | - | This variable is associated with ChannelID identifier. | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | BIS L/BIS R | ['MP70', 'MX800', 'X3']
2168 |  |  | Total Power (Ch1) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | EEG: TP1 | ['MP70', 'MX800', 'X3']
2169 |  |  | Spike Rate | - | This variable is associated with ChannelID identifier. | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | BurstL/BurstR | ['MP70', 'MX800', 'X3']
2173 |  |  | Spectral Edge Frequency - 2nd EEG | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | EEG: Frequ2 | ['MP70', 'MX800', 'X3']
2178 |  |  | Total Power (Ch2) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | EEG: TP2 | ['MP70', 'MX800', 'X3']
2182 |  |  | Narcotrend Index | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | NI | ['MP70', 'MX800', 'X3']
2183 |  |  | Sharp Transient Intensity | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | STI | ['MP70', 'MX800', 'X3']
2188 |  |  | Total Power Relative | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | TPrel | ['MP70', 'MX800', 'X3']
2217 |  |  | Manual Breath Ventilation Pressure Setting | mbar | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300(added with SW 2.30 for Babylog VN500, added with SW 2.30 for Evita Infinity V500). Manual inspiratory pressure for Neonates during NIV - Non-Invasive Ventilation, PmanInsp. | Double | Settings | ['Evita Infinity V500 [MEDIBUS.X]\xa0', 'Evita Infinity V800\xa0'] | [] | ['Evita V800', 'Evita V500']
2245 |  |  | Manual Breath Maximum Time Limit | s | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300(added with SW 2.30 for Babylog VN500, added with SW 2.30 for Evita Infinity V500). Maximum time of pressure for manual inspiration pressure for Neonates during NIV - Non-Invasive Ventilation, TmanInsp. | Double | Settings | ['Evita Infinity V500 [MEDIBUS.X]\xa0', 'Evita Infinity V800\xa0'] | [] | ['Evita V800', 'Evita V500']
2246 |  |  | O2 Flush Percentage Setting | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sO2max | ['MP70', 'MX800', 'X3']
2250 |  |  | CO2 Tidal Production | ['mL', '-'] | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. CO2 production volume per breath, VTCO2. For Evita Infinity V500, Babylog VN500: Supported through Medibus.X protocol. | Double | VitalSigns | ['IntelliVue MX800\xa0', 'Evita Infinity V500 [MEDIBUS.X]\xa0', 'IntelliVue X3\xa0', 'Evita Infinity V800\xa0', 'IntelliVue MP70\xa0'] | VCO2ti | ['X3', 'MX800', 'Evita V800', 'MP70', 'Evita V500']
2251 |  |  | CO2 Minute Production | mL/min | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Evita V300. CO2 production, V'CO2. | Double | VitalSigns | ['Evita Infinity V500 [MEDIBUS.X]\xa0', 'Evita Infinity V800\xa0'] | [] | ['Evita V800', 'Evita V500']
2254 |  |  | Inspiratory Resistance | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Rinsp | ['MP70', 'MX800', 'X3']
2262 |  |  | Pulmonary Artery BP - Systolic | - | [] | Double | VitalSigns | Common | IBP: PAP | ['MP20', 'MP70', 'MX800', 'X3']
2263 |  |  | Pulmonary Artery BP - Diastolic | - | [] | Double | VitalSigns | Common | IBP: PAP | ['MP20', 'MP70', 'MX800', 'X3']
2264 |  |  | Pulmonary Artery BP Mean | - | [] | Double | VitalSigns | Common | IBP: PAP | ['MP20', 'MP70', 'MX800', 'X3']
2271 |  |  | Right Atrial Pressure (RAP) | - | Available on IntelliVue devices with firmware revision G or later. | Double | VitalSigns | Common | IBP: RAP | ['MP20', 'MP70', 'MX800', 'X3']
2272 |  |  | Left Atrial Pressure (LAP) | - | Available on IntelliVue devices with firmware revision G or later. | Double | VitalSigns | Common | IBP: LAP | ['MP20', 'MP70', 'MX800', 'X3']
2273 |  |  | Intra Cranial Pressure (ICP) | - | Available on IntelliVue devices with firmware revision G or later. | Double | VitalSigns | Common | IBP: ICP | ['MP20', 'MP70', 'MX800', 'X3']
2317 |  |  | Pulmonary Artery Wedge Pressure | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | IBP: PAWP | ['MP70', 'MX800', 'X3']
2321 |  |  | Generic NIBP Systolic Pressure | - | This variable is same as variable 2, except that the early estimated values are marked as NonSignificant in this variable. | Double | VitalSigns | Common | NBP | ['MP20', 'MP70', 'MX800', 'X3']
2322 |  |  | Generic NIBP Diastolic Pressure | - | This variable is same as variable 3, except that the early estimated values are marked as NonSignificant in this variable. | Double | VitalSigns | Common | NBP | ['MP20', 'MP70', 'MX800', 'X3']
2323 |  |  | Generic NIBP Mean Pressure | - | This variable is same as variable 4, except that the early estimated values are marked as NonSignificant in this variable. | Double | VitalSigns | Common | NBP | ['MP20', 'MP70', 'MX800', 'X3']
2343 |  |  | Inspired O2 | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | VENT: FIO2 | ['MP70', 'MX800', 'X3']
2350 |  |  | Inspired O2 Pressure | - | This variable is marked as obsolete because the label is incorrect. Please see variable 52523 for concentration and 52527 for partial pressure. | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | O2 | ['MP70', 'MX800', 'X3']
2352 |  |  | CO2 (Gas Analyzer) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | CO2 | ['MP70', 'MX800', 'X3']
2355 |  |  | End Tidal CO2 Percentage | % | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Evita V300, Fabius Family. This variable is associated with UpperAlarmLimit attribute, which is available for Perseus A500, Atlan, Evita Infinity V500, Babylog VN500, Savina 300, Primus, Primus IE, Apollo, Zeus IE, Oxylog 3000 plus, Evita V300. This variable is associated with LowerAlarmLimit attribute, which is available for Perseus A500, Atlan, Evita Infinity V500, Babylog VN500, Savina 300, Primus, Primus IE, Apollo, Zeus IE, Oxylog 3000 plus, Evita V300. end tidal CO2 etCO2 | Double | VitalSigns | ['Evita Infinity V500 [MEDIBUS.X]\xa0', 'Evita Infinity V800\xa0'] | [] | ['Evita V800', 'Evita V500']
2357 |  |  | End Tidal N2O | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | N2O | ['MP70', 'MX800', 'X3']
2362 |  |  | End Tidal Desflurane | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | DES | ['MP70', 'MX800', 'X3']
2364 |  |  | End Tidal Sevoflurane | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | SEV | ['MP70', 'MX800', 'X3']
2366 |  |  | End Tidal Halothane | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | HAL | ['MP70', 'MX800', 'X3']
2368 |  |  | End Tidal Enflurane | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | ENF | ['MP70', 'MX800', 'X3']
2370 |  |  | End Tidal Isoflurane | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | ISO | ['MP70', 'MX800', 'X3']
2377 |  |  | Fresh Gas O2 Flow Rate | L/min | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Evita V300. O2 therapy flow, Flow. For Evita Infinity V500, Babylog VN500: Flow O2 Therapy Flow. | Double | DeviceInfo | ['Evita Infinity V500 [MEDIBUS.X]\xa0', 'Evita Infinity V800\xa0'] | [] | ['Evita V800', 'Evita V500']
2378 |  |  | Fresh Gas N2O Flow | - | Mixer parameter. Setting of fresh gas N2O flow. | Double | DeviceInfo | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sfgN2O | ['MP70', 'MX800', 'X3']
2379 |  |  | Fresh Gas Air Flow | - | [] | Double | DeviceInfo | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | fgAir | ['MP70', 'MX800', 'X3']
2391 |  |  | Transcutaneous O2 Pressure | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | TCUT: tcpO2 | ['MP70', 'MX800', 'X3']
2392 |  |  | Transcutaneous CO2 Pressure | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | TCUT: tcpCO2 | ['MP70', 'MX800', 'X3']
2399 |  |  | Arterial O2 Saturation (SaO2) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
2408 |  |  | EEG Total Power | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | EEG: TP | ['MP70', 'MX800', 'X3']
2409 |  |  | Bispectral Index (BIS) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | EEG: BIS | ['MP70', 'MX800', 'X3']
2411 |  |  | EEG Burst Suppression Ratio | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | EEG: BSI | ['MP70', 'MX800', 'X3']
2420 |  |  | Spontaneous Alveolar Minute Volume | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | SpMVal | ['MP70', 'MX800', 'X3']
2421 |  |  | Airway Dead Space | ['mL', '-'] | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. Serial dead space volume, Vds. For Evita Infinity V500, Babylog VN500: Serial dead space Vds. | Double | VitalSigns | ['IntelliVue MX800\xa0', 'Evita Infinity V500 [MEDIBUS.X]\xa0', 'IntelliVue X3\xa0', 'Evita Infinity V800\xa0', 'IntelliVue MP70\xa0'] | awVd | ['X3', 'MX800', 'Evita V800', 'MP70', 'Evita V500']
2423 |  |  | Dynamic Compliance Over Last 20% Breath | unitless | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. Ratio of compliance during the last 20 % of inspiration over dynamic compliance, C20/Cdyn. For Evita Infinity V500, Babylog VN500: C20/Cdyn. | Double | VitalSigns | ['Evita Infinity V500 [MEDIBUS.X]\xa0', 'Evita Infinity V800\xa0'] | [] | ['Evita V800', 'Evita V500']
2440 |  |  | Total Alveolar Tidal Volume | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | TValv | ['MP70', 'MX800', 'X3']
2445 |  |  | Patient Rebreathing Alarm | unitless | Consult detailed help for more information on this variable | ['Boolean\xa0', 'Long\xa0'] | Alarms | ['Monnal T60\xa0', 'Common\xa0'] | [] | ['V60', 'Monnal T60']
2447 |  |  | Respiration Waveform | - | [] | Double | Waveforms | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Waves: cmResp | ['MP70', 'MX800', 'X3']
2456 |  |  | EEG Signal Quality Index | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | EEG: SQI | ['MP70', 'MX800', 'X3']
2467 |  |  | Patient Type ID | unitless | Consult detailed help for more information on this variable | UByte | PatientInfo | ['Evita Infinity V500 [MEDIBUS.X]\xa0', 'Evita Infinity V800\xa0'] | [] | ['Evita V800', 'Evita V500']
2477 |  |  | Battery Module Status | unitless | The current status of the battery power system. | String | DeviceInfo | Bed Data Server (Common) | BED | Total Scale
2486 |  |  | NIBP Interval Time | min | [] | Long | DeviceInfo | Common | [] | 5008
2499 |  |  | Relative Blood Volume Change Limit | % | [] | Long | Settings | Common | [] | 5008
2517 |  |  | Wearable Respiration Rate | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | wRR | ['MP70', 'MX800', 'X3']
2518 |  |  | Wearable Pulse Rate | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | wPulse | ['MP70', 'MX800', 'X3']
2519 |  |  | Wearable Patient Temperature | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | wTemp | ['MP70', 'MX800', 'X3']
2520 |  |  | Respiration Signal Quality Index (Wearable) | - | [] | Double | OtherMeasuredData | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | wSQI | ['MP70', 'MX800', 'X3']
2521 |  |  | Pulse Rate Signal Quality Index (Wearable) | - | [] | Double | OtherMeasuredData | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | wSQI | ['MP70', 'MX800', 'X3']
2537 |  |  | Citrate Calcium Activation Status | unitless | Possible values: | Boolean | DeviceInfo | Multifiltrate PRO | Dialysis Settings | Multifiltrate PRO
2539 |  |  | Therapy Pause Start Time | unitless | Format: YYYYMMDDhhmmss | String | OtherMeasuredData | Multifiltrate PRO | Treatment Status | Multifiltrate PRO
2543 |  |  | Citrate Blood Concentration Setting | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sCitDo | ['MP70', 'MX800', 'X3']
2547 |  |  | Neutrophil Count | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Bands | ['MP70', 'MX800', 'X3']
2568 |  |  | Airway Peak Pressure Alarm | unitless | [] | Boolean | Alarms | Common | [] | Monnal T60
2580 |  |  | Fresh Gas Total Flow Setting | - | Mixer parameter. | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sfgFl | ['MP70', 'MX800', 'X3']
2596 |  |  | Alarm Severity Level | unitless | This variable is in multiset format. This variable contains alarm priority level, variable 6164 contains alarm codes, and variable 3878 contains alarm messages. They need to be interpreted together. One alarm code may have one or more priority values and one or more alarm messages associated with it. The possible alarms are depended on device firmware version. Please see device user manuals for details. See variables 177 and 588 and 4131 for string format. | UByte (Multiset) | Alarms | ['Evita Infinity V500 [MEDIBUS.X]\xa0', 'Evita Infinity V800\xa0'] | Alarm And Text Message | ['Evita V800', 'Evita V500']
2598 |  |  | Patient Admission Status | unitless | Possible values: "empty", "pre-admitted", "admitted", and "discharged". | String | PatientInfo | Common | Patient Demographics | ['MP20', 'MP70', 'MX800', 'X3']
2602 |  |  | ECG Lead V | - | Physiological identifier: NOM_ECG_AMPL_ST_V | Double | VitalSigns | Common | ECG: ST | ['MP20', 'MP70', 'MX800', 'X3']
2738 |  |  | Plateau Pressure | - | This variable contains parameter from Philips label "NLS_NOM_VENT_PRESS_RESP_PLAT" and observed value "NLS_NOM_VENT_PRESS_RESP_PLAT". | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | VENT: PEinsp | ['MP70', 'MX800', 'X3']
2754 |  |  | Ventilator General Status | unitless | Ventilator Standby mode status. | UByte | DeviceInfo | ['Evita Infinity V500 [MEDIBUS.X]\xa0', 'Evita Infinity V800\xa0'] | [] | ['Evita V800', 'Evita V500']
2756 |  |  | Ventilation Mode ID | unitless | This variable contains ID of basic ventilation mode ID without options. See also variable 584. | UByte | DeviceInfo | ['Evita Infinity V500 [MEDIBUS.X]\xa0', 'Evita Infinity V800\xa0'] | [] | ['Evita V800', 'Evita V500']
2758 |  |  | Ventilator Operation Mode ID | unitless | For Evita 4, Evita 2 dura, Evita XL: Firmware version 04.00 or higher For Evita 2: Variable label is not correct. Please use variable 2467 instead. Obsolete variable that may be removed in future release. | UByte | DeviceInfo | ['Evita Infinity V500 [MEDIBUS.X]\xa0', 'Evita Infinity V800\xa0'] | [] | ['Evita V800', 'Evita V500']
2773 |  |  | I:E Ratio Setting | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sIE 1: | ['MP70', 'MX800', 'X3']
2787 |  |  | APRV Low Time (Tlow) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Tlow | ['MP70', 'MX800', 'X3']
2821 |  |  | System General Status | unitless | This variable identifies the current operating mode of the device. Possible values: "not specified", "monitoring", "demo", and "service". Moreover, "- standby" or "- config" can be appended | String | DeviceInfo | Common | System Information | ['MP20', 'MP70', 'MX800', 'X3']
2845 |  |  | Frequency | Hz | This variable describes the frequency of the main power supply. Possible values: "unspecified", "50", and "60". | String | DeviceInfo | Common | System Information | ['MP20', 'MP70', 'MX800', 'X3']
2847 |  |  | Urea Distribution Volume Setting | L | [] | Double | Settings | Common | [] | 5008
2848 |  |  | Auto Ultrafiltration Control Activation Status | unitless | Possible values: 0=off, 1=on | Boolean | Settings | Common | [] | 5008
2855 |  |  | Device Status | unitless | ['This variable indicated if the water flow is through the patient wrap (garment) or is circulating internally (bypass). Possible values: 0 = Bypass 1 = Garment\xa0', 'Consult detailed help for more information on this variable\xa0', 'Possible values: "disconnected", "unassociated", "operating".\xa0'] | ['UShort\xa0', 'Boolean\xa0', 'String\xa0'] | DeviceInfo | ['Common\xa0', 'Multifiltrate PRO\xa0'] | ['System Information\xa0', 'Treatment Status\xa0'] | ['MP70', 'X3', 'MX800', 'MP20', 'Multifiltrate PRO', 'Criticool']
2880 |  |  | Maximum Apnea Time Setting | - | [] | Double | DeviceInfo | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sTAmax | ['MP70', 'MX800', 'X3']
2889 |  |  | Inspired Agent Concentration | ['%', '-'] | This Variable is associated with a SourceID identifier. | Double | VitalSigns | ['IntelliVue X3\xa0', 'Common\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | ['Agent\xa0', 'inAGT\xa0'] | ['MP70', 'Expression IP5 et MR400', 'MX800', 'X3']
2890 |  |  | Expired Agent Concentration | % | This Variable is associated with a SourceID identifier. | Double | VitalSigns | Common | Agent | Expression IP5 et MR400
2901 |  |  | Patient First Name | unitless | [] | String | PatientInfo | Common | Patient Demographics | ['X3', 'MX800', 'MP20', '5008', 'MP70']
2916 |  |  | End Diastolic Volume | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | EDV | ['MP70', 'MX800', 'X3']
2917 |  |  | End Diastolic Volume Index | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | EDVI | ['MP70', 'MX800', 'X3']
2918 |  |  | End Systolic Volume | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | ESV | ['MP70', 'MX800', 'X3']
2919 |  |  | End Systolic Volume Index | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | ESVI | ['MP70', 'MX800', 'X3']
2922 |  |  | EEG Electromyography (EMG) | - | This variable is associated with ChannelID identifier. | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | EMG L/EMG R | ['MP70', 'MX800', 'X3']
2923 |  |  | Generic EEG Wave | - | [] | Double | Waveforms | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Waves: EEG | ['MP70', 'MX800', 'X3']
2924 |  |  | ECG Pad Wave | - | [] | Double | Waveforms | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Waves: Pads | ['MP70', 'MX800', 'X3']
2925 |  |  | ECG Paddle Wave | - | [] | Double | Waveforms | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Waves: Paddle | ['MP70', 'MX800', 'X3']
2930 |  |  | Unspecific Pressure Pulse Rate | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Pulse | ['MP70', 'MX800', 'X3']
2932 |  |  | Patient Temperature 1 | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
2933 |  |  | Patient Temperature 2 | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
2952 |  |  | PS Activation State | unitless | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300, Fabius Family. A value of 1 indicates that Pressure Support (PS) is added to the current ventilation mode. A value of 0 indicates that Pressure Support (PS) is not added to the current ventilation mode. Text message displayed on device screen: "/PS" | Boolean | DeviceInfo | ['Evita Infinity V500 [MEDIBUS.X]\xa0', 'Evita Infinity V800\xa0'] | [] | ['Evita V800', 'Evita V500']
2953 |  |  | VS Activation State | unitless | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. A value of 1 indicates that Volume support (VS) is added to the current ventilation mode. A value of 0 indicates that Volume support (VS) is not added to the current ventilation mode. Text messages displayed on device screen: "/VS" | Boolean | DeviceInfo | ['Evita Infinity V500 [MEDIBUS.X]\xa0', 'Evita Infinity V800\xa0'] | [] | ['Evita V800', 'Evita V500']
2954 |  |  | APRV High Pressure (Phigh) | ['mbar', '-'] | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. Upper pressure level during APRV, Phigh. | Double | VitalSigns | ['IntelliVue MX800\xa0', 'Evita Infinity V500 [MEDIBUS.X]\xa0', 'IntelliVue X3\xa0', 'Evita Infinity V800\xa0', 'IntelliVue MP70\xa0'] | Phigh | ['X3', 'MX800', 'Evita V800', 'MP70', 'Evita V500']
2959 |  |  | APRV Low Pressure (Plow) | ['mbar', '-'] | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. Lower pressure level during APRV, Plow. | Double | VitalSigns | ['IntelliVue MX800\xa0', 'Evita Infinity V500 [MEDIBUS.X]\xa0', 'IntelliVue X3\xa0', 'Evita Infinity V800\xa0', 'IntelliVue MP70\xa0'] | Plow | ['X3', 'MX800', 'Evita V800', 'MP70', 'Evita V500']
2961 |  |  | Therapy Pause End Time | unitless | Format: YYYYMMDDhhmmss | String | OtherMeasuredData | Multifiltrate PRO | Treatment Status | Multifiltrate PRO
2969 |  |  | Disposable Lifetime | min | The resolution of this variable is 10 min. This variable is sent from the device to the DDI on a regular basis, but the device will keep on sending the same value until a minimal change of 10 min is observed. | Long | DeviceInfo | Multifiltrate PRO | Dialysis Course | Multifiltrate PRO
2984 |  |  | Plateau Pressure Setting | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sPplat | ['MP70', 'MX800', 'X3']
2991 |  |  | FiO2 Upper Alarm Limit | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | AlarmLimit: highO2 | ['MP70', 'MX800', 'X3']
2992 |  |  | FiO2 Lower Alarm Limit | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | AlarmLimit: lowO2 | ['MP70', 'MX800', 'X3']
2995 |  |  | Device Altitude | m | [] | Double | DeviceInfo | Common | System Information | ['MP20', 'MP70', 'MX800', 'X3']
3016 |  |  | Transmembrane Pressure Upper Alarm Limit | mmHg | This variable is also sent as alarm limit of variableID 3029. | Long | Settings | Multifiltrate PRO | Dialysis Settings | Multifiltrate PRO
3062 |  |  | Hb Calculation Offset | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | OffsHb | ['MP70', 'MX800', 'X3']
3070 |  |  | Patient Profile | unitless | Na profile number | UByte | PatientInfo | Common | [] | 5008
3071 |  |  | Therapy Profile | unitless | [] | UByte | PatientInfo | Common | [] | 5008
3091 |  |  | Patient Time Since Last Ward Admission | - | [] | Double | PatientInfo | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | TWrdAd | ['MP70', 'MX800', 'X3']
3092 |  |  | Patient Total Admission Time | - | [] | Double | PatientInfo | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | WardAd | ['MP70', 'MX800', 'X3']
3100 |  |  | Dialysate Sodium Start Concentration Setting | mmol/L | [] | Double | Settings | Common | [] | 5008
3101 |  |  | Patient Drainage Volume | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | DrnVol | ['MP70', 'MX800', 'X3']
3147 |  |  | Maximum Lung Elastance Setting | - | Possible units: cmH2O/L, mBar/L | Double | Settings | Common | [] | V60
3189 |  |  | Average Replacement Flow Rate | mL/min | [] | Long | VitalSigns | Common | [] | 5008
3190 |  |  | Minimum Relative Treated Blood Volume | % | [] | Double | VitalSigns | Common | [] | 5008
3194 |  |  | Auto Replacement Control Activation Status | unitless | Possible values: 0=disabled, 1=enabled | Boolean | Settings | Common | [] | 5008
3203 |  |  | Maximum Airway Resistance Setting | - | Possible units: cmH2O/L/s, mBar/L/s | Double | Settings | Common | [] | V60
3204 |  |  | Sum of Minimum Alveolar Concentration (MAC) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | MAC | ['MP70', 'MX800', 'X3']
3236 |  |  | Relative Blood Volume Change | % | [] | Double | VitalSigns | Common | [] | Artis
3237 |  |  | Relative Blood Volume Change | % | [] | Double | VitalSigns | Common | [] | 5008
3239 |  |  | Replacement Volume Setting | L | [] | Double | Settings | Common | [] | 5008
3253 |  |  | Blood Circuit Pressure Alarm | unitless | Value of "1" when alarm is active. | Boolean | Alarms | Common | [] | 5008
3263 |  |  | Recirculation Percentage | % | [] | Double | VitalSigns | Common | [] | 5008
3266 |  |  | HFV Frequency Setting | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sHFFrq | ['MP70', 'MX800', 'X3']
3268 |  |  | Ventilator Base Flow Setting | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sfgFmn | ['MP70', 'MX800', 'X3']
3269 |  |  | Filtration Ratio | % | [] | UShort | VitalSigns | Common | [] | 5008
3271 |  |  | Blood Pressure Difference | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | IBP: DeltaP | ['MP70', 'MX800', 'X3']
3272 |  |  | Generic Blood Pressure Difference 1 | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | IBP: DeltaP1 | ['MP70', 'MX800', 'X3']
3273 |  |  | Generic Blood Pressure Difference 2 | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | IBP: DeltaP2 | ['MP70', 'MX800', 'X3']
3278 |  |  | Patient Height | - | [] | Double | PatientInfo | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Height | ['MP70', 'MX800', 'X3']
3385 |  |  | Arterial O2 Content (CaO2) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | CaO2 | ['MP70', 'MX800', 'X3']
3391 |  |  | Alveolar Arterial O2 Difference (AaDO2) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | AaDO2 | ['MP70', 'MX800', 'X3']
3400 |  |  | Venous O2 Content (CvO2) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | CvO2 | ['MP70', 'MX800', 'X3']
3415 |  |  | Ionized Calcium Concentration (Ca) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | iCa | ['MP70', 'MX800', 'X3']
3417 |  |  | Ultrafiltration Remaining Time | min | Remaining ISO UF time. | Long | VitalSigns | Common | [] | 5008
3421 |  |  | Chloride Concentration (Cl-) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Cl | ['MP70', 'MX800', 'X3']
3430 |  |  | Total Replacement Flow Rate | mL/min | [] | Long | VitalSigns | Common | [] | 5008
3435 |  |  | Activated Clotting Time | - | Measured by the i-Stat module. | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | ACT | ['MP70', 'MX800', 'X3']
3446 |  |  | Generic Temperature | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | T1 | ['MP70', 'MX800', 'X3']
3447 |  |  | Generic Temperature 2 | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | T2 | ['MP70', 'MX800', 'X3']
3448 |  |  | Generic Temperature 3 | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | T3 | ['MP70', 'MX800', 'X3']
3449 |  |  | Generic Temperature 4 | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | T4 | ['MP70', 'MX800', 'X3']
3479 |  |  | Equivalent Dialysate Sodium (Na) | mmol/L | [] | Long | VitalSigns | Common | [] | Artis
3490 |  |  | TcGas Site Timer | - | [] | Double | OtherMeasuredData | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | TCUT: SitTim | ['MP70', 'MX800', 'X3']
3491 |  |  | O2 Saturation (SO2) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | SaO2 | ['MP70', 'MX800', 'X3']
3493 |  |  | Desired Equivalent Dialysate Sodium (Na) | mmol/L | [] | Long | Settings | Common | [] | Artis
3496 |  |  | TcGas Transducer Temperature | - | [] | Double | DeviceInfo | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | TCUT: SensrT | ['MP70', 'MX800', 'X3']
3502 |  |  | Peak Expiratory Flow Setting | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | s%PEF | ['MP70', 'MX800', 'X3']
3503 |  |  | Auto Flow Active Status | ['unitless', '-'] | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300. A value of 1 indicates that AutoFlow (AF) is added to the current ventilation mode. A value of 0 indicates that AutoFlow (AF) is not added to the current ventilation mode. Text message displayed on device screen: "/AF" | ['Double\xa0', 'Boolean\xa0'] | DeviceInfo | ['IntelliVue MX800\xa0', 'Evita Infinity V500 [MEDIBUS.X]\xa0', 'IntelliVue X3\xa0', 'Evita Infinity V800\xa0', 'IntelliVue MP70\xa0'] | sFlCtr | ['X3', 'MX800', 'Evita V800', 'MP70', 'Evita V500']
3506 |  |  | Inspiratory Minimum CO2 | - | [] | Double | VitalSigns | Common | CO2 | ['MP20', 'MP70', 'MX800', 'X3']
3507 |  |  | Expired Tidal Volume Indexed by Body Weight | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | TVe/kg | ['MP70', 'MX800', 'X3']
3555 |  |  | ATC Activation Status | ['unitless', '-'] | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Evita V300. A value of 1 indicates that automatic tube compensation (ATC) is added to the current ventilation mode. A value of 0 indicates that automatic tube compensation (ATC) is not added to the current ventilation mode. Text message displayed on device screen: "/ATC" For Savina 300 (MEDIBUS.X): added with SW 5.00 | ['Double\xa0', 'Boolean\xa0'] | DeviceInfo | ['IntelliVue MX800\xa0', 'Evita Infinity V500 [MEDIBUS.X]\xa0', 'IntelliVue X3\xa0', 'Evita Infinity V800\xa0', 'IntelliVue MP70\xa0'] | sTubeC | ['X3', 'MX800', 'Evita V800', 'MP70', 'Evita V500']
3570 |  |  | Skin Temperature | - | [] | Double | VitalSigns | Common | Tskin | ['MP20', 'MP70', 'MX800', 'X3']
3571 |  |  | Rectal Temperature | - | [] | Double | VitalSigns | Common | Trect | ['MP20', 'MP70', 'MX800', 'X3']
3572 |  |  | Nasopharyngeal Temperature | - | [] | Double | VitalSigns | Common | Tnaso | ['MP20', 'MP70', 'MX800', 'X3']
3573 |  |  | Esophageal Temperature | - | [] | Double | VitalSigns | Common | Tesoph | ['MP20', 'MP70', 'MX800', 'X3']
3574 |  |  | Core Temperature | - | [] | Double | VitalSigns | Common | Tcore | ['MP20', 'MP70', 'MX800', 'X3']
3575 |  |  | Arterial Temperature | ['°C', '-'] | [] | Double | VitalSigns | Common | Tart | ['X3', 'MX800', 'MP20', '5008', 'MP70']
3576 |  |  | Venous Temperature | ['°C', '-'] | [] | Double | VitalSigns | Common | Tven | ['X3', 'MX800', 'MP20', '5008', 'MP70']
3583 |  |  | Generic Core Temperature 1 | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | T1Core | ['MP70', 'MX800', 'X3']
3584 |  |  | Generic Core Temperature 2 | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | T2Core | ['MP70', 'MX800', 'X3']
3590 |  |  | Device Type | unitless | [] | String | DeviceInfo | Common | [] | 5008
3591 |  |  | Device Number | unitless | ['Device Model Number (example: "M8000")\xa0', 'Consult detailed help for more information on this variable\xa0'] | String | DeviceInfo | Common | ['Device Information\xa0', 'System Information\xa0'] | ['X3', 'MX800', 'MP20', 'Evita V800', 'MP70', 'Evita V500']
3592 |  |  | Low Airway Flow Alarm | unitless | Consult detailed help for more information on this variable | Long | Alarms | V60 (PVOI) | [] | V60
3615 |  |  | Battery Charge | unitless | 58112 => battery charged | UShort | DeviceInfo | Common | [] | Applix Smart
3626 |  |  | Therapy Ending Time | unitless | Format: YYYYMMDDhhmm | String | OtherMeasuredData | ['Common\xa0', 'Multifiltrate PRO\xa0'] | Treatment Status | ['Multifiltrate PRO', '5008']
3627 |  |  | Venous Pump Flow | mL/min | [] | Long | VitalSigns | Common | [] | Artis
3635 |  |  | Drug Delivery Rate | - | [] | Double | DeviceInfo | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sDlvRt | ['MP70', 'MX800', 'X3']
3637 |  |  | Drug Concentration | - | [] | Double | PatientInfo | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sDConc | ['MP70', 'MX800', 'X3']
3639 |  |  | Manufacturers Name | unitless | "Philips" | String | DeviceInfo | Common | System Information | ['MP20', 'MP70', 'MX800', 'X3']
3640 |  |  | Pressure Alarm Delay | - | [] | Double | DeviceInfo | ['IntelliVue X3\xa0', 'Common\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sPADel | ['V60', 'MP70', 'MX800', 'X3']
3646 |  |  | Arterial Blood Pressure (ART) | - | Available on IntelliVue devices with firmware revision G or later. Arterial Blood Pressure (NOM_PRESS_BLD_ART). | Double | VitalSigns | Common | IBP: ART | ['MP20', 'MP70', 'MX800', 'X3']
3647 |  |  | Pulmonary Artery Blood Pressure (PA) | - | Available on IntelliVue devices with firmware revision G or later. | Double | VitalSigns | Common | IBP: PAP | ['MP20', 'MP70', 'MX800', 'X3']
3704 |  |  | Arterial/Main Pump Flow | ['-', 'mL/min'] | [] | ['Double\xa0', 'Long\xa0'] | VitalSigns | ['IntelliVue X3\xa0', 'Common\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | HLMfl | ['MP70', 'Artis', 'MX800', 'X3']
3710 |  |  | Slave Pump Flow | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | SlvPfl | ['MP70', 'MX800', 'X3']
3715 |  |  | Suction Pump Flow | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | SucPfl | ['MP70', 'MX800', 'X3']
3720 |  |  | Auxiliary Pump Flow | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | AuxPfl | ['MP70', 'MX800', 'X3']
3723 |  |  | Auxiliary Pump Time Since Start of Current Plegia Session | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | AxOnTi | ['MP70', 'MX800', 'X3']
3724 |  |  | Auxiliary Pump Time Since Stop of Plegia Session | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | AxOffT | ['MP70', 'MX800', 'X3']
3725 |  |  | Auxiliary Pump Delivered Volume in Current Plegia Session | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | AxDVol | ['MP70', 'MX800', 'X3']
3726 |  |  | Auxiliary Pump Total Volume | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | AxTVol | ['MP70', 'MX800', 'X3']
3727 |  |  | Auxiliary Pump Total Plegia Time | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | AxPlTi | ['MP70', 'MX800', 'X3']
3730 |  |  | Cardioplegia Main Pump Flow | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | PlePfl | ['MP70', 'MX800', 'X3']
3733 |  |  | Cardioplegia Main Pump Delivered Volume of Current Plegia Session | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | CpDVol | ['MP70', 'MX800', 'X3']
3734 |  |  | Cardioplegia Main Pump Total Volume | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | CpTVol | ['MP70', 'MX800', 'X3']
3737 |  |  | Cardioplegia Slave Pump Flow | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | SplPfl | ['MP70', 'MX800', 'X3']
3740 |  |  | Cardioplegia Slave Pump Delivered Volume in Current Plegia Session | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | CsDVol | ['MP70', 'MX800', 'X3']
3741 |  |  | Cardioplegia Slave Pump Total Volume | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | CsTVol | ['MP70', 'MX800', 'X3']
3742 |  |  | Cycle Flow Sensitivity Setting | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sFlSeC | ['MP70', 'MX800', 'X3']
3761 |  |  | Hydrocarbon Plasma Concentration (HCO3) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | HCO3 | ['MP70', 'MX800', 'X3']
3781 |  |  | Corpuscular Mean Hemoglobin Concentration | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | MCHC | ['MP70', 'MX800', 'X3']
3785 |  |  | Arterial O2 Partial Pressure | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | PaO2 | ['MP70', 'MX800', 'X3']
3786 |  |  | Venous O2 Partial Pressure | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | PvO2 | ['MP70', 'MX800', 'X3']
3787 |  |  | Arterial CO2 Partial Pressure | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | PaCO2 | ['MP70', 'MX800', 'X3']
3793 |  |  | Hematocrit (HCT) | - | Post-Dialysis Hematocrit | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | cHCTo | ['MP70', 'MX800', 'X3']
3804 |  |  | Respiratory Rate (Flow) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | RRawf | ['MP70', 'MX800', 'X3']
3805 |  |  | Respiratory Rate (CO2) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | RRCO2 | ['MP70', 'MX800', 'X3']
3833 |  |  | Device Protocol Mode | unitless | The connectivity protocol supported by the bed communication controller. | String | DeviceInfo | Bed Data Server (Common) | BED | Total Scale
3849 |  |  | Alarm History | unitless | 16 last alarms | UShort (Multiset) | Alarms | Common | [] | Applix Smart
3854 |  |  | Patient Pace Maker Status | unitless | Possible values: "not paced", "paced gen". | String | PatientInfo | Common | Patient Demographics | ['MP20', 'MP70', 'MX800', 'X3']
3858 |  |  | Volume Not Achieved Alarm | unitless | [] | Boolean | Alarms | Common | [] | Monnal T60
3869 |  |  | Heart Rate Signal Quality | - | [] | Double | OtherMeasuredData | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | mSQI | ['MP70', 'MX800', 'X3']
3878 |  |  | Generic Alarm Message | unitless | This variable is in multiset format. This variable contains alarm messages, variable 2596 contains alarm priority level, variable 6164 contains alarm codes. They need to be interpreted together. One alarm code may have one or more priority values and one or more alarm messages associated with it. The possible alarms are dependent on device firmware version. Please see device user manuals for details. See variables 177 and 588 and 4131 for string format. The HL7 output can contain an '&' character to indicate a LOW alarm. In order to avoid issues, the '&' character will be replaced by a '?' character in the DDI output. | String (Multiset) | Alarms | ['Evita Infinity V500 [MEDIBUS.X]\xa0', 'Evita Infinity V800\xa0'] | Alarm And Text Message | ['Evita V800', 'Evita V500']
3909 |  |  | Message Code | unitless | This variable, along with variable 1197 are marked as obsolete because they use either"¦¦" or "," in the array string as delimiter, which are often not compatible with customers' applications. Variables 8366 and 8696 should be used instead. | String | Messages | Evita Infinity V500 [MEDIBUS.X] | Alarm And Text Message | Evita V500
3917 |  |  | Tracheal Pressure | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Ptrach | ['MP70', 'MX800', 'X3']
3921 |  |  | Ventilator Leakage | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Lksys | ['MP70', 'MX800', 'X3']
3963 |  |  | Warmer Mattress Temperature | - | [] | Double | OtherMeasuredData | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Tmatt | ['MP70', 'MX800', 'X3']
3998 |  |  | Auto Release Activation Status | unitless | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. A value of 1 indicates that AutoRelease is added to the current ventilation mode. A value of 0 indicates that AutoRelease is not added to the current ventilation mode. Text messages displayed on device screen: "/AutoRelease" | Boolean | DeviceInfo | ['Evita Infinity V500 [MEDIBUS.X]\xa0', 'Evita Infinity V800\xa0'] | [] | ['Evita V800', 'Evita V500']
4006 |  |  | Inspired MAC Concentration | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | MAC | ['MP70', 'MX800', 'X3']
4008 |  |  | Inspiratory 2nd Agent | ['%', '-'] | This Variable is associated with a SourceID identifier. | Double | VitalSigns | ['IntelliVue X3\xa0', 'Expression IP5\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | ['inAGTs\xa0', 'Agent (EXPRESS Plus, Express Connect Only)\xa0'] | ['MP70', 'Expression IP5 et MR400', 'MX800', 'X3']
4009 |  |  | Expiratory 2nd Agent | ['%', '-'] | ['This is the device\'s measurement with label="NLS_NOM_EMFC_etAGTs" and physio_id="NOM_CONC_AWAY_AGENT_ET".\xa0', 'This Variable is associated with a SourceID identifier.\xa0'] | Double | VitalSigns | ['IntelliVue X3\xa0', 'Expression IP5\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | ['etAGTs\xa0', 'Agent (EXPRESS Plus, Express Connect Only)\xa0'] | ['MP70', 'Expression IP5 et MR400', 'MX800', 'X3']
4010 |  |  | Minute Volume Alarm Delay | s | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. Delay for minute volume high alarm, MVhigh delay. This variable is associated with a ChannelID identifier. | Double | Settings | ['Evita Infinity V500 [MEDIBUS.X]\xa0', 'Evita Infinity V800\xa0'] | [] | ['Evita V800', 'Evita V500']
4015 |  |  | Disconnection Alarm Delay Time | s | For Evita 4, Evita 2 dura, Evita XL: Firmware version 4.00 or higher. For Savina: Tdeconnect. Available since Device SW version 02.00. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Carina, Evita V300. Disconnection alarm time, Tdiscon. For Evita Infinity V500, Babylog VN500: Tdisconnect. For Carina: Disconnection detection timeout. | Double | Settings | ['Evita Infinity V500 [MEDIBUS.X]\xa0', 'Evita Infinity V800\xa0'] | [] | ['Evita V800', 'Evita V500']
4016 |  |  | Flow Acceleration Setting | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sFlacc | ['MP70', 'MX800', 'X3']
4017 |  |  | Cumulated Halothane Consumption | - | [] | Double | OtherMeasuredData | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | AccHAL | ['MP70', 'MX800', 'X3']
4018 |  |  | Cumulated Enflurane Consumption | - | [] | Double | OtherMeasuredData | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | AccENF | ['MP70', 'MX800', 'X3']
4019 |  |  | Cumulated Isoflurane Consumption | - | [] | Double | OtherMeasuredData | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | AccISO | ['MP70', 'MX800', 'X3']
4020 |  |  | Cumulated Desflurane Consumption | - | [] | Double | OtherMeasuredData | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | AccDes | ['MP70', 'MX800', 'X3']
4021 |  |  | Cumulated Sevoflurane Consumption | - | [] | Double | OtherMeasuredData | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | AccSEV | ['MP70', 'MX800', 'X3']
4026 |  |  | Air Humidity Setting | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sHum | ['MP70', 'MX800', 'X3']
4028 |  |  | Mattress Temperature Setting | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sTmatt | ['MP70', 'MX800', 'X3']
4030 |  |  | Expired Agent Setting | - | End-Tidal agent concentration | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sEtAGT | ['MP70', 'MX800', 'X3']
4037 |  |  | Inspired Desflurane Setting | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sInDES | ['MP70', 'MX800', 'X3']
4038 |  |  | Expired Desflurane Setting | - | End-Tidal desflurane concentration | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sEtDES | ['MP70', 'MX800', 'X3']
4039 |  |  | Inspired Sevoflurane Setting | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sInSEV | ['MP70', 'MX800', 'X3']
4040 |  |  | Expired Sevoflurane Setting | - | End-Tidal sevoflurane concentration | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sEtSEV | ['MP70', 'MX800', 'X3']
4044 |  |  | Medibus Revision | unitless | [] | String | DeviceInfo | Common | Device Information | ['Evita V800', 'Evita V500']
4045 |  |  | Device Revision | unitless | Format: mm.nn | String | DeviceInfo | Common | Device Information | ['Evita V800', 'Evita V500']
4079 |  |  | Fresh Gas Sevoflurane | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | fgSEV | ['MP70', 'MX800', 'X3']
4080 |  |  | Fresh Gas Desflurane | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | fgDES | ['MP70', 'MX800', 'X3']
4081 |  |  | Ambient Temperature | - | [] | Double | OtherMeasuredData | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Tamb | ['MP70', 'MX800', 'X3']
4082 |  |  | Ambient Pressure | - | [] | Double | OtherMeasuredData | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | PB | ['MP70', 'MX800', 'X3']
4084 |  |  | Pulse Rate ( Unspecified Source ) | - | [] | Double | VitalSigns | Common | Pulse | ['MP20', 'MP70', 'MX800', 'X3']
4089 |  |  | Central Venous Pressure Diastolic | - | [] | Double | VitalSigns | Common | IBP: CVP | ['MP20', 'MP70', 'MX800', 'X3']
4093 |  |  | Central Venous Mean Pressure | - | [] | Double | VitalSigns | Common | IBP: CVP | ['MP20', 'MP70', 'MX800', 'X3']
4094 |  |  | Central Venous Pressure Systolic | - | [] | Double | VitalSigns | Common | IBP: CVP | ['MP20', 'MP70', 'MX800', 'X3']
4120 |  |  | Expiratory Volume Wave | - | [] | Double | Waveforms | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Waves: AWVex | ['MP70', 'MX800', 'X3']
4131 |  |  | Alarm Priority | unitless | This variable is same as 2596 but sent as a string, which is an array containing all alarm priorities and separated by "," delimiter. This variable, along with variables 177 and 588, are marked as obsolete because they use either"¦¦" or "," in the array string as delimiter, which are often not compatible with customers' applications. Variables 6164, 3878, and 2596 should be used instead. | String | Alarms | Evita Infinity V500 [MEDIBUS.X] | Alarm And Text Message | Evita V500
4143 |  |  | Apnea Ventilation I:E Ratio Setting (Number) | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sBIE1: | ['MP70', 'MX800', 'X3']
4145 |  |  | Dynamic Plateau Airway Pressure | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | PplDyn | ['MP70', 'MX800', 'X3']
4150 |  |  | VG Activation Status | unitless | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Carina, Evita V300. A value of 1 indicates that Volume Guarantee (VG) is added to the current ventilation mode. A value of 0 indicates that Volume Guarantee (VG) is not added to the current ventilation mode. Text message displayed on device screen: "/VG" | Boolean | DeviceInfo | ['Evita Infinity V500 [MEDIBUS.X]\xa0', 'Evita Infinity V800\xa0'] | [] | ['Evita V800', 'Evita V500']
4158 |  |  | Transcutaneous Relative Heater Power | - | [] | Double | OtherMeasuredData | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | HeatPr | ['MP70', 'MX800', 'X3']
4171 |  |  | Spontaneous Respiratory Rate Setting | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sRRmax | ['MP70', 'MX800', 'X3']
4201 |  |  | ['Product Description', 'Product Descritpion'] | unitless | ['4 characters for product identification\xa0', 'Variables 764, 765 and 4201 must be interpreted conjointly.'] | ['String (Multiset)\xa0', 'String\xa0'] | DeviceInfo | Common | System Information | ['X3', 'Applix Smart', 'MX800', 'MP20', 'MP70']
4236 |  |  | Transmembrane Pressure Setting | mmHg | [] | Long | Settings | Common | [] | Artis
4237 |  |  | Desired Blood Volume Change Percentage | % | [] | Double | Settings | Common | [] | Artis
4238 |  |  | Single Needle Max Pressure Setting | mmHg | [] | Long | Settings | Common | [] | 5008
4239 |  |  | Single Needle Min Pressure Setting | mmHg | [] | Long | Settings | Common | [] | 5008
4240 |  |  | Plasma Sodium (Na) | mmol/L | For 6008 only: This variable will contain the Online Clearance Monitoring plasma sodium concentration when sodium management is being used. | ['Double\xa0', 'Long\xa0'] | VitalSigns | Common | [] | ['5008', 'Artis']
4241 |  |  | Dilution Ratio (Qf/Qb) | unitless | [] | Double | VitalSigns | Common | [] | Artis
4274 |  |  | Service Message | unitless | The current status of the service required monitor system. | String | Group not set yet | Bed Data Server (Common) | BED | Total Scale
4281 |  |  | Nurse Call Status | unitless | The current status for the nurse call switch. | String | DeviceInfo | Bed Data Server (Common) | BED | Total Scale
4352 |  |  | Base Pressure | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
4369 |  |  | PEEP/CPAP Setting | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | VENT: sPEEP | ['MP70', 'MX800', 'X3']
4370 |  |  | Sigh Rate Setting | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | VENT: sSghR | ['MP70', 'MX800', 'X3']
4375 |  |  | Data Averaging Interval Setting | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
4376 |  |  | Chart Interval Setting | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
4388 |  |  | Apnea Inspiratory Pressure Setting | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sIPApn/sAInP | ['MP70', 'MX800', 'X3']
4393 |  |  | Patient Hose Compliance | - | [] | Double | OtherMeasuredData | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Chos | ['MP70', 'MX800', 'X3']
4401 |  |  | Oximeter Averaging Interval Setting | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
4457 |  |  | System Alarm Status | unitless | Overall alert state of the system. | Double | Alarms | Common | Alarms | ['MP20', 'MP70', 'MX800', 'X3']
4458 |  |  | Patient Alarm Priority | unitless | Overall Patient alarm priority. This is the current highest priority value of currently active Patient Alarms. | Double | Alarms | Common | Alarms | ['MP20', 'MP70', 'MX800', 'X3']
4459 |  |  | Technical Alarm Priority | unitless | Overall Technical alarm priority. This is the current highest priority value of currently active Technical Alarms. | Double | Alarms | Common | Alarms | ['MP20', 'MP70', 'MX800', 'X3']
4460 |  |  | Audio Alarm Priority | unitless | Overall audible alarm priority. This is the current highest priority value of currently active audible alarms. | Double | Alarms | Common | Alarms | ['MP20', 'MP70', 'MX800', 'X3']
4461 |  |  | Technical Alarm Source ID | unitless | [] | Double (Multiset) | Alarms | Common | Alarms | ['MP20', 'MP70', 'MX800', 'X3']
4462 |  |  | Technical Alarm Global State | unitless | Consult detailed help for more information on this variable | Double (Multiset) | Alarms | Common | Alarms | ['MP20', 'MP70', 'MX800', 'X3']
4463 |  |  | Technical Alarm Priority | unitless | Consult detailed help for more information on this variable | Double (Multiset) | Alarms | Common | Alarms | ['MP20', 'MP70', 'MX800', 'X3']
4464 |  |  | Technical Alarm Flag | unitless | Consult detailed help for more information on this variable | Double (Multiset) | Alarms | Common | Alarms | ['MP20', 'MP70', 'MX800', 'X3']
4465 |  |  | Technical Alarm Priority | unitless | [] | Double (Multiset) | Alarms | Common | Alarms | ['MP20', 'MP70', 'MX800', 'X3']
4466 |  |  | Technical Alarm Description | unitless | [] | String (Multiset) | Alarms | Common | Alarms | ['MP20', 'MP70', 'MX800', 'X3']
4467 |  |  | Patient Alarm Source ID | unitless | [] | Double (Multiset) | Alarms | Common | Alarms | ['MP20', 'MP70', 'MX800', 'X3']
4468 |  |  | Patient Alarm Global State | unitless | ['Alarm disabled status', 'Consult detailed help for more information on this variable\xa0'] | ['Double (Multiset)\xa0', 'Long\xa0'] | Alarms | ['Common\xa0', 'V60 (PVOI)\xa0'] | Alarms | ['X3', 'MX800', 'MP20', 'V60', 'MP70']
4469 |  |  | Patient Alarm Priority | unitless | Consult detailed help for more information on this variable | Double (Multiset) | Alarms | Common | Alarms | ['MP20', 'MP70', 'MX800', 'X3']
4470 |  |  | Patient Alarm Flag | unitless | Consult detailed help for more information on this variable | Double (Multiset) | Alarms | Common | Alarms | ['MP20', 'MP70', 'MX800', 'X3']
4471 |  |  | Patient Alarm Priority | unitless | [] | Double (Multiset) | Alarms | Common | Alarms | ['MP20', 'MP70', 'MX800', 'X3']
4472 |  |  | Patient Alarm Description | unitless | [] | String (Multiset) | Alarms | Common | Alarms | ['MP20', 'MP70', 'MX800', 'X3']
4473 |  |  | Device Nomenclature Revision | unitless | Format: major.minor (such as 1.0) | String | DeviceInfo | Common | System Information | ['MP20', 'MP70', 'MX800', 'X3']
4474 |  |  | Device Text Catalog Revision | unitless | Format: "MM.NN mm.nn", where "MM.NN" is text catalog revision and "mm.nn" is language revision. | String | DeviceInfo | Common | System Information | ['MP20', 'MP70', 'MX800', 'X3']
4475 |  |  | Device Data Format | unitless | Always 11 (Unicode). | String | DeviceInfo | Common | System Information | ['MP20', 'MP70', 'MX800', 'X3']
4476 |  |  | Device Operating Area | unitless | Possible values: "unspecified", "operating room", "intensive care unit", "neonatal intensive care unit", and "cardiology care unit". | String | DeviceInfo | Common | System Information | ['MP20', 'MP70', 'MX800', 'X3']
4496 |  |  | Patient O2 Consumption (VO2) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | VO2 | ['MP70', 'MX800', 'X3']
4497 |  |  | Patient CO2 Production (VCO2) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | bbvCO2 | ['MP70', 'MX800', 'X3']
4507 |  |  | Arterial O2 Saturation (Left) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | SpO2 l | ['MP70', 'MX800', 'X3']
4508 |  |  | Arterial O2 Saturation (Right) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | SpO2 r | ['MP70', 'MX800', 'X3']
4509 |  |  | SpO2 Difference | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | DeltaSpO2 | ['MP70', 'MX800', 'X3']
4510 |  |  | Perfusion Indicator | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Perf | ['MP70', 'MX800', 'X3']
4511 |  |  | Relative Perfusion (Left) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Perf l | ['MP70', 'MX800', 'X3']
4512 |  |  | Relative Perfusion (Right) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Perf r | ['MP70', 'MX800', 'X3']
4513 |  |  | Ao Systolic | - | [] | Double | VitalSigns | Common | IBP: Ao | ['MP20', 'MP70', 'MX800', 'X3']
4514 |  |  | Ao Diastolic | - | [] | Double | VitalSigns | Common | IBP: Ao | ['MP20', 'MP70', 'MX800', 'X3']
4515 |  |  | Ao Mean | - | [] | Double | VitalSigns | Common | IBP: Ao | ['MP20', 'MP70', 'MX800', 'X3']
4516 |  |  | RAP Systolic | - | [] | Double | VitalSigns | Common | IBP: RAP | ['MP20', 'MP70', 'MX800', 'X3']
4517 |  |  | RAP Diastolic | - | [] | Double | VitalSigns | Common | IBP: RAP | ['MP20', 'MP70', 'MX800', 'X3']
4518 |  |  | RAP Mean | - | [] | Double | VitalSigns | Common | IBP: RAP | ['MP20', 'MP70', 'MX800', 'X3']
4519 |  |  | Intracranial Pressure ICP Systolic | - | [] | Double | VitalSigns | Common | IBP: ICP | ['MP20', 'MP70', 'MX800', 'X3']
4520 |  |  | Intracranial Pressure ICP Diastolic | - | [] | Double | VitalSigns | Common | IBP: ICP | ['MP20', 'MP70', 'MX800', 'X3']
4521 |  |  | ICP Mean | - | [] | Double | VitalSigns | Common | IBP: ICP | ['MP20', 'MP70', 'MX800', 'X3']
4522 |  |  | UAP Systolic | - | [] | Double | VitalSigns | Common | IBP: UAP | ['MP20', 'MP70', 'MX800', 'X3']
4523 |  |  | UAP Diastolic | - | [] | Double | VitalSigns | Common | IBP: UAP | ['MP20', 'MP70', 'MX800', 'X3']
4524 |  |  | UAP Mean | - | [] | Double | VitalSigns | Common | IBP: UAP | ['MP20', 'MP70', 'MX800', 'X3']
4525 |  |  | UVP Systolic | - | [] | Double | VitalSigns | Common | IBP: UVP | ['MP20', 'MP70', 'MX800', 'X3']
4526 |  |  | UVP Diastolic | - | [] | Double | VitalSigns | Common | IBP: UVP | ['MP20', 'MP70', 'MX800', 'X3']
4527 |  |  | UVP Mean | - | [] | Double | VitalSigns | Common | IBP: UVP | ['MP20', 'MP70', 'MX800', 'X3']
4528 |  |  | Unspecific Pressure Systolic | - | [] | Double | VitalSigns | Common | IBP: P | ['MP20', 'MP70', 'MX800', 'X3']
4529 |  |  | Unspecific Pressure Diastolic | - | [] | Double | VitalSigns | Common | IBP: P | ['MP20', 'MP70', 'MX800', 'X3']
4530 |  |  | Unspecific Pressure Mean | - | [] | Double | VitalSigns | Common | IBP: P | ['MP20', 'MP70', 'MX800', 'X3']
4531 |  |  | Cerebral Perfusion Pressure (CPP) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | IBP: CPP | ['MP70', 'MX800', 'X3']
4533 |  |  | Index of Left Ventricular Contractility | - | This variable contains measurement from label "NLS_NOM_GRAD_PRESS_BLD_AORT_POS_MAX" and observed value "NOM_GRAD_PRESS_BLD_AORT_POS_MAX". | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | dPmax | ['MP70', 'MX800', 'X3']
4556 |  |  | Aortic Blood Pressure (Ao) Wave | - | [] | Double | Waveforms | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Waves: Ao | ['MP70', 'MX800', 'X3']
4560 |  |  | Auxiliary Waveform | - | [] | Double | Waveforms | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Waves: AUX | ['MP70', 'MX800', 'X3']
4561 |  |  | Intermittent Mandatory Ventilation Status | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | VENT: IMV | ['MP70', 'MX800', 'X3']
4562 |  |  | Left Cardiac Work | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | LCW | ['MP70', 'MX800', 'X3']
4564 |  |  | Right Cardiac Work | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | RCW | ['MP70', 'MX800', 'X3']
4565 |  |  | Serum Creatinine | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | SCrea | ['MP70', 'MX800', 'X3']
4566 |  |  | Urine Plasma Osmolarity Ratio | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | U/POsm | ['MP70', 'MX800', 'X3']
4567 |  |  | Urine Sodium Potassium Ratio | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | UrNa/K | ['MP70', 'MX800', 'X3']
4568 |  |  | Urine Volume | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | UrVol | ['MP70', 'MX800', 'X3']
4570 |  |  | Dialyzer Name | unitless | [] | String | DeviceInfo | Common | [] | 5008
4573 |  |  | Dialysis Filter Lifetime | min | Filter Lifetime | Long | DeviceInfo | Multifiltrate PRO | Dialysis Course | Multifiltrate PRO
4574 |  |  | Left Cardiac Work Index | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | LCWI | ['MP70', 'MX800', 'X3']
4576 |  |  | Right Cardiac Work Index | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | RCWI | ['MP70', 'MX800', 'X3']
4577 |  |  | Light Intensity (SvO2 Module) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | LI | ['MP70', 'MX800', 'X3']
4578 |  |  | SpO2 SvO2 Difference | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Sp-vO2 | ['MP70', 'MX800', 'X3']
4579 |  |  | Dialysate Name | unitless | [] | String | PatientInfo | Common | [] | 5008
4581 |  |  | Normalized Post Replacement Volume | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | VdpoNh | ['MP70', 'MX800', 'X3']
4582 |  |  | Normalized Effluent Volume | - | This variable is associated with ChannelID identifier. | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Veff/EffDn | ['MP70', 'MX800', 'X3']
4584 |  |  | Normalized Total Replacement Volume (PBP) | - | This variable is associated with ChannelID identifier. | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | VrePN/VrePNh | ['MP70', 'MX800', 'X3']
4585 |  |  | Normalized Replacement Volume (Pre) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | VdprNh | ['MP70', 'MX800', 'X3']
4586 |  |  | Normalized Ultrafiltration Volume | - | This variable is associated with ChannelID identifier. | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | VufrN/VufrNh | ['MP70', 'MX800', 'X3']
4587 |  |  | Patient Fluid Balance Limit Setting | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sEPFL | ['MP70', 'MX800', 'X3']
4588 |  |  | Total Replacement Volume (PBP) | - | This variable is associated with ChannelID identifier. | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | VreP/VreP/h | ['MP70', 'MX800', 'X3']
4589 |  |  | Cumulated Liquid Volume (Pre) | - | This variable is associated with ChannelID identifier. | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Vdpr/Vdpr/h | ['MP70', 'MX800', 'X3']
4592 |  |  | Dialysate Average Flow Rate | mL/min | [] | ULong | VitalSigns | Common | [] | 5008
4593 |  |  | Filtration Fraction Setting | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sFufrP | ['MP70', 'MX800', 'X3']
4596 |  |  | System Pressure High Alarm | unitless | Possible values: 0 = Active 1 = Not Active | Boolean | Alarms | Common | [] | Criticool
4601 |  |  | Dialysate Sodium (Na+) Concentration Setting | mmol/L | [] | Double | Settings | Common | [] | 5008
4605 |  |  | Dialysate Bicarbonate (HCO3) Concentration Setting | mmol/L | [] | Double | Settings | Common | [] | 5008
4608 |  |  | Urea Clearance | mL/min | [] | Short | VitalSigns | Common | [] | 5008
4609 |  |  | Fluid Removal Average Rate | mL/h | [] | Long | VitalSigns | Common | [] | 5008
4610 |  |  | Anticoagulant Stopping Interval Setting | s | [] | Long | Settings | Common | [] | 5008
4614 |  |  | Target Treatement Effectiveness (Kt/V) Setting | unitless | [] | Double | Settings | Common | [] | 5008
4615 |  |  | Initial Hematocrit | % | [] | Double | VitalSigns | Common | [] | 5008
4616 |  |  | Diff.-Temperature between Tblood and Tskin | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
4617 |  |  | Diff.-Temperature between Tblood and Trect | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
4618 |  |  | Diff.-Temperature between Tblood and Tnaso | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
4619 |  |  | Diff.-Temperature between Tblood and Tesop | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
4620 |  |  | Diff.-Temperature between Tblood and Tcore | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
4621 |  |  | Diff.-Temperature between Tblood and Tart | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
4622 |  |  | Diff.-Temperature between Tblood and Tven | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
4623 |  |  | Diff.-Temperature between Tskin and Trect | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
4624 |  |  | Diff.-Temperature between Tskin and Tnaso | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
4625 |  |  | Diff.-Temperature between Tskin and Tesop | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
4626 |  |  | Diff.-Temperature between Tskin and Tcore | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
4627 |  |  | Diff.-Temperature between Tskin and Tart | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
4628 |  |  | Diff.-Temperature between Tskin and Tven | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
4629 |  |  | Diff.-Temperature between Trect and Tnaso | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
4630 |  |  | Diff.-Temperature between Trect and Tesop | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
4631 |  |  | Diff.-Temperature between Trect and Tcore | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
4632 |  |  | Diff.-Temperature between Trect and Tart | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
4633 |  |  | Diff.-Temperature between Trect and Tven | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
4634 |  |  | Diff.-Temperature between Tnaso and Tesop | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
4635 |  |  | Diff.-Temperature between Tnaso and Tcore | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
4636 |  |  | Diff.-Temperature between Tnaso and Tart | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
4637 |  |  | Diff.-Temperature between Tnaso and Tven | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
4638 |  |  | Diff.-Temperature between Tesop and Tcore | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
4639 |  |  | Diff.-Temperature between Tesop and Tart | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
4640 |  |  | Diff.-Temperature between Tesop and Tven | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
4641 |  |  | Diff.-Temperature between Tcore and Tart | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
4642 |  |  | Diff.-Temperature between Tcore and Tven | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
4643 |  |  | Diff.-Temperature between Tart and Tven | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
4644 |  |  | Newborn Body Length | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Length | ['MP70', 'MX800', 'X3']
4646 |  |  | Newborn Gestational Age | - | [] | Double | PatientInfo | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | G.Age | ['MP70', 'MX800', 'X3']
4647 |  |  | Calories | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
4648 |  |  | Input/Output Totals | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
4649 |  |  | Newborn Apgar Scores | - | [] | Double | PatientInfo | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
4650 |  |  | BSA Formula Boyd | - | [] | Double | PatientInfo | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | BSA(B) | ['MP70', 'MX800', 'X3']
4651 |  |  | BSA Formula Dubois | - | [] | Double | PatientInfo | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | BSA(D) | ['MP70', 'MX800', 'X3']
4653 |  |  | Pupil Reaction (Left Eye) | - | Light reaction of left eye's pupil. | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | PRL | ['MP70', 'MX800', 'X3']
4654 |  |  | Pupil Reaction (Right Eye) | - | Light reaction of right eye's pupil. | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | PRR | ['MP70', 'MX800', 'X3']
4661 |  |  | Pleth Gain | - | [] | Double | DeviceInfo | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | PlGain | ['MP70', 'MX800', 'X3']
4662 |  |  | Airway Temperature | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Tairwy | ['MP70', 'MX800', 'X3']
4663 |  |  | Ventilation Perfusion Index | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | VENT: VQI | ['MP70', 'MX800', 'X3']
4664 |  |  | O2 Extraction Index | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | O2EI | ['MP70', 'MX800', 'X3']
4665 |  |  | Duration Above Base Pressure | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | VENT: DPosP | ['MP70', 'MX800', 'X3']
4666 |  |  | Inspiration Rise Time | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | VENT: RiseTi | ['MP70', 'MX800', 'X3']
4667 |  |  | Continuous Positive Airway Pressure (CPAP) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | CPAP | ['MP70', 'MX800', 'X3']
4668 |  |  | HFV Amplitude | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | VENT: HFVAmp | ['MP70', 'MX800', 'X3']
4669 |  |  | Urine Urea | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | UrUrea | ['MP70', 'MX800', 'X3']
4670 |  |  | Urine pH Value | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | UrpH | ['MP70', 'MX800', 'X3']
4671 |  |  | Total of CO2 (tCO2) | - | Result of Blood gas Analysis. | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | tCO2 | ['MP70', 'MX800', 'X3']
4672 |  |  | Total Bilirubin | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | tBili | ['MP70', 'MX800', 'X3']
4673 |  |  | Serum Glucose | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | SerGlu | ['MP70', 'MX800', 'X3']
4674 |  |  | Urine Glucose | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | UrGlu | ['MP70', 'MX800', 'X3']
4675 |  |  | Direct Bilirubin | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | dBili | ['MP70', 'MX800', 'X3']
4676 |  |  | Serum Calcium | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | SerCa | ['MP70', 'MX800', 'X3']
4677 |  |  | Total Serum Calcium | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | tSerCa | ['MP70', 'MX800', 'X3']
4678 |  |  | Serum Magnesium | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | SerMg | ['MP70', 'MX800', 'X3']
4679 |  |  | Serum Phosphat | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | SerPho | ['MP70', 'MX800', 'X3']
4680 |  |  | Serum Kalium (Potassium) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | SerK | ['MP70', 'MX800', 'X3']
4681 |  |  | Serum Clorid | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | SerCl | ['MP70', 'MX800', 'X3']
4682 |  |  | Serum Albumine | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | SerAlb | ['MP70', 'MX800', 'X3']
4683 |  |  | Urine Clorid | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | UrCl | ['MP70', 'MX800', 'X3']
4684 |  |  | Serum Globulin | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | SerGlo | ['MP70', 'MX800', 'X3']
4685 |  |  | Serum Protein | ['-', 'g/dL'] | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'Common\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | SerPro | ['MP70', '5008', 'MX800', 'X3']
4686 |  |  | Serum Urea | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | SrUrea | ['MP70', 'MX800', 'X3']
4687 |  |  | Partial Thromboplastin Time | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | PTT | ['MP70', 'MX800', 'X3']
4688 |  |  | Prothrombin Time | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | PT | ['MP70', 'MX800', 'X3']
4689 |  |  | Thrombin Time | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | TT | ['MP70', 'MX800', 'X3']
4690 |  |  | Thromboplastine Time (Quick Time) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Quick | ['MP70', 'MX800', 'X3']
4691 |  |  | Alkalische Phosphatase | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | AP | ['MP70', 'MX800', 'X3']
4692 |  |  | Alpha Amylase | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | alphaA | ['MP70', 'MX800', 'X3']
4693 |  |  | Cholesterinesterase | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | CHE | ['MP70', 'MX800', 'X3']
4694 |  |  | Creatinin Kinase | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | SerCK | ['MP70', 'MX800', 'X3']
4695 |  |  | Creatine Kinase of Type Muscle-Brain | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | CK-MB | ['MP70', 'MX800', 'X3']
4696 |  |  | Creatine Kinase of Type Muscle | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | CK-MM | ['MP70', 'MX800', 'X3']
4697 |  |  | Gamma Glutamyl Transpeptidase | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | GGT | ['MP70', 'MX800', 'X3']
4698 |  |  | Glutamic Oxaloacetic Transaminase | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | GOT | ['MP70', 'MX800', 'X3']
4699 |  |  | Glutamic Pyruvic Transaminase | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | GPT | ['MP70', 'MX800', 'X3']
4700 |  |  | Cholesterin | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Chol | ['MP70', 'MX800', 'X3']
4701 |  |  | Triglyzeride | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | TGL | ['MP70', 'MX800', 'X3']
4702 |  |  | Urine Protein | - | (Total) Protein in Urine | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | UrPro | ['MP70', 'MX800', 'X3']
4703 |  |  | Urine Calcium | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | UrCa | ['MP70', 'MX800', 'X3']
4704 |  |  | Carboxy Hemoglobin | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | CO-Hb | ['MP70', 'MX800', 'X3']
4705 |  |  | Total Protein | ['-', 'g/dL'] | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'Common\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | tPro | ['MP70', '5008', 'MX800', 'X3']
4706 |  |  | Lactate Dehydrogenase | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | LDH | ['MP70', 'MX800', 'X3']
4707 |  |  | Aspartin Aminotransferase | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | AST | ['MP70', 'MX800', 'X3']
4708 |  |  | Alveolarproteinose (Alkaline Phosphatase) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | ALP | ['MP70', 'MX800', 'X3']
4709 |  |  | Reticulocyte Count | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | RC | ['MP70', 'MX800', 'X3']
4710 |  |  | Coagulation Time | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | CT | ['MP70', 'MX800', 'X3']
4711 |  |  | Erithrocyte Sedimentation Rate | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | ESR | ['MP70', 'MX800', 'X3']
4712 |  |  | Packed Cell Volume | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
4713 |  |  | Kaolin Cephalin Time | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | KCT | ['MP70', 'MX800', 'X3']
4715 |  |  | Minimum Pressure During Breathing Cycle | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
4716 |  |  | Infusion Pump Accumulated Volume | - | [] | Double | OtherMeasuredData | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | AccVol | ['MP70', 'MX800', 'X3']
4717 |  |  | Transthoracic Impedance | - | [] | Double | VitalSigns | Common | T.I. | ['MP20', 'MP70', 'MX800', 'X3']
4718 |  |  | Inspired Expired N2O Difference | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | VENT: i-eN2O | ['MP70', 'MX800', 'X3']
4719 |  |  | Inspired Expired Halothane Difference | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | VENT: i-eHAL | ['MP70', 'MX800', 'X3']
4720 |  |  | Inspired Expired Enfluran Difference | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | VENT: i-eENF | ['MP70', 'MX800', 'X3']
4721 |  |  | Inspired Expired Isofluran Difference | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | VENT: i-eISO | ['MP70', 'MX800', 'X3']
4722 |  |  | Inspired Expired Sevofluran Difference | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | VENT: i-eSEV | ['MP70', 'MX800', 'X3']
4723 |  |  | Inspired Expired Desfluran Difference | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | VENT: i-eDES | ['MP70', 'MX800', 'X3']
4724 |  |  | Inspired Expired Agent Difference | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | VENT: i-eAGT | ['MP70', 'MX800', 'X3']
4725 |  |  | Inspired Expired O2 Difference | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
4726 |  |  | Patient Circuit O2 | - | O2 measured in the Patient Circuit | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | VENT: cktO2 | ['MP70', 'MX800', 'X3']
4727 |  |  | Barometric Pressure in Sample Chamber | - | This variable is obsolete that may be removed in a future release. It is duplicated with variable 4082. | Double | OtherMeasuredData | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Obsolete Variables | ['MP70', 'MX800', 'X3']
4728 |  |  | Minute Volume | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | MINVOL | ['MP70', 'MX800', 'X3']
4729 |  |  | Mean Airway Pressure | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | MnAwP | ['MP70', 'MX800', 'X3']
4730 |  |  | Airway Respiration Rate | - | [] | Double | VitalSigns | Common | awRR | ['MP20', 'MP70', 'MX800', 'X3']
4731 |  |  | Inspiratory Airway Pressure Waveform | - | [] | Double | Waveforms | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Waves: AWPin | ['MP70', 'MX800', 'X3']
4732 |  |  | Inspiratory Flow Waveform | ['L/min', '-'] | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. | Double | Waveforms | ['IntelliVue MX800\xa0', 'Evita Infinity V500 [MEDIBUS.X]\xa0', 'IntelliVue X3\xa0', 'Evita Infinity V800\xa0', 'IntelliVue MP70\xa0'] | Waves: AWFin | ['X3', 'MX800', 'Evita V800', 'MP70', 'Evita V500']
4733 |  |  | HFV Inspired Mandatory Minute Volume | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | HFMVin | ['MP70', 'MX800', 'X3']
4734 |  |  | HFV Gas Transport Coefficient Value | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | DCO2 | ['MP70', 'MX800', 'X3']
4735 |  |  | HFV Inspired Tidal Volume | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | VENT: HFTVin | ['MP70', 'MX800', 'X3']
4736 |  |  | HFV Tidal Volume | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | VENT: HFVTV | ['MP70', 'MX800', 'X3']
4738 |  |  | Generic Heart Rate | - | Heart Rate received from an external device | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | extHR | ['MP70', 'MX800', 'X3']
4739 |  |  | PERF from the 2nd SpO2/PLETH module | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
4741 |  |  | Percent Alpha (Left) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | EEG: LT %AL | ['MP70', 'MX800', 'X3']
4742 |  |  | Percent Beta (Left) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | EEG: LT %BE | ['MP70', 'MX800', 'X3']
4745 |  |  | Absolute Alpha (Left) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | EEG: LT AL | ['MP70', 'MX800', 'X3']
4746 |  |  | Absolute Beta (Left) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | EEG: LT BE | ['MP70', 'MX800', 'X3']
4747 |  |  | Absolute Delta (Left) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | EEG: LT DL | ['MP70', 'MX800', 'X3']
4748 |  |  | Absolute Theta (Left) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | EEG: LT TH | ['MP70', 'MX800', 'X3']
4749 |  |  | Mean Dominant Frequency (Left) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | EEG: LT MDF | ['MP70', 'MX800', 'X3']
4750 |  |  | Median Power Frequency (Left) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | EEG: LT MPF | ['MP70', 'MX800', 'X3']
4751 |  |  | Peak Power Frequency (Left) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | EEG: LT PPF | ['MP70', 'MX800', 'X3']
4752 |  |  | Spectral Edge Frequency (Left) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | EEG: LT SEF | ['MP70', 'MX800', 'X3']
4753 |  |  | Total Power (Left) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | EEG: LT TP | ['MP70', 'MX800', 'X3']
4754 |  |  | Scale of the Left Channel EEG Wave | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | EEG: LSCALE | ['MP70', 'MX800', 'X3']
4755 |  |  | Percent Alpha (Right) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | EEG: RT %AL | ['MP70', 'MX800', 'X3']
4756 |  |  | Percent Beta (Right) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | EEG: RT %BE | ['MP70', 'MX800', 'X3']
4759 |  |  | Absolute Alpha (Right) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | EEG: RT AL | ['MP70', 'MX800', 'X3']
4760 |  |  | Absolute Beta (Right) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | EEG: RT BE | ['MP70', 'MX800', 'X3']
4761 |  |  | Absolute Delta (Right) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | EEG: RT DL | ['MP70', 'MX800', 'X3']
4762 |  |  | Absolute Theta (Right) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | EEG: RT TH | ['MP70', 'MX800', 'X3']
4763 |  |  | Mean Dominant Frequency (Right) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | EEG: RT MDF | ['MP70', 'MX800', 'X3']
4764 |  |  | Median Power Frequency (Right) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | EEG: RT MPF | ['MP70', 'MX800', 'X3']
4765 |  |  | Peak Power Frequency (Right) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | EEG: RT PPF | ['MP70', 'MX800', 'X3']
4766 |  |  | Spectral Edge Frequency (Right) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | EEG: RT SEF | ['MP70', 'MX800', 'X3']
4767 |  |  | Total Power (Right) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | EEG: RT TP | ['MP70', 'MX800', 'X3']
4768 |  |  | Scale of the Right Channel EEG Wave | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | EEG: RSCALE | ['MP70', 'MX800', 'X3']
4769 |  |  | Duration of Positive Pressure | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
4772 |  |  | Fresh Gas N2O Concentration | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | fgN2O | ['MP70', 'MX800', 'X3']
4774 |  |  | Liquid Level (Anesthetic Agent Bottle) | - | [] | Double | OtherMeasuredData | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | AGTLev | ['MP70', 'MX800', 'X3']
4775 |  |  | Liquid Level (ISOflurane Bottle) | - | [] | Double | OtherMeasuredData | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | ISOLev | ['MP70', 'MX800', 'X3']
4776 |  |  | Liquid Level (ENFlurane Bottle) | - | [] | Double | OtherMeasuredData | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | ENFLev | ['MP70', 'MX800', 'X3']
4777 |  |  | Liquid Level (HALothane Bottle) | - | [] | Double | OtherMeasuredData | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | HALLev | ['MP70', 'MX800', 'X3']
4778 |  |  | Liquid Level (DESflurane Bottle) | - | [] | Double | OtherMeasuredData | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | DESLev | ['MP70', 'MX800', 'X3']
4779 |  |  | Liquid Level (SEVoflurane Bottle) | - | [] | Double | OtherMeasuredData | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | SEVLev | ['MP70', 'MX800', 'X3']
4780 |  |  | Invasive Blood Pressure Wave | - | [] | Double | Waveforms | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Waves: P | ['MP70', 'MX800', 'X3']
4781 |  |  | Urine Shift Volume | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | UrVSht | ['MP70', 'MX800', 'X3']
4782 |  |  | Urine Flow Rate | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | UrFl | ['MP70', 'MX800', 'X3']
4786 |  |  | Mean Dominant Frequency (CH1) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | EEG: Frequ1 | ['MP70', 'MX800', 'X3']
4787 |  |  | Mean Dominant Frequency (CH2) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | EEG: Frequ2 | ['MP70', 'MX800', 'X3']
4788 |  |  | Peak Power Frequency (CH1) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | EEG: Frequ1 | ['MP70', 'MX800', 'X3']
4789 |  |  | Peak Power Frequency (CH2) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | EEG: Frequ2 | ['MP70', 'MX800', 'X3']
4794 |  |  | % of Power in the Alpha Band (CH1) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | EEG: Prcnt1 | ['MP70', 'MX800', 'X3']
4795 |  |  | % of Power in the Alpha Band (CH2) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | EEG: Prcnt2 | ['MP70', 'MX800', 'X3']
4796 |  |  | % of Power in the Beta Band (1st EEG) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | EEG: Prcnt1 | ['MP70', 'MX800', 'X3']
4797 |  |  | % of Power in the Beta Band (2nd EEG) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | EEG: Prcnt2 | ['MP70', 'MX800', 'X3']
4800 |  |  | Hemoglobin | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Hb | ['MP70', 'MX800', 'X3']
4801 |  |  | Patient Temperature | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
4802 |  |  | Capillaries pH Value | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | pHc | ['MP70', 'MX800', 'X3']
4803 |  |  | Adjusted pH (at Patient Temperature) | - | Adjusted pH at Patient Temperature | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | &pH | ['MP70', 'MX800', 'X3']
4804 |  |  | Adjusted Capillaries pH Value | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | &pHc | ['MP70', 'MX800', 'X3']
4805 |  |  | Partial O2 in the Capillaries | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | PcO2 | ['MP70', 'MX800', 'X3']
4806 |  |  | Adjusted PO2 (at Patient Temperature) | - | Adjusted PO2 at Patient Temperature | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | &PO2 | ['MP70', 'MX800', 'X3']
4807 |  |  | Adjusted PcO2 (at Patient Temperature) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | &PcO2 | ['MP70', 'MX800', 'X3']
4808 |  |  | Venous CO2 Partial Pressure (PvCO2) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | PvCO2 | ['MP70', 'MX800', 'X3']
4809 |  |  | Partial CO2 in the Capillaries | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | PcCO2 | ['MP70', 'MX800', 'X3']
4811 |  |  | Adjusted PcCO2 (at Patient Temperature) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | &PcCO2 | ['MP70', 'MX800', 'X3']
4812 |  |  | Calculated Total CO2 (tCO2) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | CALC: 'tCO2 | ['MP70', 'MX800', 'X3']
4813 |  |  | Calculated SO2 | - | Calculated SaO2 | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | CALC: 'SaO2 | ['MP70', 'MX800', 'X3']
4815 |  |  | Calculated SvO2 | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | CALC: 'SvO2 | ['MP70', 'MX800', 'X3']
4816 |  |  | Calculated ScO2 | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | CALC: 'ScO2 | ['MP70', 'MX800', 'X3']
4817 |  |  | Calculated HCO3 | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | CALC: HCO3 | ['MP70', 'MX800', 'X3']
4818 |  |  | Base Excess of Extra Cellular Fluid | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | CALC: Beecf | ['MP70', 'MX800', 'X3']
4819 |  |  | Calculated Anion Gap | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | CALC: AnGap | ['MP70', 'MX800', 'X3']
4821 |  |  | Urea | - | Measured by the i-Stat module. | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Urea | ['MP70', 'MX800', 'X3']
4822 |  |  | Manually Entered FiO2 | - | In the Blood Review window. | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
4823 |  |  | Calculated Arterial Blood Base Excess | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | BE,B | ['MP70', 'MX800', 'X3']
4824 |  |  | Ionized Magnesium | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | iMg | ['MP70', 'MX800', 'X3']
4826 |  |  | Desired Module Temperature | - | It's the selected/adjusted Temp in i-Stat Window. | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
4827 |  |  | Calculated BUN/Creatinine Ratio | - | Calculated by the i-Stat module. | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | B/Cre | ['MP70', 'MX800', 'X3']
4828 |  |  | Calculated Urea/Creatinine Ratio | - | Calculated by the i-Stat module. | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | U/Cre | ['MP70', 'MX800', 'X3']
4829 |  |  | Lactate | - | Measured value by the i-Stat module. | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Lact | ['MP70', 'MX800', 'X3']
4830 |  |  | Device Simulator Status | - | Device data with the values: Passed and Failed. | Double | DeviceInfo | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
4835 |  |  | Pressure of Sensor 1 | - | Pressure value of sensor 1. Device specific data. | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
4836 |  |  | Pressure of Sensor 2 | - | Pressure value of sensor 2. Device specific data. | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
4837 |  |  | Module Temperature | - | [] | Double | OtherMeasuredData | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
4843 |  |  | Air Temperature | - | [] | Double | OtherMeasuredData | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
4844 |  |  | Micro Environment Air Temperature Setting | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sTemp | ['MP70', 'MX800', 'X3']
4847 |  |  | Weight of the Urine Disposable Bag | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | BagWgt | ['MP70', 'MX800', 'X3']
4848 |  |  | Total Urine Volume of the Current Measurement Period | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | tUrVol | ['MP70', 'MX800', 'X3']
4849 |  |  | Urine Fluid Density | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | UrDens | ['MP70', 'MX800', 'X3']
4850 |  |  | Urine Fluid Temperature | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
4854 |  |  | Ventilation Activated Status | - | Parameter to Informs whether the Patient is Ventilated | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | VENT: PtVent | ['MP70', 'MX800', 'X3']
4856 |  |  | Simplified Acuity Physiologic Score | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
4857 |  |  | Value Normal | - | Missing parameters in SAPS parameter set which are considered in normal range. | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
4858 |  |  | Aspartate Aminotransferase | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
4859 |  |  | Spontaneous Respiration Rate | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | SpRR | ['MP70', 'MX800', 'X3']
4864 |  |  | Current Fluid in Urine Bag | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | BagVol | ['MP70', 'MX800', 'X3']
4873 |  |  | NMT TOF Count | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | NMT: TOFcnt | ['MP70', 'MX800', 'X3']
4874 |  |  | NMT TOF Ratio | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | NMT: TOFrat | ['MP70', 'MX800', 'X3']
4875 |  |  | Twitch Height of the 1Hz/0.1Hz Stimulation Response | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | NMT: Twitch | ['MP70', 'MX800', 'X3']
4876 |  |  | Post Tetatic Count | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | NMT: PTC | ['MP70', 'MX800', 'X3']
4877 |  |  | Remaining Time to Next Stimulation | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | NMT: RemTi | ['MP70', 'MX800', 'X3']
4878 |  |  | TOF Response TOF1 | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | NMT: TOF1 | ['MP70', 'MX800', 'X3']
4879 |  |  | TOF Response TOF2 | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | NMT: TOF2 | ['MP70', 'MX800', 'X3']
4880 |  |  | TOF Response TOF3 | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | NMT: TOF3 | ['MP70', 'MX800', 'X3']
4881 |  |  | TOF Response TOF4 | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | NMT: TOF4 | ['MP70', 'MX800', 'X3']
4883 |  |  | Activated Partial Thromboplastin Time (aPTT Whole Blood Time) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | aPTTWB | ['MP70', 'MX800', 'X3']
4884 |  |  | Activated Partial Thromboplastin Time (aPTT Plasma Equivalent Time) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | aPTTPE | ['MP70', 'MX800', 'X3']
4886 |  |  | Prothrombin Time (Whole Blood Time) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | PT WB | ['MP70', 'MX800', 'X3']
4887 |  |  | Prothrombin Time (Plasma Equivalent Time) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | PT PE | ['MP70', 'MX800', 'X3']
4888 |  |  | Prothrombin Time Ratio | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | PTrat | ['MP70', 'MX800', 'X3']
4889 |  |  | Spontaneous Inspiration Time | ['s', '-'] | For Evita 4, Evita 2 dura, Evita XL: Available only with EvitaXL firmware version 7.02 or higher. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. | Double | VitalSigns | ['IntelliVue MX800\xa0', 'Evita Infinity V500 [MEDIBUS.X]\xa0', 'IntelliVue X3\xa0', 'Evita Infinity V800\xa0', 'IntelliVue MP70\xa0'] | InsTi | ['X3', 'MX800', 'Evita V800', 'MP70', 'Evita V500']
4890 |  |  | Overdistension Index | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | C20/C | ['MP70', 'MX800', 'X3']
4891 |  |  | Lung Time Constant | ['s', '-'] | For Evita 4, Evita 2 dura, Evita XL: Available only with EvitaXL firmware version 7.00 or higher. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. | Double | VitalSigns | ['IntelliVue MX800\xa0', 'Evita Infinity V500 [MEDIBUS.X]\xa0', 'IntelliVue X3\xa0', 'Evita Infinity V800\xa0', 'IntelliVue MP70\xa0'] | TC | ['X3', 'MX800', 'Evita V800', 'MP70', 'Evita V500']
4893 |  |  | Rate Volume Ratio | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | RVrat | ['MP70', 'MX800', 'X3']
4894 |  |  | Normalized Ionized Calcium (Ca) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | iCa(N) | ['MP70', 'MX800', 'X3']
4956 |  |  | Peak Inspiratory Flow Setting | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | s%PIF | ['MP70', 'MX800', 'X3']
4957 |  |  | Trigger Sensitivity Setting | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'Common\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['V60', 'MP70', 'MX800', 'X3']
4958 |  |  | Plateau Time Setting | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sPltTi | ['MP70', 'MX800', 'X3']
4959 |  |  | Multiple Sigh Number Setting | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | VENT: sSghNr | ['MP70', 'MX800', 'X3']
4960 |  |  | Apnea Tidal Volume Setting | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sATV | ['MP70', 'MX800', 'X3']
4961 |  |  | Apnea Respiration Rate Setting | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sARR | ['MP70', 'MX800', 'X3']
4962 |  |  | Apnea Peak Flow Setting | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sAPkFl | ['MP70', 'MX800', 'X3']
4963 |  |  | Apnea Inspired O2 Setting | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | VENT: sAFIO2 | ['MP70', 'MX800', 'X3']
4964 |  |  | PSV Pressure Setting | - | Pressure Support Ventilation Setting | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | VENT: sPSV | ['MP70', 'MX800', 'X3']
4966 |  |  | Sigh Activation Status | ['unitless', '-'] | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. A value of 1 indicates that the sigh function has been added to the current ventilation mode. A value of 0 indicates that the sigh function has not been added to the current ventilation mode. Text message displayed on device screen: "/Sigh" | ['Double\xa0', 'Boolean\xa0'] | DeviceInfo | ['IntelliVue MX800\xa0', 'Evita Infinity V500 [MEDIBUS.X]\xa0', 'IntelliVue X3\xa0', 'Evita Infinity V800\xa0', 'IntelliVue MP70\xa0'] | VENT: sEnSgh | ['X3', 'MX800', 'Evita V800', 'MP70', 'Evita V500']
4967 |  |  | Suction O2 Concentration Setting | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | VENT: sO2Suc | ['MP70', 'MX800', 'X3']
4968 |  |  | PCV Control Pressure Setting | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | VENT: sPin | ['MP70', 'MX800', 'X3']
4969 |  |  | PCV Inspiratory Time Setting | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | VENT: sPVinT | ['MP70', 'MX800', 'X3']
4970 |  |  | PCV Apnea Respiration Rate Setting | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | VENT: sAPVRR | ['MP70', 'MX800', 'X3']
4971 |  |  | PCV Apnea Inspiratory Time Setting | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | VENT: sAPVTi | ['MP70', 'MX800', 'X3']
4972 |  |  | PCV Apnea O2 Setting | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | VENT: sAPVO2 | ['MP70', 'MX800', 'X3']
4973 |  |  | PCV Apnea High Airway Pressure Setting | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | VENT: sAPVhP | ['MP70', 'MX800', 'X3']
4975 |  |  | Inspiratory Component of I:E Ratio Setting | - | Pressure Ventilation I component of I:E Ratio setting | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sPVI | ['MP70', 'MX800', 'X3']
4976 |  |  | Expiratory Component of I:E Ratio Setting | - | Pressure Ventilation E component of I:E Ratio setting | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sPVE | ['MP70', 'MX800', 'X3']
4977 |  |  | Inspiratory Component of Apnea Ventilation I:E Ratio Setting | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sAPVI | ['MP70', 'MX800', 'X3']
4978 |  |  | Expiratory Component of Apnea Ventilation I:E Ratio Setting | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sAPVE | ['MP70', 'MX800', 'X3']
4979 |  |  | Setting: Pressure Ventilation Control | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
4981 |  |  | Cycle Time Setting | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sCycTi | ['MP70', 'MX800', 'X3']
4982 |  |  | IPPV Ventilation Respiration Rate Setting | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | VENT: sIPPV | ['MP70', 'MX800', 'X3']
4983 |  |  | IMV Ventilation Respiration Rate Setting | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | VENT: sIMV | ['MP70', 'MX800', 'X3']
4984 |  |  | Pressure Support PEEP Setting | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | VENT: sSPEEP | ['MP70', 'MX800', 'X3']
4987 |  |  | Apnea Ventilation Delay | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
4989 |  |  | O2 Monitoring Mode | - | [] | Double | DeviceInfo | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | VENT: sO2Mon | ['MP70', 'MX800', 'X3']
4990 |  |  | O2 Calibration Mode | - | [] | Double | DeviceInfo | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | VENT: sO2Cal | ['MP70', 'MX800', 'X3']
4993 |  |  | Inspiratory Time Setting | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | VENT: sInsTi | ['MP70', 'MX800', 'X3']
4994 |  |  | Expiratory Time Setting | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | VENT: sExpTi | ['MP70', 'MX800', 'X3']
4996 |  |  | Alarm Percentage on Rise Time Setting | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | VENT: sALMRT | ['MP70', 'MX800', 'X3']
4998 |  |  | Ventilator Flow Setting | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | VENT: sFlow | ['MP70', 'MX800', 'X3']
5000 |  |  | Low Inspiratory Pressure Setting | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sPmin | ['MP70', 'MX800', 'X3']
5001 |  |  | HFV Flow Setting | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sHFVFl | ['MP70', 'MX800', 'X3']
5002 |  |  | HFV Respiration Rate Setting | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sHFVRR | ['MP70', 'MX800', 'X3']
5005 |  |  | Mechanical Ventilation Activated Status | - | Controlled Mechanical Ventilation Setting | Double | DeviceInfo | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | VENT: sCMV | ['MP70', 'MX800', 'X3']
5006 |  |  | SIMV Mode | - | Synchronized Intermittent Mandatory Ventilation Setting | Double | DeviceInfo | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | VENT: sSIMV | ['MP70', 'MX800', 'X3']
5007 |  |  | Mandatory Inspired Minute Volume Setting | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | VENT: sMMV | ['MP70', 'MX800', 'X3']
5009 |  |  | Infusion Pump Delivery Rate Setting | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sDRate | ['MP70', 'MX800', 'X3']
5011 |  |  | Airway Respiration Rate Setting | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sAWRR | ['MP70', 'MX800', 'X3']
5012 |  |  | Trigger Flow/Volume Setting | - | This is the device's measurement with label="NLS_NOM_SETT_VENT_VOL_LUNG_TRAPD" and physio_id="NOM_SETT_VENT_VOL_LUNG_TRAPD". | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | VENT: sTrVol | ['MP70', 'MX800', 'X3']
5013 |  |  | Apnea Alarm Delay | - | [] | Double | DeviceInfo | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sAADel | ['MP70', 'MX800', 'X3']
5014 |  |  | HFV Amplitude Setting | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sHFVAm | ['MP70', 'MX800', 'X3']
5015 |  |  | Volume Alarm Delay | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sMVDel | ['MP70', 'MX800', 'X3']
5016 |  |  | Pressure Increase Rate Setting | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | VENT: sPincR | ['MP70', 'MX800', 'X3']
5017 |  |  | Volume Warning Setting | - | Delivered by the Evita 2 VueLink Driver. | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | VENT: sVmax | ['MP70', 'MX800', 'X3']
5026 |  |  | Airway Inspiratory Pressure Lower Alarm Limit | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | AlarmLimit: loPmax | ['MP70', 'MX800', 'X3']
5028 |  |  | Apnea Alarm Delay Setting | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
5030 |  |  | Applied Tidal Volume Setting | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sTVap | ['MP70', 'MX800', 'X3']
5031 |  |  | Assist Pressure Sensitivity Setting | - | Assist Sensitivity. Used by the Bear 1000 ventilator. | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sSens | ['MP70', 'MX800', 'X3']
5032 |  |  | High Inspiratory Pressure Setting | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
5033 |  |  | Background Flow Setting | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | VENT: sBkgFl | ['MP70', 'MX800', 'X3']
5035 |  |  | Vaporizer Agent Concentration Setting | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sAGT | ['MP70', 'MX800', 'X3']
5036 |  |  | Vaporizer Isoflurane Concentration Setting | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sISO | ['MP70', 'MX800', 'X3']
5037 |  |  | Vaporizer Enflurane Concentration Setting | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sENF | ['MP70', 'MX800', 'X3']
5038 |  |  | Vaporizer Halothane Concentration Setting | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sHAL | ['MP70', 'MX800', 'X3']
5039 |  |  | Vaporizer Desflurane Concentration Setting | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sDES | ['MP70', 'MX800', 'X3']
5040 |  |  | Vaporizer Sevoflurane Concentration Setting | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sSEV | ['MP70', 'MX800', 'X3']
5041 |  |  | Gas Measurement Sensor Position | - | [] | Double | DeviceInfo | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | VENT: sGasPr | ['MP70', 'MX800', 'X3']
5042 |  |  | O2 Measurement Sensor Position | - | [] | Double | DeviceInfo | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | VENT: sO2Pr | ['MP70', 'MX800', 'X3']
5045 |  |  | Incubator/Warmer Temperature Setting | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | [] | ['MP70', 'MX800', 'X3']
5046 |  |  | Preset Period of UrVol Setting | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sUrTi | ['MP70', 'MX800', 'X3']
5050 |  |  | Low Time Setting | - | Setting variable for Evita 4 Airway Pressure Release Ventilation Mode. | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | VENT: sTlow | ['MP70', 'MX800', 'X3']
5051 |  |  | High Time Setting | - | Setting variable for Evita 4 Airway Pressure Release Ventilation Mode. | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | VENT: sThigh | ['MP70', 'MX800', 'X3']
5052 |  |  | Low Pressure Setting | - | Setting variable for Evita 4 Airway Pressure Release Ventilation Mode. | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | VENT: sPlow | ['MP70', 'MX800', 'X3']
5053 |  |  | High Pressure Setting | - | Setting variable for Evita 4 Airway Pressure Release Ventilation Mode. | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | VENT: sPhigh | ['MP70', 'MX800', 'X3']
5054 |  |  | Volume Assist Level Setting | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | VENT: sVolas | ['MP70', 'MX800', 'X3']
5055 |  |  | CPAP Flow Assist Level Setting | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | VENT: sFlas | ['MP70', 'MX800', 'X3']
5056 |  |  | Preset Stimulation Current Setting | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sCurnt | ['MP70', 'MX800', 'X3']
5057 |  |  | Preset Stimulation Charge Setting | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sChrge | ['MP70', 'MX800', 'X3']
5058 |  |  | Preset Stimulation Impulse Duration Setting | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sPulsD | ['MP70', 'MX800', 'X3']
5059 |  |  | NMT TOF Repetition Time | - | Preset Train of Four (Slow TOF) Repetition Time Setting | Double | DeviceInfo | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | NMT: sRepTi | ['MP70', 'MX800', 'X3']
5060 |  |  | Panting Limit Setting | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | VENT: sfmax | ['MP70', 'MX800', 'X3']
5065 |  |  | Airway Inspiratory Pressure Upper Alarm Limit | - | Alarm Limit: High Pressure | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | AlarmLimit: highP | ['MP70', 'MX800', 'X3']
5066 |  |  | Sigh Volume | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Vsigh | ['MP70', 'MX800', 'X3']
5069 |  |  | Sustained Pressure Alarm Limit | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | AlarmLimit: sustP | ['MP70', 'MX800', 'X3']
5079 |  |  | Mixed Venous O2 Saturation | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | SvO2 | ['MP70', 'MX800', 'X3']
5080 |  |  | Generic Term for the Transcutaneous Gases | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | TCUT: tcGas | ['MP70', 'MX800', 'X3']
5083 |  |  | Intrinsic PEEP | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | iPEEP | ['MP70', 'MX800', 'X3']
5086 |  |  | Positive Peak Inspiratory Pressure | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | PIP | ['MP70', 'MX800', 'X3']
5087 |  |  | Dead Space to Tidal Volume Ratio | ['%', '-'] | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. Ratio of serial dead space volume to expiratory tidal volume, Vds/VTe. For Evita Infinity V500, Babylog VN500: Dead space ventilation Vds/VTe. | Double | VitalSigns | ['IntelliVue MX800\xa0', 'Evita Infinity V500 [MEDIBUS.X]\xa0', 'IntelliVue X3\xa0', 'Evita Infinity V800\xa0', 'IntelliVue MP70\xa0'] | Vd/Vt | ['X3', 'MX800', 'Evita V800', 'MP70', 'Evita V500']
5088 |  |  | Lung Volume Trapped | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | VENT: TrpVol | ['MP70', 'MX800', 'X3']
5089 |  |  | Lung Vital Capacity | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | VC | ['MP70', 'MX800', 'X3']
5090 |  |  | Generic Lung Compliance | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | COMP | ['MP70', 'MX800', 'X3']
5092 |  |  | EEG Electromyography | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | EEG: EMG | ['MP70', 'MX800', 'X3']
5093 |  |  | Suppression Ratio | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | EEG: SR | ['MP70', 'MX800', 'X3']
5095 |  |  | EEG Peak Power Frequency | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | EEG: PPF | ['MP70', 'MX800', 'X3']
5101 |  |  | End Tidal N2 | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | N2 | ['MP70', 'MX800', 'X3']
5102 |  |  | Inspired N2 | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | N2 | ['MP70', 'MX800', 'X3']
5109 |  |  | Generic Agent Wave | - | [] | Double | Waveforms | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Waves: AGT | ['MP70', 'MX800', 'X3']
5110 |  |  | End Tidal MAC Concentration | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | MAC | ['MP70', 'MX800', 'X3']
5111 |  |  | Left Ventricular Stroke Volume Index | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | LVSWI | ['MP70', 'MX800', 'X3']
5112 |  |  | Left Ventricular Stroke Volume | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | LVSW | ['MP70', 'MX800', 'X3']
5113 |  |  | Right Ventricular Stroke Volume | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | RVSW | ['MP70', 'MX800', 'X3']
5114 |  |  | Glasgow Coma Score | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | GCS | ['MP70', 'MX800', 'X3']
5115 |  |  | SubScore of the GCS - Eye Response | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | EyeRsp | ['MP70', 'MX800', 'X3']
5116 |  |  | SubScore of the GCS - Motoric Response | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | MotRsp | ['MP70', 'MX800', 'X3']
5117 |  |  | SubScore of the GCS - Verbal Response | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | VblRsp | ['MP70', 'MX800', 'X3']
5118 |  |  | Head Circumference | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | HC | ['MP70', 'MX800', 'X3']
5119 |  |  | pH in Arterial Blood | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Ph | ['MP70', 'MX800', 'X3']
5120 |  |  | Arterial Blood Hemoglobin (HB) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Hb | ['MP70', 'MX800', 'X3']
5121 |  |  | pH in Venous Blood | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | pHv | ['MP70', 'MX800', 'X3']
5122 |  |  | Urine Natrium | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | UrNa | ['MP70', 'MX800', 'X3']
5123 |  |  | Natrium in Serum | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | SerNa | ['MP70', 'MX800', 'X3']
5124 |  |  | pH in the Blood Plasma | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | pH | ['MP70', 'MX800', 'X3']
5125 |  |  | Blood Natrium (Sodium) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Na | ['MP70', 'MX800', 'X3']
5126 |  |  | Kalium (Potassium) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | K | ['MP70', 'MX800', 'X3']
5127 |  |  | Generic Glucose Parameter | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Glu | ['MP70', 'MX800', 'X3']
5128 |  |  | Partial CO2 | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | PCO2 | ['MP70', 'MX800', 'X3']
5129 |  |  | Partial O2 | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | PO2 | ['MP70', 'MX800', 'X3']
5130 |  |  | Base Excess of Blood | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | BE | ['MP70', 'MX800', 'X3']
5131 |  |  | Percent Alveolarvenous Shunt Qs/Qt | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Qs/Qt | ['MP70', 'MX800', 'X3']
5132 |  |  | Extravascular Thermal Volume Index | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | ETVI | ['MP70', 'MX800', 'X3']
5133 |  |  | VPB Rate | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
5134 |  |  | ECG Lead V1 Wave | - | [] | Double | Waveforms | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Waves: V1 | ['MP70', 'MX800', 'X3']
5135 |  |  | ECG Lead V2 Wave | - | [] | Double | Waveforms | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Waves: V2 | ['MP70', 'MX800', 'X3']
5136 |  |  | ECG Lead V3 Wave | - | [] | Double | Waveforms | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Waves: V3 | ['MP70', 'MX800', 'X3']
5137 |  |  | ECG Lead V4 Wave | - | [] | Double | Waveforms | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Waves: V4 | ['MP70', 'MX800', 'X3']
5138 |  |  | ECG Lead V5 Wave | - | [] | Double | Waveforms | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Waves: V5 | ['MP70', 'MX800', 'X3']
5139 |  |  | ECG Lead V6 Wave | - | [] | Double | Waveforms | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Waves: V6 | ['MP70', 'MX800', 'X3']
5140 |  |  | ECG Lead I Wave | - | [] | Double | Waveforms | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Waves: I | ['MP70', 'MX800', 'X3']
5143 |  |  | ECG Lead aVR Wave | - | [] | Double | Waveforms | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Waves: aVR | ['MP70', 'MX800', 'X3']
5144 |  |  | ECG Lead aVL Wave | - | [] | Double | Waveforms | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Waves: aVL | ['MP70', 'MX800', 'X3']
5145 |  |  | ECG Lead aVF Wave | - | [] | Double | Waveforms | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Waves: aVF | ['MP70', 'MX800', 'X3']
5185 |  |  | Mean Dominant Frequency (CH3) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | MDF3 | ['MP70', 'MX800', 'X3']
5186 |  |  | Mean Dominant Frequency (CH4) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | MDF4 | ['MP70', 'MX800', 'X3']
5189 |  |  | NMT DBS Count | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | DBScnt | ['MP70', 'MX800', 'X3']
5191 |  |  | Peak Power Frequency (CH3) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | PPF3 | ['MP70', 'MX800', 'X3']
5192 |  |  | Peak Power Frequency (CH4) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | PPF4 | ['MP70', 'MX800', 'X3']
5235 |  |  | Low Water Level Alarm | unitless | Low water level alarm Possible values: 0 = On 1 = Off | Boolean | Alarms | Common | [] | Criticool
5318 |  |  | Blood Temperature Alarm | unitless | Value of "1" when alarm is active. | Boolean | Alarms | Common | [] | 5008
5374 |  |  | Expiratory Time Constant (RCexp) | s | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. | Double | VitalSigns | ['Evita Infinity V500 [MEDIBUS.X]\xa0', 'Evita Infinity V800\xa0'] | [] | ['Evita V800', 'Evita V500']
5507 |  |  | Remaining Time Counter | - | Time to Elapse Counter used by the i-Stat module. | Double | OtherMeasuredData | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Elapse | ['MP70', 'MX800', 'X3']
5508 |  |  | Current Volume Infused | mL | [] | UShort | OtherMeasuredData | Common | [] | Applix Smart
5598 |  |  | APRV Low Pressure Synchronisation Time Setting | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sTloS | ['MP70', 'MX800', 'X3']
5603 |  |  | APRV High Pressure Synchronisation Time Setting | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sThiS | ['MP70', 'MX800', 'X3']
5614 |  |  | Acoustic Respiration Rate | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | RRac | ['MP70', 'MX800', 'X3']
5616 |  |  | Regional O2 Saturation (rSO2 2) | - | [] | Double | VitalSigns | Common | [] | ['MP20', 'MP70', 'MX800', 'X3']
5617 |  |  | Regional O2 Saturation (rSO2 3) | - | [] | Double | VitalSigns | Common | [] | ['MP20', 'MP70', 'MX800', 'X3']
5620 |  |  | Regional O2 Saturation (rSO2 4) | - | [] | Double | VitalSigns | Common | [] | ['MP20', 'MP70', 'MX800', 'X3']
5632 |  |  | Regional O2 Saturation (rSO2 5) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | SO2 5 | ['MP70', 'MX800', 'X3']
5633 |  |  | Regional O2 Saturation (rSO2 6) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | SO2 6 | ['MP70', 'MX800', 'X3']
5637 |  |  | O2 Delivery Index (Thermodilution) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | iDO2I | ['MP70', 'MX800', 'X3']
5659 |  |  | Noninvasive Venous Hemoglobin (SpHbv) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | SpHbv | ['MP70', 'MX800', 'X3']
5660 |  |  | Noninvasive Arterial Hemoglobin (SpHba) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | SpHba | ['MP70', 'MX800', 'X3']
5661 |  |  | Generic Pulse Rate | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | iAvgPR | ['MP70', 'MX800', 'X3']
5680 |  |  | Corrected Total Hemoglobin | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | cSpHb | ['MP70', 'MX800', 'X3']
5681 |  |  | Corrected Arterial Total Hemoglobin | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | cSpHb | ['MP70', 'MX800', 'X3']
5682 |  |  | Corrected Venous Total Hemoglobin | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | cSpHbv | ['MP70', 'MX800', 'X3']
5694 |  |  | Systemic Vascular Resistance (Thermodilution) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | iSVR | ['MP70', 'MX800', 'X3']
5695 |  |  | Systemic Vascular Resistance Index (Thermodilution) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | iSVRI | ['MP70', 'MX800', 'X3']
5707 |  |  | Pulse Rate (Abdominal) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Pulse | ['MP70', 'MX800', 'X3']
5719 |  |  | ECG Lead V3R | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | ECG: ST | ['MP70', 'MX800', 'X3']
5720 |  |  | ECG Lead V4R | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | ECG: ST | ['MP70', 'MX800', 'X3']
5721 |  |  | ECG Lead V5R | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | ECG: ST | ['MP70', 'MX800', 'X3']
5728 |  |  | EtCO2 Phase III Increment | - | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. Possible units: mmHg/L, kPa/L, Vol%/L. Increase of measured CO2 value in phase III of the capnogram, CO2 slope. | Double (Multiset) | VitalSigns | ['Evita Infinity V500 [MEDIBUS.X]\xa0', 'Evita Infinity V800\xa0'] | [] | ['Evita V800', 'Evita V500']
5763 |  |  | Inspired NO | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | NO | ['MP70', 'MX800', 'X3']
5767 |  |  | ECG Lead V9 | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | ECG: ST | ['MP70', 'MX800', 'X3']
5774 |  |  | ECG Lead V7 | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | ECG: ST | ['MP70', 'MX800', 'X3']
5775 |  |  | ECG Lead V8 | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | ECG: ST | ['MP70', 'MX800', 'X3']
5807 |  |  | Apnea Ventilation I:E Ratio Setting | unitless | This variable is associated with ChannelID identifier. This variable is calculated from variable 4977 and 4978, which is sent as a string in the format "1:xx.xx". | String | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Calculated Variables | ['MP70', 'MX800', 'X3']
5815 |  |  | Spontaneous I:E Ratio | unitless | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. Format: 1:x if E > I, x:1 if I > E, or 1:1 if I = E. The useless leading or ending “0” or “.” is not transmitted, and the maximum digits after the floating point is 2. For Evita Infinity V500, Babylog VN500: Format: 1:x if E > I, x:1 if I > E, or 1:1 if I = E. The useless leading or ending “0” or “.” is not transmitted, and the maximum digits after the floating point is 2. | String | VitalSigns | ['Evita Infinity V500 [MEDIBUS.X]\xa0', 'Evita Infinity V800\xa0'] | [] | ['Evita V800', 'Evita V500']
5817 |  |  | Assisted Respiration Rate | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Rrasst | ['MP70', 'MX800', 'X3']
5818 |  |  | Correlation Coefficient Squared | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | r2 | ['MP70', 'MX800', 'X3']
5837 |  |  | ECG Lead V6R | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | ECG: ST | ['MP70', 'MX800', 'X3']
5845 |  |  | Total Replacement Cumulated Volume | mL | Cumulated total substitution volume (pre + post). | Long | VitalSigns | Common | [] | 5008
5846 |  |  | Single Needle Blood Pump Rate | mL/min | [] | Long | VitalSigns | Common | [] | 5008
5850 |  |  | Anticoagulant Cumulated Volume | mL | [] | Double | VitalSigns | Common | [] | ['5008', 'Artis']
5853 |  |  | Delivered Treatment Effectiveness (Kt/V) | unitless | [] | Double | VitalSigns | Common | [] | ['5008', 'Artis']
5858 |  |  | Disinfect Mode Active Status | unitless | Consult detailed help for more information on this variable | UByte | DeviceInfo | Common | [] | 5008
5859 |  |  | Dialysis Mode Active Status | unitless | Consult detailed help for more information on this variable | Long | DeviceInfo | Common | [] | 5008
5861 |  |  | Effective Blood Volume Processed (Kt) | L | [] | Double | VitalSigns | Common | [] | ['5008', 'Artis']
5888 |  |  | Dialysate pH | pH | [] | Double | VitalSigns | Common | [] | Artis
5889 |  |  | Ionic Clearance | mL/min | Ionic Dialysance | Long | VitalSigns | Common | [] | Artis
5890 |  |  | Ionic Mass Balance | mmol/L | [] | Long | VitalSigns | Common | [] | Artis
5892 |  |  | Bicarbonate Conductivity (8.4%) | mS/cm | [] | Double | VitalSigns | Common | [] | Artis
5899 |  |  | Actual Anticoagulant Rate | mL/h | [] | Double | VitalSigns | Common | [] | ['5008', 'Artis']
5905 |  |  | Therapy Processed Percentage | % | [] | Byte | VitalSigns | Common | [] | 5008
5949 |  |  | Backup Respiration Rate Setting | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sRRbak | ['MP70', 'MX800', 'X3']
5958 |  |  | Inspired Expired Pressure Difference | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sPi-e | ['MP70', 'MX800', 'X3']
5959 |  |  | Backup Inspired Expired Pressure Difference Setting | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sPi-eB | ['MP70', 'MX800', 'X3']
5961 |  |  | PEEP/CPAP Alarm | unitless | [] | Boolean | Alarms | Common | [] | Monnal T60
6063 |  |  | Central Venous Pressure Lower Alarm Limit | mmHg | This variable is also sent as alarm limit of variableID 3248. | Long | Settings | Multifiltrate PRO | Dialysis Settings | Multifiltrate PRO
6064 |  |  | Central Venous Pressure Upper Alarm Limit | mmHg | This variable is also sent as alarm limit of variableID 3248. | Long | Settings | Multifiltrate PRO | Dialysis Settings | Multifiltrate PRO
6124 |  |  | Femoral Arterial Pressure Mean | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | IBP: FAP | ['MP70', 'MX800', 'X3']
6125 |  |  | Femoral Arterial Pressure Systolic | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | IBP: FAP | ['MP70', 'MX800', 'X3']
6126 |  |  | Femoral Arterial Pressure Diastolic | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | IBP: FAP | ['MP70', 'MX800', 'X3']
6153 |  |  | ET Tube Blocked Alarm | unitless | [] | Boolean | Alarms | Common | [] | Monnal T60
6322 |  |  | ECG Lead V7 Wave | - | [] | Double | Waveforms | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Waves: V7 | ['MP70', 'MX800', 'X3']
6323 |  |  | ECG Lead V8 Wave | - | [] | Double | Waveforms | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Waves: V8 | ['MP70', 'MX800', 'X3']
6324 |  |  | ECG Lead V9 Wave | - | [] | Double | Waveforms | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Waves: V9 | ['MP70', 'MX800', 'X3']
6326 |  |  | Airway Aux Wave | - | [] | Double | Waveforms | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Waves: AWPx | ['MP70', 'MX800', 'X3']
6328 |  |  | Blood Plasma Inlet Pressure | mmHg | [] | Short | VitalSigns | Common | [] | 5008
6331 |  |  | Fetal DECG Wave | - | [] | Double | Waveforms | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Waves: DECG | ['MP70', 'MX800', 'X3']
6334 |  |  | Temperature Module General Status | unitless | Possible values: 0=off, 1=on | Boolean | DeviceInfo | Common | [] | 5008
6339 |  |  | Patient BSA Formula | unitless | Consult detailed help for more information on this variable | Double | PatientInfo | Common | Patient Demographics | ['MP20', 'MP70', 'MX800', 'X3']
6343 |  |  | Platelets (Thrombocyte Count) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Plts | ['MP70', 'MX800', 'X3']
6388 |  |  | Endotracheal Tube Type | - | [] | Double | DeviceInfo | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sTube | ['MP70', 'MX800', 'X3']
6390 |  |  | Minute Volume Upper Alarm Limit | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | AlarmLimit: highMV | ['MP70', 'MX800', 'X3']
6394 |  |  | Pulse Rate from Plethysmogram | - | [] | Double | VitalSigns | Common | Pulse | ['MP20', 'MP70', 'MX800', 'X3']
6396 |  |  | ECG Ectopic Status | - | Consult detailed help for more information on this variable | Double | VitalSigns | Common | Ectopic Status (on IntelliVue) | ['MP20', 'MP70', 'MX800', 'X3']
6399 |  |  | ST Lead ES | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | ECG: ST | ['MP70', 'MX800', 'X3']
6400 |  |  | ST Lead AS | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | ECG: ST | ['MP70', 'MX800', 'X3']
6401 |  |  | ST Lead AI | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | ECG: ST | ['MP70', 'MX800', 'X3']
6402 |  |  | ST Index | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | ECG: Stindx | ['MP70', 'MX800', 'X3']
6409 |  |  | Pulse Pressure | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | PP | ['MP70', 'MX800', 'X3']
6427 |  |  | Last Desinfection Time | unitless | Format: YYYYMMDDhhmm | String | DeviceInfo | Common | [] | 5008
6429 |  |  | Last Therapy End Time | unitless | Format: YYYYMMDDhhmm | String | DeviceInfo | Common | [] | 5008
6430 |  |  | Dialysis Pump Stop Status | unitless | Consult detailed help for more information on this variable | UShort | DeviceInfo | Common | [] | 5008
6450 |  |  | Alveolar Dead Space | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Vdalv | ['MP70', 'MX800', 'X3']
6451 |  |  | Tympanic Temperature | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Ttymp | ['MP70', 'MX800', 'X3']
6452 |  |  | Injectate Volume (CO) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | InjVol | ['MP70', 'MX800', 'X3']
6453 |  |  | SpO2 (Telemetry) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | %SpO2T | ['MP70', 'MX800', 'X3']
6454 |  |  | Pulse Rate (Telemetry) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | PulseT | ['MP70', 'MX800', 'X3']
6455 |  |  | Intracranial Pressure ICP1 (Systolic) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | IBP: IC2 | ['MP70', 'MX800', 'X3']
6456 |  |  | Intracranial Pressure ICP1 (Diastolic) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | IBP: IC1 | ['MP70', 'MX800', 'X3']
6457 |  |  | Intracranial Pressure ICP1 (Mean) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | IBP: IC3 | ['MP70', 'MX800', 'X3']
6458 |  |  | Intracranial Pressure ICP2 Systolic | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | IBP: IC2 | ['MP70', 'MX800', 'X3']
6459 |  |  | Intracranial Pressure ICP2 Diastolic | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | IBP: IC2 | ['MP70', 'MX800', 'X3']
6460 |  |  | Intracranial Pressure ICP2 Mean | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | IBP: IC2 | ['MP70', 'MX800', 'X3']
6461 |  |  | Brachial Arterial Pressure (BA) Systolic | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | IBP: BAP | ['MP70', 'MX800', 'X3']
6462 |  |  | Brachial Arterial Pressure (BA) Diastolic | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | IBP: BAP | ['MP70', 'MX800', 'X3']
6463 |  |  | Brachial Arterial Pressure (BA) Mean | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | IBP: BAP | ['MP70', 'MX800', 'X3']
6464 |  |  | Vesical Temperature (Bladder) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Tvesic | ['MP70', 'MX800', 'X3']
6465 |  |  | Cerebral Temperature | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Tcereb | ['MP70', 'MX800', 'X3']
6466 |  |  | ST Level for Lead I | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | ECG: Rf-I | ['MP70', 'MX800', 'X3']
6467 |  |  | ST Level for Lead II | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | ECG: Rf-II | ['MP70', 'MX800', 'X3']
6468 |  |  | ST Level for Lead III | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | ECG: Rf-III | ['MP70', 'MX800', 'X3']
6469 |  |  | ST Level for Lead aVR | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | ECG: Rf-aVR | ['MP70', 'MX800', 'X3']
6470 |  |  | ST Level for Lead aVL | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | ECG: Rf-aVL | ['MP70', 'MX800', 'X3']
6471 |  |  | ST Level for Lead aVF | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | ECG: Rf-aVF | ['MP70', 'MX800', 'X3']
6472 |  |  | ST Level for Lead V1 | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | ECG: Rf-V1 | ['MP70', 'MX800', 'X3']
6473 |  |  | ST Level for Lead V2 | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | ECG: Rf-V2 | ['MP70', 'MX800', 'X3']
6474 |  |  | ST Level for Lead V3 | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | ECG: Rf-V3 | ['MP70', 'MX800', 'X3']
6475 |  |  | ST Level for Lead V4 | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | ECG: Rf-V4 | ['MP70', 'MX800', 'X3']
6476 |  |  | ST Level for Lead V5 | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | ECG: Rf-V5 | ['MP70', 'MX800', 'X3']
6477 |  |  | ST Level for Lead V6 | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | ECG: Rf-V6 | ['MP70', 'MX800', 'X3']
6478 |  |  | Prothrombin Time (International Normalized Ratio) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | PT INR | ['MP70', 'MX800', 'X3']
6479 |  |  | PSV Tidal Volume | - | Tidal Volume in Pressure Support Ventilation Mode | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | TVPSV | ['MP70', 'MX800', 'X3']
6480 |  |  | Negative Inspiratory Force Index | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
6481 |  |  | A-Line ARX Index | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | EEG: AAI | ['MP70', 'MX800', 'X3']
6484 |  |  | Impedance Cardiography (ICG) Wave | - | [] | Double | Waveforms | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Waves: ICG | ['MP70', 'MX800', 'X3']
6485 |  |  | Thoracic Fluid Content | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | TFC | ['MP70', 'MX800', 'X3']
6486 |  |  | Thoracic Fluid Content Index | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | TFI | ['MP70', 'MX800', 'X3']
6487 |  |  | Accelerated Cardiac Index | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | ACI | ['MP70', 'MX800', 'X3']
6488 |  |  | Heart Contractility Index | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | HI | ['MP70', 'MX800', 'X3']
6489 |  |  | O2 Saturation (Left) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | SO2 l | ['MP70', 'MX800', 'X3']
6490 |  |  | O2 Saturation (Right) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | SO2 r | ['MP70', 'MX800', 'X3']
6491 |  |  | Pulse Pressure Variation | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | PPV | ['MP70', 'MX800', 'X3']
6492 |  |  | Downslope Time | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
6502 |  |  | ART Pulse Rate | - | ART pulse rate | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Pulse | ['MP70', 'MX800', 'X3']
6543 |  |  | Umbilical Arterial Pressure Pulse Rate | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Pulse | ['MP70', 'MX800', 'X3']
6544 |  |  | Patient Visit Number | unitless | Patient encounter ID contains the ID of the current visit of the patient | String | PatientInfo | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Patient Demographics | ['MP70', 'MX800', 'X3']
6551 |  |  | End Tidal CO2 Pressure | - | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Evita V300, Fabius Family. This variable is associated with UpperAlarmLimit attribute, which is available for Perseus A500, Atlan, Evita Infinity V500, Babylog VN500, Savina 300, Primus, Primus IE, Apollo, Zeus IE, Oxylog 3000 plus, Evita V300. This variable is associated with LowerAlarmLimit attribute, which is available for Perseus A500, Atlan, Evita Infinity V500, Babylog VN500, Savina 300, Primus, Primus IE, Apollo, Zeus IE, Oxylog 3000 plus, Evita V300. Possible units: mmHg, kPa. end tidal CO2 etCO2 | Double | VitalSigns | ['Evita Infinity V500 [MEDIBUS.X]\xa0', 'Evita Infinity V800\xa0'] | [] | ['Evita V800', 'Evita V500']
6555 |  |  | Patient Height at Entered Time | - | [] | Double | PatientInfo | Common | Patient Demographics | ['MP20', 'MP70', 'MX800', 'X3']
6557 |  |  | Patient Weight at Entered Time | - | [] | Double | PatientInfo | Common | Patient Demographics | ['MP20', 'MP70', 'MX800', 'X3']
6567 |  |  | Patient Body Surface Area (BSA) | - | [] | Double | PatientInfo | Common | Patient Demographics | ['MP20', 'MP70', 'MX800', 'X3']
6643 |  |  | Substitution Fluid And Blood Ratio (Qi/Qb) | unitless | [] | Long | Measurements | Common | [] | Artis
6685 |  |  | Inspired O2 Setting (FiO2) | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | VENT: sFIO2 | ['MP70', 'MX800', 'X3']
6736 |  |  | Rapid Shallow Breathing Index (SBI) | - | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: Obsolete variable that may be removed in future release. Variable is in Multiset format. This variable is same as 1422 but DDI buffer and send all values from different data packages together. | Double (Multiset) | VitalSigns | Evita Infinity V500 [MEDIBUS.X] | [] | Evita V500
6739 |  |  | Backup Inspiratory Time Setting | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sInTiB | ['MP70', 'MX800', 'X3']
6744 |  |  | Pressure Support Level Above Phigh | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sdPin | ['MP70', 'MX800', 'X3']
6772 |  |  | Heater Power | - | [] | Double | DeviceInfo | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sPower | ['MP70', 'MX800', 'X3']
6861 |  |  | Capillary Refill Time | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | CRT | ['MP70', 'MX800', 'X3']
6891 |  |  | Expiratory Flow Sensor Failure Alarm | unitless | [] | Boolean | Alarms | Common | [] | Monnal T60
6935 |  |  | Estimated Patient Circuit Leakage | - | Possible units: mL/min, L/min | Double | OtherMeasuredData | ['IntelliVue X3\xa0', 'Common\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | LkPat | ['V60', 'MP70', 'MX800', 'X3']
6936 |  |  | Estimated Ventilation Total Leakage | - | [] | Double | OtherMeasuredData | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | LkTot | ['MP70', 'MX800', 'X3']
6967 |  |  | Ventilation Mode Additional Information | - | [] | Double | DeviceInfo | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sVSupp | ['MP70', 'MX800', 'X3']
6995 |  |  | Backup Plateau Time Setting | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sPltTA | ['MP70', 'MX800', 'X3']
7000 |  |  | ECG Lead V3R Wave | - | [] | Double | Waveforms | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Waves: V3R | ['MP70', 'MX800', 'X3']
7001 |  |  | ECG Lead V4R Wave | - | [] | Double | Waveforms | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Waves: V4R | ['MP70', 'MX800', 'X3']
7002 |  |  | ECG Lead V5R Wave | - | [] | Double | Waveforms | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Waves: V5R | ['MP70', 'MX800', 'X3']
7003 |  |  | ECG Lead V6R Wave | - | [] | Double | Waveforms | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Waves: V6R | ['MP70', 'MX800', 'X3']
7019 |  |  | Signal Quality Index (BIS) | - | This variable is associated with ChannelID identifier. | Double | OtherMeasuredData | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | SQI L/SQI R | ['MP70', 'MX800', 'X3']
7033 |  |  | Patient Demand High Alarm | unitless | [] | Boolean | Alarms | Common | [] | Monnal T60
7038 |  |  | High Inspired Minute Volume Alarm | unitless | [] | Boolean | Alarms | Common | [] | Monnal T60
7039 |  |  | Low Inspired Minute Volume Alarm | unitless | [] | Boolean | Alarms | Common | [] | Monnal T60
7076 |  |  | Generic Ventilation Mode | unitless | Consult detailed help for more information on this variable | String | Settings | ['Evita Infinity V500 [MEDIBUS.X]\xa0', 'Evita Infinity V800\xa0'] | [] | ['Evita V800', 'Evita V500']
7077 |  |  | Ventilation Mode Primary | unitless | Consult detailed help for more information on this variable | String | Settings | ['Evita Infinity V500 [MEDIBUS.X]\xa0', 'Evita Infinity V800\xa0'] | [] | ['Evita V800', 'Evita V500']
7078 |  |  | Ventilation Mode Options | unitless | Consult detailed help for more information on this variable | String | Settings | ['Evita Infinity V500 [MEDIBUS.X]\xa0', 'Evita Infinity V800\xa0'] | [] | ['Evita V800', 'Evita V500']
7106 |  |  | Continuous NIBP Systolic | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | cNBP SYS | ['MP70', 'MX800', 'X3']
7107 |  |  | Continuous NIBP Diastolic | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | cNBP DIA | ['MP70', 'MX800', 'X3']
7108 |  |  | Continuous NIBP Mean | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | cNBP MEAN | ['MP70', 'MX800', 'X3']
7118 |  |  | Next Intervention Time | - | [] | Double | OtherMeasuredData | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | InTtvTi | ['MP70', 'MX800', 'X3']
7315 |  |  | Total Power (CH3) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | TP3 | ['MP70', 'MX800', 'X3']
7316 |  |  | Total Power (CH4) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | TP4 | ['MP70', 'MX800', 'X3']
7318 |  |  | Variable PS Activation Status | unitless | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800. A value of 1 indicates that Variable PS (VPS) has been added to the current ventilation mode. A value of 0 indicates that Variable PS (VPS) has not been added to the current ventilation mode. Text message displayed on device screen: "VariablePS" | Boolean | Settings | ['Evita Infinity V500 [MEDIBUS.X]\xa0', 'Evita Infinity V800\xa0'] | [] | ['Evita V800', 'Evita V500']
7335 |  |  | Total Power (Left) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | TP L | ['MP70', 'MX800', 'X3']
7336 |  |  | Total Power (Right) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | TP R | ['MP70', 'MX800', 'X3']
7351 |  |  | Isolated Ultrafiltration Volume | mL | [] | Long | VitalSigns | Common | [] | 5008
7352 |  |  | Single Needle Stroke Volume | mL | Average single needle stroke volume | Long | VitalSigns | Common | [] | 5008
7353 |  |  | Initial Ultrafiltration Factor Setting | unitless | [] | UShort | Settings | Common | [] | 5008
7364 |  |  | Patient Operation Mode | unitless | Possible values: 1 for Invasive, 2 for Noninvasive | UByte | PatientInfo | ['Evita Infinity V500 [MEDIBUS.X]\xa0', 'Evita Infinity V800\xa0'] | [] | ['Evita V800', 'Evita V500']
7414 |  |  | PCV I:E Ratio Setting | unitless | This variable is associated with ChannelID identifier. This variable is calculated from variable 4975 and 4976, which is sent as a string in the format "1:xx.xx". | String | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Calculated Variables | ['MP70', 'MX800', 'X3']
7434 |  |  | Inspiratory CO2 (FiCO2) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | VENT: FICO2 | ['MP70', 'MX800', 'X3']
7452 |  |  | Device Operation Mode | unitless | Consult detailed help for more information on this variable | String | DeviceInfo | Common | [] | Criticool
7532 |  |  | End Tidal NO2 | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | NO2 | ['MP70', 'MX800', 'X3']
7535 |  |  | Fresh Gas Total Flow | - | [] | Double | OtherMeasuredData | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | fgFlow | ['MP70', 'MX800', 'X3']
7545 |  |  | Bolus Volume Setting | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sVBol | ['MP70', 'MX800', 'X3']
7557 |  |  | Generic NIBP Pressure | - | Available on IntelliVue devices with firmware revision G or later. | Double | VitalSigns | Common | NBP | ['MP20', 'MP70', 'MX800', 'X3']
7587 |  |  | Endotracheal Tube Diameter | - | [] | Double | DeviceInfo | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sTubeD | ['MP70', 'MX800', 'X3']
7641 |  |  | Mandatory Minute Volume | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | VENT: MMV | ['MP70', 'MX800', 'X3']
7644 |  |  | N2O Consumption | - | [] | Double | OtherMeasuredData | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | AccN2O | ['MP70', 'MX800', 'X3']
7765 |  |  | Device Power Up Time | s | Number of seconds since the Monitor device's last power-on. | Double | DeviceInfo | Common | System Information | ['MP20', 'MP70', 'MX800', 'X3']
7768 |  |  | O2 Consumption | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | AccO2 | ['MP70', 'MX800', 'X3']
7770 |  |  | Air Consumption | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | AccAir | ['MP70', 'MX800', 'X3']
7773 |  |  | Arterial Blood Pressure (ABP) | - | Available on IntelliVue devices with firmware revision G or later. Arterial Blood Pressure (NOM_PRESS_BLD_ART_ABP). | Double | VitalSigns | Common | IBP: ABP | ['MP20', 'MP70', 'MX800', 'X3']
7777 |  |  | ABP Pulse Rate | - | ABP pulse rate | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Pulse | ['MP70', 'MX800', 'X3']
7797 |  |  | Signal Noise Ratio (CCO) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | SNR | ['MP70', 'MX800', 'X3']
7803 |  |  | Spontaneous Airway Respiration Rate | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | SpAWRR | ['MP70', 'MX800', 'X3']
7804 |  |  | Synchronized Respiration Rate | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | RRsync | ['MP70', 'MX800', 'X3']
7805 |  |  | Mechanical Respiration Rate | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
7806 |  |  | Daily Urine Output | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | U/O | ['MP70', 'MX800', 'X3']
7809 |  |  | Activated Partial Thromboplastin Time Ratio | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | PTTrat | ['MP70', 'MX800', 'X3']
7810 |  |  | Global Ejection Fraction | - | This is the device's measurement with label="NLS_NOM_FRACT_EJECT" and physio_id="NOM_FRACT_EJECT". See also 8790. | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | GEF | ['MP70', 'MX800', 'X3']
7811 |  |  | Generic CO2 Parameter | - | This variable contains "PECO2, Partial expired CO2", which is measured from label="NLS_NOM_VENT_AWAY_CO2_EXP", physio_id="NOM_VENT_AWAY_CO2_EXP". | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | VENT: PECO2 | ['MP70', 'MX800', 'X3']
7835 |  |  | Ventilation Comfort Mode Flex Setting | - | [] | Double | Settings | Common | [] | V60
7858 |  |  | Suppression Ratio (BIS) | - | This variable is associated with ChannelID identifier. | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | SR L/SR R | ['MP70', 'MX800', 'X3']
7865 |  |  | Right Ventricular Stroke Work Index | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | RVSWI | ['MP70', 'MX800', 'X3']
7870 |  |  | Base Excess of Extracellular Fluid | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | BEecf | ['MP70', 'MX800', 'X3']
7872 |  |  | CO2 Partial Pressure Waveform | - | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Evita V300. Possible units: mmHg, kPa. Unit "mmHg" is available for Perseus A500, Atlan, Evita Infinity V500, Babylog VN500, Savina 300, Primus, Primus IE, Apollo, Zeus IE, Oxylog 3000 plus, Evita V300. Unit "kPa" is available for Perseus A500, Atlan, Savina 300, Primus, Primus IE, Apollo, Zeus IE, Oxylog 3000 plus. | Double | Waveforms | ['Evita Infinity V500 [MEDIBUS.X]\xa0', 'Evita Infinity V800\xa0'] | [] | ['Evita V800', 'Evita V500']
7878 |  |  | Rise Time Setting | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sRmpTi | ['MP70', 'MX800', 'X3']
7879 |  |  | Time Duration at Low Pressure (Tlow) | ['s', '-'] | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. Effective expiratory time during APRV/AutoRelease, Tlow. For Evita Infinity V500, Babylog VN500: Tlow (resulting Tlow from expiratory flow cycling). | Double | VitalSigns | ['IntelliVue MX800\xa0', 'Evita Infinity V500 [MEDIBUS.X]\xa0', 'IntelliVue X3\xa0', 'Evita Infinity V800\xa0', 'IntelliVue MP70\xa0'] | Th/TL | ['X3', 'MX800', 'Evita V800', 'MP70', 'Evita V500']
7880 |  |  | Time Duration at High Pressure (Thigh) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Th/TL | ['MP70', 'MX800', 'X3']
7904 |  |  | Spontaneous Tidal Volume | - | [] | ['Double\xa0', 'Double (Multiset)\xa0'] | VitalSigns | ['IntelliVue MX800\xa0', 'Evita Infinity V500 [MEDIBUS.X]\xa0', 'IntelliVue X3\xa0', 'Evita Infinity V800\xa0', 'IntelliVue MP70\xa0'] | VENT: SpTV | ['X3', 'MX800', 'Evita V800', 'MP70', 'Evita V500']
7905 |  |  | Cardiac Troponin l | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | cTnI | ['MP70', 'MX800', 'X3']
7906 |  |  | Cardiopulmonary Bypass Status | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | CPB | ['MP70', 'MX800', 'X3']
7907 |  |  | Cardiac Brain Natriuretic Peptide | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | BNP | ['MP70', 'MX800', 'X3']
7908 |  |  | Mean Dominant Frequency | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | EEG: MDF | ['MP70', 'MX800', 'X3']
7980 |  |  | Dialysate Conductivity Alarm | unitless | Value of "1" when alarm is active. | Boolean | Alarms | Common | [] | 5008
7988 |  |  | Plasma Conductivity | mS/cm | [] | Double | VitalSigns | Common | [] | Artis
8012 |  |  | Plateau Airway Pressure Alarm | unitless | [] | Boolean | Alarms | Monnal T60 | [] | Monnal T60
8153 |  |  | IPAP Minimum Pressure Setting | - | Possible units: cmH2O, hPa, kPa, mBar, mmHg | Double | Settings | Common | [] | V60
8189 |  |  | Cerebral State Index (CSI) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | CSI | ['MP70', 'MX800', 'X3']
8201 |  |  | EEG Amplitude | - | This variable is associated with ChannelID identifier. | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | aEEG1/aEEG2/aEEG3/aEEG4 | ['MP70', 'MX800', 'X3']
8244 |  |  | Bolus Manual Enable Status | unitless | Possible values: 0=disabled, 1=enabled | Boolean | DeviceInfo | Common | [] | 5008
8261 |  |  | Patient Expiratory Port Type | unitless | Consult detailed help for more information on this variable | Long | DeviceInfo | Common | [] | V60
8262 |  |  | Plasma Filtration Fraction Setting | % | [] | UShort | Settings | Common | [] | 5008
8301 |  |  | Nebulizer Error | unitless | [] | Boolean | Alarms | Monnal T60 | [] | Monnal T60
8339 |  |  | Patient Middle Names | unitless | The Middle Name attribute contains the middle name of the patient. | String | PatientInfo | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Patient Demographics | ['MP70', 'MX800', 'X3']
8340 |  |  | Patient Family Name | unitless | [] | String | PatientInfo | Common | [] | 5008
8350 |  |  | Ao Pulse Rate | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Pulse | ['MP70', 'MX800', 'X3']
8351 |  |  | PAP Pulse Rate | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Pulse | ['MP70', 'MX800', 'X3']
8352 |  |  | CVP Pulse Rate | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Pulse | ['MP70', 'MX800', 'X3']
8353 |  |  | RAP Pulse Rate | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Pulse | ['MP70', 'MX800', 'X3']
8354 |  |  | LAP Pulse Rate | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Pulse | ['MP70', 'MX800', 'X3']
8355 |  |  | ICP Pulse Rate | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Pulse | ['MP70', 'MX800', 'X3']
8356 |  |  | Umbilical Venous Pressure Pulse Rate | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Pulse | ['MP70', 'MX800', 'X3']
8357 |  |  | FAP Pulse Rate | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Pulse | ['MP70', 'MX800', 'X3']
8358 |  |  | BAP Pulse Rate | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Pulse | ['MP70', 'MX800', 'X3']
8359 |  |  | IC1 Pulse Rate | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Pulse | ['MP70', 'MX800', 'X3']
8360 |  |  | Intracranial Pressure ICP2 Pulse Rate | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Pulse | ['MP70', 'MX800', 'X3']
8361 |  |  | Pulse Rate (P1) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Pulse | ['MP70', 'MX800', 'X3']
8362 |  |  | Pulse Rate (P2) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Pulse | ['MP70', 'MX800', 'X3']
8363 |  |  | Pulse Rate (P3) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Pulse | ['MP70', 'MX800', 'X3']
8364 |  |  | Pulse Rate (P4) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Pulse | ['MP70', 'MX800', 'X3']
8365 |  |  | Pulse Rate from Plethysmogram (Left) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Pulse | ['MP70', 'MX800', 'X3']
8366 |  |  | Generic Message Code | unitless | This variable is in multiset format. This variable, along with variable 8696, contains text message code and alarm text respectively. They need to be interpreted together. | UShort (Multiset) | Messages | ['Evita Infinity V500 [MEDIBUS.X]\xa0', 'Evita Infinity V800\xa0'] | Alarm And Text Message | ['Evita V800', 'Evita V500']
8367 |  |  | Pleth Wave (Right) | - | [] | Double | Waveforms | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Waves: PLETHr | ['MP70', 'MX800', 'X3']
8368 |  |  | Pleth Wave (Left) | - | [] | Double | Waveforms | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Waves: PLETHl | ['MP70', 'MX800', 'X3']
8369 |  |  | Calculated SO2 | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | CALC: 'SO2 | ['MP70', 'MX800', 'X3']
8370 |  |  | Airway Respiration Rate (Ohmeda) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | VENT: RRaw | ['MP70', 'MX800', 'X3']
8371 |  |  | Airway Respiration Rate Setting (Ohmeda) | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | VENT: sRRaw | ['MP70', 'MX800', 'X3']
8372 |  |  | Fresh Gas O2 Concentration | - | [] | Double | OtherMeasuredData | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | fgO2 | ['MP70', 'MX800', 'X3']
8373 |  |  | Manually entered FIO2 or Flow | - | This variable contains measurement from label "NLS_NOM_CONC_AWAY_O2" and observed value "NOM_CONC_AWAY_O2". | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | O2 | ['MP70', 'MX800', 'X3']
8374 |  |  | Generic O2 Parameter | - | This is the device's measurement with label="NLS_NOM_SETT_CONC_AWAY_O2" and physio_id="NOM_CONC_AWAY_O2". | Double | [] | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sO2 | ['MP70', 'MX800', 'X3']
8375 |  |  | Fresh Gas O2 Flow Setting | - | Mixer parameter. | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sfgO2 | ['MP70', 'MX800', 'X3']
8376 |  |  | Airway Pressure Numeric Value (Ohmeda) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | AWP | ['MP70', 'MX800', 'X3']
8377 |  |  | Inspiratory Peak Flow | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | PIF | ['MP70', 'MX800', 'X3']
8378 |  |  | Maximum Pressure During Breathing Cycle | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | VENT: Pmax | ['MP70', 'MX800', 'X3']
8380 |  |  | Peak Inspiratory Flow Setting | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sPIF | ['MP70', 'MX800', 'X3']
8381 |  |  | Airway Maximum Pressure Setting | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | VENT: sPmax | ['MP70', 'MX800', 'X3']
8382 |  |  | Peak Inspiratory Pressure Setting | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sPIP | ['MP70', 'MX800', 'X3']
8383 |  |  | Setting: Inspiratory Pressure (Ohmeda) | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
8384 |  |  | Mean Airway Pressure (Ohmeda) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
8385 |  |  | EEG Wave (Left) | - | [] | Double | Waveforms | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Waves: LT EEG | ['MP70', 'MX800', 'X3']
8386 |  |  | EEG Wave (Right) | - | [] | Double | Waveforms | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Waves: RT EEG | ['MP70', 'MX800', 'X3']
8387 |  |  | % of Power in the Delta Band - 1st EEG | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | EEG: Prcnt1 | ['MP70', 'MX800', 'X3']
8388 |  |  | % of Power in the Delta Band - 2nd EEG | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | EEG: Prcnt2 | ['MP70', 'MX800', 'X3']
8389 |  |  | Percent Delta - Left Side | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | EEG: LT %DL | ['MP70', 'MX800', 'X3']
8390 |  |  | Percent Delta - Right Side | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | EEG: RT %DL | ['MP70', 'MX800', 'X3']
8391 |  |  | % of Power in the Theta Band (1st EEG) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | EEG: Prcnt1 | ['MP70', 'MX800', 'X3']
8392 |  |  | % of Power in the Theta Band (2nd EEG) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | EEG: Prcnt2 | ['MP70', 'MX800', 'X3']
8393 |  |  | Percent Theta - Left Side | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | EEG: LT %TH | ['MP70', 'MX800', 'X3']
8394 |  |  | Percent Theta - Right Side | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | EEG: RT %TH | ['MP70', 'MX800', 'X3']
8395 |  |  | Generic Calculation Constant | - | This variable contains measurement from Philips label "NLS_NOM_NUM_CALC_CONST" and observed value "NOM_METRIC_NOS". See also variable 8493. | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | CompCt | ['MP70', 'MX800', 'X3']
8396 |  |  | Pressure Ventilation Control Pressure | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | VENT: PVcP | ['MP70', 'MX800', 'X3']
8397 |  |  | Right Heart Ejection Fraction | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | REF | ['MP70', 'MX800', 'X3']
8398 |  |  | White Blood Count (Leucocyte Count) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | WBC | ['MP70', 'MX800', 'X3']
8399 |  |  | Red Blood Count (Erithrocyte Count) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | RBC | ['MP70', 'MX800', 'X3']
8400 |  |  | Mean Corpuscular Volume | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | MCV | ['MP70', 'MX800', 'X3']
8401 |  |  | Corpuscular Mean Hemoglobin Content | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | MCH | ['MP70', 'MX800', 'X3']
8402 |  |  | Ferrum | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Fe | ['MP70', 'MX800', 'X3']
8403 |  |  | Fetal Hemoglobin | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | HbF | ['MP70', 'MX800', 'X3']
8404 |  |  | MetHemoglobin | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Met-Hb | ['MP70', 'MX800', 'X3']
8406 |  |  | Expiratory Peak Flow | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | PEF | ['MP70', 'MX800', 'X3']
8407 |  |  | Expiratory Peak Flow | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | eeFlow | ['MP70', 'MX800', 'X3']
8408 |  |  | The V stands for Volume | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
8409 |  |  | The R stands for Rate | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
8410 |  |  | Adjusted PaCO2 (at Patient Temperature) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | &PaCO2 | ['MP70', 'MX800', 'X3']
8411 |  |  | Incubator Heating Power | - | Power Required to set Air&Pat Temp in the Incubator | Double | OtherMeasuredData | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | INCUBATOR: Power | ['MP70', 'MX800', 'X3']
8412 |  |  | Actual Patient Age | - | Actual patient age, provided by a module, measured in years. Please note that this Patient Age will be available only when the corresponding module is connected to the Monitor. Variable 169 contains the patient age as provided by the Monitor itself (calculated from the birth date entered on the Monitor device.) | Double | PatientInfo | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Age | ['MP70', 'MX800', 'X3']
8413 |  |  | The P stands for Pressure | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
8414 |  |  | PCV Apnea Control Pressure Setting | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | VENT: sAPVcP | ['MP70', 'MX800', 'X3']
8415 |  |  | Fresh Gas Air Flow Setting | - | Mixer parameter. | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sfgAir | ['MP70', 'MX800', 'X3']
8416 |  |  | PEEP/CPAP Lower Alarm Limit | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | AlarmLimit: loPEEP | ['MP70', 'MX800', 'X3']
8417 |  |  | Tidal Volume Upper Alarm Limit | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | AlarmLimit: highTV | ['MP70', 'MX800', 'X3']
8418 |  |  | T-Tre Temperature difference between Temp and Trect | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
8419 |  |  | T-Tbl Temperature difference between Temp and Tblood | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
8420 |  |  | T-Tco Temperature difference between Temp and Tcore | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
8421 |  |  | T-Tsk Temperature difference between Temp and Tskin | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
8422 |  |  | T-Tes Temperature difference between Temp and Tesoph | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
8423 |  |  | T-Tna Temperature difference between Temp and Tnaso | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
8424 |  |  | T-Tar Temperature difference between Temp and Tart | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
8425 |  |  | T-Tve Temperature difference between Temp and Tven | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
8426 |  |  | T-Taw Temperature difference between Temp and Tairwy | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
8427 |  |  | Tre-T Temperature difference between Trect and Temp | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
8428 |  |  | Tre-bl Temperature difference between Trect and Tblood | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
8429 |  |  | Tre-sk Temperature difference between Trect and Tskin | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
8430 |  |  | Tre-aw Temperature difference between Trect and Tairwy | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
8431 |  |  | Temperature difference between Tblood and Temp | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
8433 |  |  | Tbl-aw Temperature difference between Tblood and Tairwy | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
8434 |  |  | Tco-T Temperature difference between Tcore and Temp | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
8435 |  |  | Tco-re Temperature difference between Tcore and Trect | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
8436 |  |  | Tco-bl Temperature difference between Tcore and Tblood | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
8437 |  |  | Tco-sk Temperature difference between Tcore and Tskin | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
8438 |  |  | Tco-es Temperature difference between Tcore and Tesoph | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
8439 |  |  | Tco-na Temperature difference between Tcore and Tnaso | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
8440 |  |  | Tco-aw Temperature difference between Tcore and Tairwy | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
8441 |  |  | Tsk-T Temperature difference between Tskin and Temp | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
8442 |  |  | Tsk-bl Temperature difference between Tskin and Tblood | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
8443 |  |  | Tsk-aw Temperature difference between Tskin and Tairwy | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
8444 |  |  | Tes-T Temperature difference between Tesoph and Temp | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
8445 |  |  | Tes-re Temperature difference between Tesoph and Trect | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
8446 |  |  | Tes-bl Temperature difference between Tesoph and Tblood | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
8447 |  |  | Tes-sk Temperature difference between Tesoph and Tskin | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
8448 |  |  | Tes-na Temperature difference between Tesoph and Tnaso | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
8449 |  |  | Tes-aw Temperature difference between Tesoph and Tairwy | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
8450 |  |  | Tna-T Temperature difference between Tnaso and Temp | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
8451 |  |  | Tna-re Temperature difference between Tnaso and Trect | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
8452 |  |  | Tna-bl Temperature difference between Tnaso and Tblood | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
8453 |  |  | Tna-sk Temperature difference between Tnaso and Tskin | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
8454 |  |  | Tna-aw Temperature difference between Tnaso and Tairwy | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
8455 |  |  | Tar-T Temperature difference between Tart and Temp | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
8456 |  |  | Tar-re Temperature difference between Tart and Trect | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
8457 |  |  | Tar-bl Temperature difference between Tart and Tblood | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
8458 |  |  | Tar-co Temperature difference between Tart and Tcore | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
8459 |  |  | Tar-sk Temperature difference between Tart and Tskin | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
8460 |  |  | Tar-es Temperature difference between Tart and Tesoph | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
8461 |  |  | Tar-na Temperature difference between Tart and Tnaso | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
8462 |  |  | Tar-aw Temperature difference between Tart and Tairwy | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
8463 |  |  | Tve-T Temperature difference between Tven and Temp | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
8464 |  |  | Tve-re Temperature difference between Tven and Trect | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
8465 |  |  | Tve-bl Temperature difference between Tven and Tblood | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
8466 |  |  | Tve-co Temperature difference between Tven and Tcore | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
8467 |  |  | Tve-sk Temperature difference between Tven and Tskin | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
8468 |  |  | Tve-es Temperature difference between Tven and Tesoph | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
8469 |  |  | Tve-na Temperature difference between Tven and Tnaso | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
8470 |  |  | MAC Concentration (Awake) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | MAC | ['MP70', 'MX800', 'X3']
8471 |  |  | Tve-ar Temperature difference between Tven and Tart | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
8472 |  |  | Tve-aw Temperature difference between Tven and Tairwy | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
8473 |  |  | Taw-T Temperature difference between Tairwy and Temp | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
8474 |  |  | Taw-re Temperature difference between Tairwy and Trect | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
8475 |  |  | Taw-bl Temperature difference between Tairwy and Tblood | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
8476 |  |  | Taw-co Temperature difference between Tairwy and Tcore | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
8477 |  |  | Taw-sk Temperature difference between Tairwy and Tskin | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
8478 |  |  | Taw-es Temperature difference between Tairwy and Tesoph | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
8479 |  |  | Taw-na Temperature difference between Tairwy and Tnaso | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
8480 |  |  | Taw-ar Temperature difference between Tairwy and Tart | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
8481 |  |  | Taw-ve Temperature difference between Tairwy and Tven | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
8483 |  |  | Generic Agent 1 Wave | - | [] | Double | Waveforms | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Waves: AGT1 | ['MP70', 'MX800', 'X3']
8484 |  |  | End Tidal Generic Agent 1 | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | AGT1 | ['MP70', 'MX800', 'X3']
8485 |  |  | Inspired Generic Agent 1 | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | AGT1 | ['MP70', 'MX800', 'X3']
8486 |  |  | Generic Agent 2 Wave | - | [] | Double | Waveforms | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Waves: AGT2 | ['MP70', 'MX800', 'X3']
8487 |  |  | End Tidal Generic Agent 2 | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | AGT2 | ['MP70', 'MX800', 'X3']
8488 |  |  | Inspired Generic Agent 2 | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | AGT2 | ['MP70', 'MX800', 'X3']
8489 |  |  | Adjusted Arterial Blood pH Value | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | &pHa | ['MP70', 'MX800', 'X3']
8490 |  |  | Adjusted Venous Blood pH Value | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | &pHv | ['MP70', 'MX800', 'X3']
8491 |  |  | Adjusted PvCO2 (at Patient Temperature) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | &PvCO2 | ['MP70', 'MX800', 'X3']
8492 |  |  | Adjusted PvO2 (at Patient Temperature) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | &PvO2 | ['MP70', 'MX800', 'X3']
8493 |  |  | Generic Calculation Constant | - | This variable contains parameter from Philips label "NLS_NOM_NUM_CATHETER_CONST" and observed value "NOM_NUM_CATHETER_CONST". See also variable 8395. | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | CathCt | ['MP70', 'MX800', 'X3']
8501 |  |  | ECG Lead V Wave | - | [] | Double | Waveforms | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Waves: V | ['MP70', 'MX800', 'X3']
8502 |  |  | ECG Lead MCL Wave | - | [] | Double | Waveforms | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Waves: MCL | ['MP70', 'MX800', 'X3']
8503 |  |  | ECG Lead MCL1 Wave | - | [] | Double | Waveforms | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Waves: MCL1 | ['MP70', 'MX800', 'X3']
8510 |  |  | MAC Concentration | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | MAC | ['MP70', 'MX800', 'X3']
8511 |  |  | Pleth Wave (Telemetry) | - | [] | Double | Waveforms | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Waves: PlethT | ['MP70', 'MX800', 'X3']
8513 |  |  | NIBP Systolic (Telemetry) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | NBP_T SYS | ['MP70', 'MX800', 'X3']
8514 |  |  | NIBP Diastolic (Telemetry) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | NBP_T DIA | ['MP70', 'MX800', 'X3']
8515 |  |  | NIBP Mean (Telemetry) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | NBP_T MEAN | ['MP70', 'MX800', 'X3']
8517 |  |  | Generic Pressure 1 (P1) Systolic | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
8518 |  |  | Generic Pressure 1 (P1) Diastolic | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
8519 |  |  | Generic Pressure 1 (P1) Mean | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
8521 |  |  | Generic Pressure 2 (P2) Systolic | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
8522 |  |  | Generic Pressure 2 (P2) Diastolic | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
8523 |  |  | Generic Pressure 2 (P2) Mean | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
8525 |  |  | Generic Pressure 3 (P3) Systolic | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
8526 |  |  | Generic Pressure 3 (P3) Diastolic | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
8527 |  |  | Generic Pressure 3 (P3) Mean | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
8529 |  |  | Generic Pressure 4 (P4) Systolic | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
8530 |  |  | Generic Pressure 4 (P4) Diastolic | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
8531 |  |  | Generic Pressure 4 (P4) Mean | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
8535 |  |  | Incubator Humidity | - | [] | Double | OtherMeasuredData | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | INCUBATOR: Hum | ['MP70', 'MX800', 'X3']
8536 |  |  | Predictive Oral Temperature | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | pToral | ['MP70', 'MX800', 'X3']
8537 |  |  | Predictive Rectal Temperature | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | pTrect | ['MP70', 'MX800', 'X3']
8538 |  |  | Predictive Axillary Temperature | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | pTaxil | ['MP70', 'MX800', 'X3']
8539 |  |  | Perfusion (Telemetry) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Perf T | ['MP70', 'MX800', 'X3']
8541 |  |  | Right to Left Heart Shunt | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | RLShnt | ['MP70', 'MX800', 'X3']
8542 |  |  | Adjusted PcO2 (at Patient Temperature) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | &PCO2 | ['MP70', 'MX800', 'X3']
8543 |  |  | Adjusted Arterial PaO2 (at Patient Temperature) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | &PaO2 | ['MP70', 'MX800', 'X3']
8544 |  |  | Pulse Rate from Plethysmogram (Right) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Pulse | ['MP70', 'MX800', 'X3']
8549 |  |  | Patient O2 Consumption per Body Surface | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | VO2m² | ['MP70', 'MX800', 'X3']
8550 |  |  | Patient O2 Consumption per Body Weight | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | VO2kg | ['MP70', 'MX800', 'X3']
8560 |  |  | Ideal Patient Weight | - | [] | Double | PatientInfo | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | IBW | ['MP70', 'MX800', 'X3']
8562 |  |  | End Flow Setting | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sEndFl | ['MP70', 'MX800', 'X3']
8564 |  |  | PSV Backup Delay Setting | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sPSVbd | ['MP70', 'MX800', 'X3']
8590 |  |  | Vt Not Achieved Alarm | unitless | Consult detailed help for more information on this variable | Long | Alarms | Common | [] | V60
8614 |  |  | Ventilator Failure Alarm | unitless | [] | Boolean | Alarms | Common | [] | Monnal T60
8640 |  |  | Technical Alarm Internal Source ID | unitless | Consult detailed help for more information on this variable | Double (Multiset) | Alarms | Common | Alarms | ['MP20', 'MP70', 'MX800', 'X3']
8641 |  |  | Technical Alarm Source Type | unitless | [] | Double (Multiset) | Alarms | Common | Alarms | ['MP20', 'MP70', 'MX800', 'X3']
8642 |  |  | Device Technical Alarm | unitless | Consult detailed help for more information on this variable | Double (Multiset) | Alarms | Common | Alarms | ['MP20', 'MP70', 'MX800', 'X3']
8643 |  |  | Patient Alarm Internal Source ID | unitless | Consult detailed help for more information on this variable | Double (Multiset) | Alarms | Common | Alarms | ['MP20', 'MP70', 'MX800', 'X3']
8644 |  |  | Patient Alarm Source Type | unitless | [] | Double (Multiset) | Alarms | Common | Alarms | ['MP20', 'MP70', 'MX800', 'X3']
8645 |  |  | Patient Alarm | unitless | Consult detailed help for more information on this variable | Double (Multiset) | Alarms | Common | Alarms | ['MP20', 'MP70', 'MX800', 'X3']
8646 |  |  | Patient Interface Type | unitless | Consult detailed help for more information on this variable | Long | DeviceInfo | Common | [] | V60
8653 |  |  | Tidal Volume Exceeded Alarm | unitless | Consult detailed help for more information on this variable | Long | Alarms | Common | [] | V60
8656 |  |  | Accoustic Respiration Rate Wave | - | [] | Double | Waveforms | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Waves: acRESP | ['MP70', 'MX800', 'X3']
8678 |  |  | General Temperature Alarm | unitless | Consult detailed help for more information on this variable | Double | Alarms | Common | [] | Criticool
8703 |  |  | Blood Temperature Wave | - | [] | Double | Waveforms | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Waves: Tblood | ['MP70', 'MX800', 'X3']
8714 |  |  | Non-specific Pressure 1 Wave | - | [] | Double | Waveforms | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Waves: P1 | ['MP70', 'MX800', 'X3']
8715 |  |  | Generic Invasive Blood Pressure P1 | - | Available on IntelliVue devices. | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | IBP: P1 | ['MP70', 'MX800', 'X3']
8716 |  |  | Generic Invasive Blood Pressure P1 Systolic | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | IBP: P1 | ['MP70', 'MX800', 'X3']
8717 |  |  | Generic Invasive Blood Pressure P1 Diastolic | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | IBP: P1 | ['MP70', 'MX800', 'X3']
8718 |  |  | Generic Invasive Blood Pressure P1 Mean | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | IBP: P1 | ['MP70', 'MX800', 'X3']
8719 |  |  | Non-specific Pressure 2 Wave | - | [] | Double | Waveforms | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Waves: P2 | ['MP70', 'MX800', 'X3']
8720 |  |  | Generic Invasive Blood Pressure P2 | - | Available on IntelliVue devices. | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | IBP: P2 | ['MP70', 'MX800', 'X3']
8721 |  |  | Generic Invasive Blood Pressure P2 Systolic | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | IBP: P2 | ['MP70', 'MX800', 'X3']
8722 |  |  | Generic Invasive Blood Pressure P2 Diastolic | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | IBP: P2 | ['MP70', 'MX800', 'X3']
8723 |  |  | Generic Invasive Blood Pressure P2 Mean | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | IBP: P2 | ['MP70', 'MX800', 'X3']
8724 |  |  | Non-specific Pressure 3 Wave | - | [] | Double | Waveforms | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Waves: P3 | ['MP70', 'MX800', 'X3']
8725 |  |  | Generic Invasive Blood Pressure P3 | - | Available on IntelliVue devices. | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | IBP: P3 | ['MP70', 'MX800', 'X3']
8726 |  |  | Generic Invasive Blood Pressure P3 Systolic | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | IBP: P3 | ['MP70', 'MX800', 'X3']
8727 |  |  | Generic Invasive Blood Pressure P3 Diastolic | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | IBP: P3 | ['MP70', 'MX800', 'X3']
8728 |  |  | Generic Invasive Blood Pressure P3 Mean | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | IBP: P3 | ['MP70', 'MX800', 'X3']
8729 |  |  | Non-specific Pressure 4 Wave | - | [] | Double | Waveforms | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Waves: P4 | ['MP70', 'MX800', 'X3']
8730 |  |  | Generic Invasive Blood Pressure P4 | - | Available on IntelliVue devices. | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | IBP: P4 | ['MP70', 'MX800', 'X3']
8731 |  |  | Generic Invasive Blood Pressure P4 Systolic | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | IBP: P4 | ['MP70', 'MX800', 'X3']
8732 |  |  | Generic Invasive Blood Pressure P4 Diastolic | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | IBP: P4 | ['MP70', 'MX800', 'X3']
8733 |  |  | Generic Invasive Blood Pressure P4 Mean | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | IBP: P4 | ['MP70', 'MX800', 'X3']
8734 |  |  | Non-specific Pressure 5 Wave | - | [] | Double | Waveforms | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Waves: P_5 | ['MP70', 'MX800', 'X3']
8735 |  |  | Generic Invasive Blood Pressure P5 | - | Available on IntelliVue devices. | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | IBP: P5 | ['MP70', 'MX800', 'X3']
8736 |  |  | Generic Invasive Blood Pressure P5 Systolic | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | IBP: P5 | ['MP70', 'MX800', 'X3']
8737 |  |  | Generic Invasive Blood Pressure P5 Diastolic | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | IBP: P5 | ['MP70', 'MX800', 'X3']
8738 |  |  | Generic Invasive Blood Pressure P5 Mean | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | IBP: P5 | ['MP70', 'MX800', 'X3']
8739 |  |  | Non-specific Pressure 6 Wave | - | [] | Double | Waveforms | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Waves: P_6 | ['MP70', 'MX800', 'X3']
8740 |  |  | Generic Invasive Blood Pressure P6 | - | Available on IntelliVue devices. | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | IBP: P6 | ['MP70', 'MX800', 'X3']
8741 |  |  | Generic Invasive Blood Pressure P6 Systolic | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | IBP: P6 | ['MP70', 'MX800', 'X3']
8742 |  |  | Generic Invasive Blood Pressure P6 Diastolic | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | IBP: P6 | ['MP70', 'MX800', 'X3']
8743 |  |  | Generic Invasive Blood Pressure P6 Mean | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | IBP: P6 | ['MP70', 'MX800', 'X3']
8744 |  |  | Non-specific Pressure 7 Wave | - | [] | Double | Waveforms | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Waves: P_7 | ['MP70', 'MX800', 'X3']
8745 |  |  | Generic Invasive Blood Pressure P7 | - | Available on IntelliVue devices. | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | IBP: P7 | ['MP70', 'MX800', 'X3']
8746 |  |  | Generic Invasive Blood Pressure P7 Systolic | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | IBP: P7 | ['MP70', 'MX800', 'X3']
8747 |  |  | Generic Invasive Blood Pressure P7 Diastolic | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | IBP: P7 | ['MP70', 'MX800', 'X3']
8748 |  |  | Generic Invasive Blood Pressure P7 Mean | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | IBP: P7 | ['MP70', 'MX800', 'X3']
8749 |  |  | Non-specific Pressure 8 Wave | - | [] | Double | Waveforms | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Waves: P_8 | ['MP70', 'MX800', 'X3']
8750 |  |  | Generic Invasive Blood Pressure P8 | - | Available on IntelliVue devices. | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | IBP: P8 | ['MP70', 'MX800', 'X3']
8751 |  |  | Generic Invasive Blood Pressure P8 Systolic | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | IBP: P8 | ['MP70', 'MX800', 'X3']
8752 |  |  | Generic Invasive Blood Pressure P8 Diastolic | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | IBP: P8 | ['MP70', 'MX800', 'X3']
8753 |  |  | Generic Invasive Blood Pressure P8 Mean | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | IBP: P8 | ['MP70', 'MX800', 'X3']
8754 |  |  | Intra-Uterine Pressure (IUP) Wave | - | [] | Double | Waveforms | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Waves: IUP | ['MP70', 'MX800', 'X3']
8755 |  |  | Intra-Uterine Pressure | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | IBP: IUP | ['MP70', 'MX800', 'X3']
8756 |  |  | Intra-Uterine Pressure Systolic | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
8757 |  |  | Intra-Uterine Pressure Diastolic | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
8758 |  |  | Intra-Uterine Pressure Mean | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
8760 |  |  | Unspecified Blood Pressure | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | IBP: P | ['MP70', 'MX800', 'X3']
8761 |  |  | Unspecified Blood Pressure Systolic | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
8762 |  |  | Unspecified Blood Pressure Diastolic | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
8763 |  |  | Unspecified Blood Pressure Mean | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
8788 |  |  | Airway Flow Wave | - | [] | Double | Waveforms | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Waves: AWF | ['MP70', 'MX800', 'X3']
8789 |  |  | Auxiliary Wave/Parameter | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
8790 |  |  | Global Ejection Fraction (2) | - | This is the device's measurement with label="NLS_NOM_EMFC_GEF" and physio_id="NOM_METRIC_NOS". See also 7810. | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
8791 |  |  | Pulmonary Vascular Permeability Index (2) | - | This is the device's measurement with label="NLS_NOM_EMFC_PVPI" and physio_id="NOM_METRIC_NOS". See also variable 6493. | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
8812 |  |  | Airway Volume | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Available only in revF or earlier | ['MP70', 'MX800', 'X3']
8858 |  |  | Arterial Blood Pressure (ART) Wave | - | [] | Double | Waveforms | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Waves: ART | ['MP70', 'MX800', 'X3']
8859 |  |  | Pulmonary Arterial Pressure (PAP) Wave | - | [] | Double | Waveforms | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Waves: PAP | ['MP70', 'MX800', 'X3']
8860 |  |  | Central Venous Pressure (CVP) Wave | - | [] | Double | Waveforms | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Waves: CVP | ['MP70', 'MX800', 'X3']
8861 |  |  | Right Atrial Pressure (RAP) Wave | - | [] | Double | Waveforms | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Waves: RAP | ['MP70', 'MX800', 'X3']
8862 |  |  | Left Atrial Pressure (LAP) Wave | - | [] | Double | Waveforms | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Waves: LAP | ['MP70', 'MX800', 'X3']
8863 |  |  | Intra-cranial Pressure (ICP) Wave | - | [] | Double | Waveforms | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Waves: ICP | ['MP70', 'MX800', 'X3']
8864 |  |  | Umbilical Arterial Pressure (UAP) Wave | - | [] | Double | Waveforms | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Waves: UAP | ['MP70', 'MX800', 'X3']
8865 |  |  | Umbilical Venous Pressure (UVP) Wave | - | [] | Double | Waveforms | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Waves: UVP | ['MP70', 'MX800', 'X3']
8866 |  |  | Airway CO2 Level Waveform | - | [] | Double | Waveforms | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Waves: CO2 | ['MP70', 'MX800', 'X3']
8867 |  |  | Airway O2 Wave | - | [] | Double | Waveforms | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Waves: O2 | ['MP70', 'MX800', 'X3']
8868 |  |  | Airway N2 Wave | - | [] | Double | Waveforms | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Waves: N2 | ['MP70', 'MX800', 'X3']
8869 |  |  | Generic Agent Wave (2nd) | - | [] | Double | Waveforms | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Waves: AGTs | ['MP70', 'MX800', 'X3']
8871 |  |  | Brachial Arterial Pressure (BAP) Wave | - | [] | Double | Waveforms | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Waves: BAP | ['MP70', 'MX800', 'X3']
8872 |  |  | Intracranial Pressure 1 (IC1) Wave | - | [] | Double | Waveforms | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Waves: IC1 | ['MP70', 'MX800', 'X3']
8873 |  |  | Intracranial Pressure 2 (IC2) Wave | - | [] | Double | Waveforms | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Waves: IC2 | ['MP70', 'MX800', 'X3']
8874 |  |  | Femoral Arterial Pressure (FAP) Wave | - | [] | Double | Waveforms | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Waves: FAP | ['MP70', 'MX800', 'X3']
8875 |  |  | Vector ECG Wave | - | [] | Double | Waveforms | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Waves: vECG | ['MP70', 'MX800', 'X3']
8889 |  |  | Transcutaneous Heater Power | - | [] | Double | OtherMeasuredData | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | HeatPa | ['MP70', 'MX800', 'X3']
8939 |  |  | BIS Standard Deviation | - | This variable is associated with ChannelID identifier. | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sdBISl/sdBISr | ['MP70', 'MX800', 'X3']
8940 |  |  | EMG Standard Deviation (BIS) | - | This variable is associated with ChannelID identifier. | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sdEMGl/sdEMGr | ['MP70', 'MX800', 'X3']
8941 |  |  | Asymmetric Index | - | This variable is associated with ChannelID identifier. | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | ASYM/ASYM L/ASYM R | ['MP70', 'MX800', 'X3']
9037 |  |  | O2 Consumption Index Estimate | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | iVO2Ie | ['MP70', 'MX800', 'X3']
9110 |  |  | Aortic Blood Pressure (Ao) | - | Available on IntelliVue devices with firmware revision G or later. Arterial Blood Pressure in the Aorta (NOM_PRESS_BLD_AORT). | Double | VitalSigns | Common | IBP: Ao | ['MP20', 'MP70', 'MX800', 'X3']
9111 |  |  | Umbilical Arterial Pressure (UAP) | - | Available on IntelliVue devices with firmware revision G or later. | Double | VitalSigns | Common | IBP: UAP | ['MP20', 'MP70', 'MX800', 'X3']
9112 |  |  | Umbilical Venous Pressure (UVP) | - | Available on IntelliVue devices with firmware revision G or later. | Double | VitalSigns | Common | IBP: UVP | ['MP20', 'MP70', 'MX800', 'X3']
9163 |  |  | Tracheal Pressure Waveform | mbar | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Evita V300. For Evita Infinity V500, Babylog VN500: Ptrach. | Double | Waveforms | ['Evita Infinity V500 [MEDIBUS.X]\xa0', 'Evita Infinity V800\xa0'] | [] | ['Evita V800', 'Evita V500']
9165 |  |  | Edi Diaphragm Electrical Activity Waveform | - | [] | Double | Waveforms | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Waves: EDI | ['MP70', 'MX800', 'X3']
9166 |  |  | Peak Diaphragm Electrical Activity (Edi) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Edimax | ['MP70', 'MX800', 'X3']
9167 |  |  | Minimum Diaphragm Electrical Activity (Edi) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Edimin | ['MP70', 'MX800', 'X3']
9168 |  |  | % Edi Trigger In NAVA | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | %EdiTr | ['MP70', 'MX800', 'X3']
9169 |  |  | % Edi Cycle Off In NAVA | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | %EdiCy | ['MP70', 'MX800', 'X3']
9170 |  |  | NAVA Level Setting | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sNAVA | ['MP70', 'MX800', 'X3']
9171 |  |  | Edi Trigger Setting | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sEdiTr | ['MP70', 'MX800', 'X3']
9196 |  |  | Patient Body Mass Index (BMI) | - | [] | Double | PatientInfo | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | BMI | ['MP70', 'MX800', 'X3']
9221 |  |  | Error Response From Device | unitless | Error code sent by the pump when an incorrect transmission happens. | UByte | DeviceInfo | Common | [] | Applix Smart
9267 |  |  | Device Pre-alarm Status | unitless | The alarms code are output as decimal values. Consult detailed help for more information on this variable | UShort | Alarms | Common | [] | Applix Smart
9270 |  |  | Maximum Positive Auxiliary Pressure | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Pxpeak | ['MP70', 'MX800', 'X3']
9271 |  |  | Mean Positive Auxiliary Pressure | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Pxmean | ['MP70', 'MX800', 'X3']
9272 |  |  | Minimum Positive Auxiliary Pressure | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Pxmin | ['MP70', 'MX800', 'X3']
9286 |  |  | Patient CO2 Production Indexed By Body Surface | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | VCO2m² | ['MP70', 'MX800', 'X3']
9287 |  |  | Patient CO2 Production Indexed By Body Weight | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | VCO2kg | ['MP70', 'MX800', 'X3']
9288 |  |  | Tidal Volume Indexed By Body Weight | ['mL/kg', '-'] | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. Tidal volume per kg body weight, VT/Wt. The leakage is determined by the difference between the delivered flow on the inspiratory side and the measured flow on the expiratory side, which is displayed as the leakage minute volume MVleak and relative as % leak (MVleak to MV). For Evita Infinity V500, Babylog VN500: VT/kg BW. | Double | VitalSigns | ['IntelliVue MX800\xa0', 'Evita Infinity V500 [MEDIBUS.X]\xa0', 'IntelliVue X3\xa0', 'Evita Infinity V800\xa0', 'IntelliVue MP70\xa0'] | TV/kg | ['X3', 'MX800', 'Evita V800', 'MP70', 'Evita V500']
9314 |  |  | Functional Residual Capacity | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | FRC | ['MP70', 'MX800', 'X3']
9315 |  |  | Corrected QT Interval (QTc) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | ECG: QTc | ['MP70', 'MX800', 'X3']
9317 |  |  | Airway N2 Concentration | - | Measurement from VueLink | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | N2 | ['MP70', 'MX800', 'X3']
9318 |  |  | Airway Agent Concentration | - | Measurement from VueLink. This variable is associated with a SourceID identifier. | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | AGT | ['MP70', 'MX800', 'X3']
9321 |  |  | Respiration Rate (Spirometry) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | RRspir | ['MP70', 'MX800', 'X3']
9323 |  |  | QT Heart Rate (QT-HR) | - | QT heart rate | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | ECG: QT-HR | ['MP70', 'MX800', 'X3']
9324 |  |  | Delta QTc | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | ECG: DeltaQTc | ['MP70', 'MX800', 'X3']
9329 |  |  | Humidification Type Setting | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sHumid | ['MP70', 'MX800', 'X3']
9336 |  |  | Normalized Rapid Shallow Breathing Index | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | RSBInm | ['MP70', 'MX800', 'X3']
9337 |  |  | Percent of Spontaneous Inspiratory Time | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Tin/Tt | ['MP70', 'MX800', 'X3']
9408 |  |  | Energy Balance Calculated by Metabolics | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | dE | ['MP70', 'MX800', 'X3']
9437 |  |  | Alveolar Volume | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Valv | ['MP70', 'MX800', 'X3']
9438 |  |  | Apparatus Dead Space | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | appVd | ['MP70', 'MX800', 'X3']
9439 |  |  | Pleth (Pre-Ductal) | - | Pulse Rate from Plethysmogram (pre ductal) | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Pulse | ['MP70', 'MX800', 'X3']
9440 |  |  | Pleth (Post-Ductal) | - | Pulse Rate from Plethysmogram (post ductal) | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Pulse | ['MP70', 'MX800', 'X3']
9441 |  |  | Relative Perfusion (Pre-Ductal) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | PerfPr | ['MP70', 'MX800', 'X3']
9442 |  |  | Relative Perfusion (Post-Ductal) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | PerfPo | ['MP70', 'MX800', 'X3']
9444 |  |  | Cardioplegia Main Pump Time of Current Plegia Session | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | CpOnTi | ['MP70', 'MX800', 'X3']
9445 |  |  | Cardioplegia Main Pump Time Since Stop of Plegia Session | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | CpOffT | ['MP70', 'MX800', 'X3']
9446 |  |  | Cardioplegia Main Pump Total Plegia Time | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | CpPlTi | ['MP70', 'MX800', 'X3']
9447 |  |  | Cardioplegia Slave Pump Time Since Start of Current Plegia Session | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | CsOnTi | ['MP70', 'MX800', 'X3']
9448 |  |  | Cardioplegia Slave Pump Time Since Stop of Plegia Session | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | CsOffT | ['MP70', 'MX800', 'X3']
9449 |  |  | Cardioplegia Slave Pump Total Plegia Time | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | CsPlTi | ['MP70', 'MX800', 'X3']
9453 |  |  | PAV-based Lung Compliance | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Cpav | ['MP70', 'MX800', 'X3']
9454 |  |  | PAV-based Patient Resistance | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Rpav | ['MP70', 'MX800', 'X3']
9455 |  |  | PAV-based Total Resistance Estimate | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Rtot | ['MP70', 'MX800', 'X3']
9456 |  |  | PAV-based Lung Elastance | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | Epav | ['MP70', 'MX800', 'X3']
9467 |  |  | Backup Inspiratory Pressure Setting | - | [] | Double | Settings | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | sPIbak | ['MP70', 'MX800', 'X3']
9479 |  |  | SmartCare Target Delta Support Pressure | mbar | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. Dpressure support goal in SmartCare, SC-DPsupp goal. For Evita Infinity V500: SC-Target ΔPsupp ΔPsupp goal. | Double | VitalSigns | ['Evita Infinity V500 [MEDIBUS.X]\xa0', 'Evita Infinity V800\xa0'] | [] | ['Evita V800', 'Evita V500']
9480 |  |  | SmartCare Rated Delta Support Pressure | mbar | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. Rated Dpressure support by internal controller in SmartCare, SC-DPsupp rated. For Evita Infinity V500: SC-Rated ΔPsupp SC-Δpsupp. | Double | VitalSigns | ['Evita Infinity V500 [MEDIBUS.X]\xa0', 'Evita Infinity V800\xa0'] | [] | ['Evita V800', 'Evita V500']
9481 |  |  | SmartCare Duration of Patient Session | - | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. Duration of session (hours and minutes) in SmartCare, SC-duration. Variable is in Multiset format (need combine the first value in hours with the second value in minutes to obtain the exact duration). Please use variable 4346 instead. For Evita Infinity V500: Multiset variable (need combine the first value in hours with the second value in minutes to obtain the exact duration). | Double (Multiset) | VitalSigns | ['Evita Infinity V500 [MEDIBUS.X]\xa0', 'Evita Infinity V800\xa0'] | [] | ['Evita V800', 'Evita V500']
9482 |  |  | SmartCare Spontaneous Respiration Rate | 1/min | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. Spontaneous respiratory rate in SmartCare, SC-Rrspon. For Evita Infinity V500: SC-Rrspon. | Double | VitalSigns | ['Evita Infinity V500 [MEDIBUS.X]\xa0', 'Evita Infinity V800\xa0'] | [] | ['Evita V800', 'Evita V500']
9483 |  |  | SmartCare Tidal Volume | mL | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. Tidal volume in SmartCare, SC-VT. For Evita Infinity V500: SC-VT. | Double | VitalSigns | ['Evita Infinity V500 [MEDIBUS.X]\xa0', 'Evita Infinity V800\xa0'] | [] | ['Evita V800', 'Evita V500']
9484 |  |  | SmartCare EtCO2 | mmHg | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. End-tidal CO2 concentration in SmartCare, SC-etCO2. For Evita Infinity V500: SC-etCO2. | Double | VitalSigns | ['Evita Infinity V500 [MEDIBUS.X]\xa0', 'Evita Infinity V800\xa0'] | [] | ['Evita V800', 'Evita V500']
9486 |  |  | Inspiratory Flow Termination Level Setting | % | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Savina 300, Oxylog 3000 plus, Evita V300. Inspiratory termination criterion based on peak inspiratory flow, Insp. term. For Evita Infinity V500: Inspiratory termination. | Double | Settings | ['Evita Infinity V500 [MEDIBUS.X]\xa0', 'Evita Infinity V800\xa0'] | [] | ['Evita V800', 'Evita V500']
9487 |  |  | Expiratory Flow Termination Setting | % | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. Expiratory termination criterion during APRV/AutoRelease, Exp. term. For Evita Infinity V500, Babylog VN500: Expiratory termination. | Double | Settings | ['Evita Infinity V500 [MEDIBUS.X]\xa0', 'Evita Infinity V800\xa0'] | [] | ['Evita V800', 'Evita V500']
9488 |  |  | SmartCare Body Weight | kg | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. Ideal body weight in SmartCare, SC-Wt. For Evita Infinity V500: SC-Body weight. | Double | Settings | ['Evita Infinity V500 [MEDIBUS.X]\xa0', 'Evita Infinity V800\xa0'] | [] | ['Evita V800', 'Evita V500']
9489 |  |  | SmartCare Start of Night Rest | h | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. Start of night rest in SmartCare, SC-startrest. For Evita Infinity V500: SC-Start of night rest. | Double | Settings | ['Evita Infinity V500 [MEDIBUS.X]\xa0', 'Evita Infinity V800\xa0'] | [] | ['Evita V800', 'Evita V500']
9490 |  |  | SmartCare End of Night Rest | h | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. End of night rest in SmartCare, SC-endrest. For Evita Infinity V500: SC-End of night rest. | Double | Settings | ['Evita Infinity V500 [MEDIBUS.X]\xa0', 'Evita Infinity V800\xa0'] | [] | ['Evita V800', 'Evita V500']
9526 |  |  | Bed Head Angle | ° | The articulated angle of the bed’s head section in degrees. -360.0 -- 360.0 | Double | DeviceInfo | Bed Data Server (Common) | BED | Total Scale
9529 |  |  | Current Patient Weight | kg | The last patient weight captured by the dynamic weight scale system displayed in kilograms. -9999.9 -- 9999.9 | Double | VitalSigns | Bed Data Server (Common) | BED | Total Scale
9530 |  |  | Bed Brake Status | unitless | The current status for the brake switch. | String | DeviceInfo | Bed Data Server (Common) | BED | Total Scale
9536 |  |  | SmartBed Working Mode | unitless | The current configuration mode of the patient positioning monitor (PPM) alarm system. | String | Settings | Bed Data Server (Common) | BED | Total Scale
9537 |  |  | SmartBed General Alarm | unitless | The current status of the patient positioning monitor (PPM) alarm system. | Byte | Alarms | Bed Data Server (Common) | BED | Total Scale
9543 |  |  | SmartBed Patient Position State | unitless | The current level of patient activity as detected by the patient positioning monitor (PPM) alarm system. | String | OtherMeasuredData | Bed Data Server (Common) | BED | Total Scale
9585 |  |  | Plateau Time | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | PlatTi | ['MP70', 'MX800', 'X3']
9586 |  |  | QT Baseline | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | ECG: QT Bsl | ['MP70', 'MX800', 'X3']
9587 |  |  | QT Heart Rate Baseline | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | ECG: QTHRBl | ['MP70', 'MX800', 'X3']
9588 |  |  | High Density Lipoprotein | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | HDL | ['MP70', 'MX800', 'X3']
9589 |  |  | Low Density Lipoprotein | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | LDL | ['MP70', 'MX800', 'X3']
9590 |  |  | C-reactive Protein | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | CRP | ['MP70', 'MX800', 'X3']
9591 |  |  | Urine Hemoglobin | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | UrHb | ['MP70', 'MX800', 'X3']
9595 |  |  | Noninvasive Total Hemoglobin (SpHb) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | SpHb | ['MP70', 'MX800', 'X3']
9596 |  |  | Noninvasive Oxygen Content (SpOC) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | SpOC | ['MP70', 'MX800', 'X3']
9597 |  |  | Noninvasive Carboxyhemoglobin (SpCO) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | SpCO | ['MP70', 'MX800', 'X3']
9598 |  |  | Noninvasive Methemoglobin (SpMet) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | SpMet | ['MP70', 'MX800', 'X3']
9600 |  |  | Pleth Variability Index (PVI) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | PVI | ['MP70', 'MX800', 'X3']
9603 |  |  | Regional O2 Saturation (rSO2) | - | [] | Double | VitalSigns | Common | [] | ['MP20', 'MP70', 'MX800', 'X3']
9605 |  |  | rSO2 Area Under Curve (AUC) | - | This variable is associated with ChannelID identifier. | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | AUC_1/AUC_2/AUC_3/AUC_4/AUC_5/AUC_6 | ['MP70', 'MX800', 'X3']
9610 |  |  | Generic O2 Saturation (3) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | SO2 3 | ['MP70', 'MX800', 'X3']
9611 |  |  | Generic O2 Saturation (4) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | SO2 4 | ['MP70', 'MX800', 'X3']
9612 |  |  | Tissue O2 Saturation (StO2) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | StO2 | ['MP70', 'MX800', 'X3']
9613 |  |  | Nitrogen Balance (NsLoss) | - | [] | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | NsLoss | ['MP70', 'MX800', 'X3']
9642 |  |  | Replacement Flow Rate | mL/h | [] | Long | VitalSigns | Common | [] | Artis
9643 |  |  | Blood Flow Rate | mL/min | Actual blood flow This variable contains value "0" when blood pump is stopped. | Long | VitalSigns | ['Common\xa0', 'Multifiltrate PRO\xa0'] | Dialysis Course | ['Multifiltrate PRO', '5008']
9648 |  |  | Transmembrane Pressure Lower Alarm Limit | mmHg | This variable is also sent as alarm limit of variableID 3029. | Long | Settings | Multifiltrate PRO | Dialysis Settings | Multifiltrate PRO
9836 |  |  | CO2 Sensor Failure Alarm | unitless | [] | Boolean | Alarms | Monnal T60 | [] | Monnal T60
9882 |  |  | Check Maximum Pressure Alarm | unitless | Consult detailed help for more information on this variable | Long | Alarms | Common | [] | V60
9955 |  |  | Maximum IPAP Setting | - | Possible units: cmH2O, hPa, kPa, mBar, mmHg | Double | Settings | Common | [] | V60
31065 |  |  | Low Ventilator Total Respiration Rate Alarm | unitless | Consult detailed help for more information on this variable | Long | Alarms | Common | [] | V60
32211 |  |  | Low Expired Tidal Volume Alarm | unitless | Consult detailed help for more information on this variable | Long | Alarms | Common | [] | V60
32221 |  |  | Low Expired Minute Volume Alarm | unitless | Consult detailed help for more information on this variable | Long | Alarms | Common | [] | V60
32311 |  |  | Low Peak Inspiratory Pressure Alarm | unitless | Consult detailed help for more information on this variable | Long | Alarms | Common | [] | V60
50021 |  |  | Device Alarm Inactivation State | unitless | [] | String | DeviceInfo | ['Evita Infinity V500 [MEDIBUS.X]\xa0', 'Evita Infinity V800\xa0'] | [] | ['Evita V800', 'Evita V500']
50201 |  |  | Device Type | unitless | This variable contains a fixed value "48 (0x30)" for Infusion Pump. | UByte | DeviceInfo | Common | Recommended Variables | Applix Smart
50203 |  |  | Device Model | unitless | [] | String | DeviceInfo | Common | [] | V60
50212 |  |  | Device Software Revision | unitless | [] | String | DeviceInfo | Common | [] | V60
51065 |  |  | Ventilator Total Respiration Rate | 1/min | [] | Double | VitalSigns | Common | [] | V60
52127 |  |  | Inspiration Time Percentage | % | [] | Double | VitalSigns | Common | [] | V60
52211 |  |  | Expired Tidal Volume | mL | [] | Double | VitalSigns | Common | [] | V60
52221 |  |  | Expired Minute Volume | L/min | [] | Double | VitalSigns | Common | [] | V60
52311 |  |  | Peak Inspiratory Pressure | cmH2O | [] | Double | VitalSigns | Common | [] | V60
52523 |  |  | Inspired O2 Concentration | % | See variable 52527 for partial pressure | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MP70\xa0', 'IntelliVue MX800\xa0'] | O2 | ['MP70', 'MX800', 'X3']
52527 |  |  | Inspired O2 Partial Pressure | mmHg | This variable is calculated from variable 52523 and will output inspired O2 partial pressure. | Double | VitalSigns | ['IntelliVue X3\xa0', 'IntelliVue MX800\xa0'] | O2 | ['MP70', 'MX800', 'X3']
52911 |  |  | Airway Flow Waveform | L/min | [] | Double | Waveforms | Common | [] | V60
52912 |  |  | Airway Volume Waveform | mL | [] | Double | Waveforms | Common | [] | V60
52913 |  |  | Airway Pressure Waveform | cmH2O | [] | Double | Waveforms | Common | [] | V60
53001 |  |  | Infusion Pump Type | unitless | Consult detailed help for more information on this variable | UByte | DeviceInfo | Common | Recommended Variables | Applix Smart
53012 |  |  | Infusion Pump Serial Number | unitless | 8 characters string | String | DeviceInfo | Common | Recommended Variables | Applix Smart
53021 |  |  | Infusion Pump Rack Index | unitless | The variable has a fixed value "1" because it is a standalone pump. | UByte | DeviceInfo | Common | Recommended Variables | Applix Smart
53022 |  |  | Infusion Pump Slot Index | unitless | The variable has a fixed value "1" because it is a standalone pump. | UByte | DeviceInfo | Common | Recommended Variables | Applix Smart
53101 |  |  | Generic Infusion Pump Mode | unitless | Consult detailed help for more information on this variable | UShort | OtherMeasuredData | Common | Recommended Variables | Applix Smart
53111 |  |  | Actual Infusion Volumetric Rate | mL/h | [] | Double | VitalSigns | Common | Recommended Variables | Applix Smart
53121 |  |  | Total Volume Infused | mL | This variable is reset to "0" only if users do so manually on the pump panel. It is not automatically reset if the infusion has been interrupted and/or the pump has been switched off. | Double | VitalSigns | Common | Recommended Variables | Applix Smart
53211 |  |  | Infusion Drug Name | unitless | This variable contains a fixed nonsignificant value "NA" | String | PatientInfo | Common | Recommended Variables | Applix Smart
53221 |  |  | Infusion Volumetric Rate Setting | mL/h | [] | Double | Settings | Common | Recommended Variables | Applix Smart
53801 |  |  | Infusion Event | unitless | Consult detailed help for more information on this variable | UByte | Messages | Common | Recommended Variables | Applix Smart
54011 |  |  | Ventilation Standby State | unitless | [] | Boolean | VitalSigns | Common | [] | V60
54013 |  |  | Ventilation Mode | unitless | Consult detailed help for more information on this variable | String | Settings | Common | [] | V60
54021 |  |  | Ventilation Mode Primary | unitless | Consult detailed help for more information on this variable | String | Settings | Common | [] | V60
54022 |  |  | Ventilation Mode Options | unitless | Consult detailed help for more information on this variable | String | Settings | Common | [] | V60
54051 |  |  | Inspired O2 (FiO2) Setting | % | [] | Double | Settings | Common | [] | V60
54061 |  |  | Ventilation Leakage Minute Volume | L/min | [] | Double | VitalSigns | Common | [] | V60
54111 |  |  | Mandatory Respiratory Rate Setting | 1/min | [] | Double | Settings | Common | [] | V60
54121 |  |  | Inspiratory Time Setting | s | [] | Double | Settings | Common | [] | V60
54123 |  |  | lnspiratory Rise Time Setting | s | [] | Double | Settings | Common | [] | V60
54151 |  |  | Inspired Tidal Volume Setting | mL | [] | Double | Settings | Common | [] | V60
54172 |  |  | CPAP Setting | cmH2O | [] | Double | Settings | Common | [] | V60
54231 |  |  | Inspiratory Flow Termination Setting | L/min | Auto-Trak sensitivity: set flow for cycling to exhalation | Double | Settings | Common | [] | V60
54351 |  |  | Proportional Support Level Setting | % | [] | Double | Settings | Common | [] | V60
54381 |  |  | EPAP Setting | cmH2O | [] | Double | Settings | Common | [] | V60
54382 |  |  | IPAP Setting | cmH2O | [] | Double | Settings | Common | [] | V60
54624 |  |  | Fresh Gas Total Flow Rate Setting | L/min | [] | Double | Settings | V60 (PVOI) | [] | V60
71065 |  |  | High Ventilator Total Respiration Rate Alarm | unitless | Consult detailed help for more information on this variable | Long | Alarms | Common | [] | V60
72211 |  |  | High Expired Tidal Volume Alarm | unitless | Consult detailed help for more information on this variable | Long | Alarms | Common | [] | V60
72311 |  |  | High Peak Inspiratory Pressure Alarm | unitless | Consult detailed help for more information on this variable | Long | Alarms | Common | [] | V60
90714 |  |  | Battery Low Alarm | unitless | This variable is associated with a ChannelID identifier. | Long | Alarms | Common | [] | V60
90719 |  |  | Main Power Disconnected Alarm | unitless | Consult detailed help for more information on this variable | Long | Alarms | Common | [] | V60
92021 |  |  | Ventilator Obstruction Alarm | unitless | Consult detailed help for more information on this variable | Long | Alarms | Common | [] | V60
92405 |  |  | Pressure Sensor Disconnected Alarm | unitless | Consult detailed help for more information on this variable | Long | Alarms | Common | [] | V60
92612 |  |  | O2 Supply Alarm | unitless | Consult detailed help for more information on this variable | Long | Alarms | Common | [] | V60
92631 |  |  | O2 Supply High Alarm | unitless | Consult detailed help for more information on this variable | Long | Alarms | Common | [] | V60
92632 |  |  | O2 Supply Low Alarm | unitless | Consult detailed help for more information on this variable | Long | Alarms | Common | [] | V60