## Sheet1
ID | Label | Type | Group | Unit | Description | Device | Source
28 | Active Alarms | Double (Multiset) | Alarms | unitless | See "Multiset Alarms" on help file. | Common | Alarms
29 | Latched Alarms | Double (Multiset) | Alarms | unitless | See "Multiset Alarms" on help file. | Common | Alarms
33 | Inactive Alarms | Double (Multiset) | Alarms | unitless | See "Multiset Alarms" on help file. | Common | Alarms
177 | Alarm Data | String | Alarms | unitless | This variable is same as 3878 but sent as a string, which is an array containing all alarm messages and separated by delimiter "¦¦". This variable, along with variables 588 and 4131, are marked as obsolete because they use either"¦¦" or "," in the array string as delimiter, which are often not compatible with customers' applications. Variables 6164, 3878, and 2596 should be used instead. The HL7 output can contain an '&' character to indicate a LOW alarm. In order to avoid issues, the '&' character will be replaced by a '?' character in the DDI output. | Evita Infinity V500 [MEDIBUS.X] | Alarm And Text Message
356 | Generic Dialysate Alarm | Boolean | Alarms | unitless | Value of "1" when alarm is active. | Common
357 | Urea Clearance Alarm | Boolean | Alarms | unitless | Value of "1" when alarm is active. | Common
358 | Blood Volume Alarm | Boolean | Alarms | unitless | Value of "1" when alarm is active. | Common
359 | Mixed Dilution Alarm | Boolean | Alarms | unitless | Value of "1" when alarm is active. | Common
540 | Low Venous Line Pressure Alarm | Boolean | Alarms | unitless | Value of "1" when alarm is active. | Common
543 | High Venous Line Pressure Alarm | Boolean | Alarms | unitless | Value of "1" when alarm is active. | Common
546 | High Arterial Line Pressure Alarm | Boolean | Alarms | unitless | Value of "1" when alarm is active. | Common
547 | Low Arterial Line Pressure Alarm | Boolean | Alarms | unitless | Value of "1" when alarm is active. | Common
558 | Heparin Pump Alarm | Boolean | Alarms | unitless | Value of "1" when alarm is active. | Common
588 | Alarm Message | String | Alarms | unitless | This variable is same as 6164 but sent as a string, which is an array containing all alarm codes and separated by delimiter ",". This variable, along with variables 177 and 4131, are marked as obsolete because they use either"¦¦" or "," in the array string as delimiter, which are often not compatible with customers' applications. Variables 6164, 3878, and 2596 should be used instead. | Evita Infinity V500 [MEDIBUS.X] | Alarm And Text Message
768 | Generic Error Code | Short | Alarms | unitless | TD measurement error, "0" is transmitted if no error occurs. | Common | Error
955 | Battery Running Alarm | Boolean | Alarms | unitless | Value of "1" when alarm is active. | Common
1187 | Low Inspired Tidal Volume Alarm | Boolean | Alarms | unitless |  | Common
1188 | Low Expired Tidal Volume Alarm | Boolean | Alarms | unitless |  | Common
1228 | High Inspired Tidal Volume Alarm | Boolean | Alarms | unitless |  | Common
1370 | Battery Low Alarm | Boolean | Alarms | unitless |  | Common
1538 | O2 Sensor Alarm | Boolean | Alarms | unitless |  | Common
1803 | Device Alarm Global Status | String | Alarms | unitless | Format: Bit 2: 0 --> Alarm | Common
1821 | Generic ICG Error | Short | Alarms | unitless | ICG measurement error, "0" is transmitted if no error occurs. | PiCCO2 | Error
1822 | Blood Pump Alarm | Boolean | Alarms | unitless | Value of "1" when alarm is active. | Common
1831 | Bed Head Angle Alarm | Byte | Alarms | unitless | The current status for the head angle alarm system. | Bed Data Server (Common) | BED
2032 | Volume Not Delivered Alarm | Long | Alarms | unitless | Consult detailed help for more information on this variable | Common
2120 | Battery Alarm | Boolean | Alarms | unitless | This alarm is raised if the internal battery is discharged. | Common
2132 | AC Power Supply Failure Alarm | ULong | Alarms | unitless | This variable is 32 bits long. | Common
2445 | Patient Rebreathing Alarm | Long | Alarms | unitless | Consult detailed help for more information on this variable | Common
2462 | SvO2 Alarm | Short | Alarms | unitless | High/Low Alarm of ScvO2 Consult detailed help for more information on this variable | PiCCO2 | Alarm
2546 | Low Expired Minute Volume Alarm | Boolean | Alarms | unitless |  | Common
2568 | Airway Peak Pressure Alarm | Boolean | Alarms | unitless |  | Common
2590 | Generic Device Alarm | String | Alarms | unitless | Format: list of maximum 16 alarm IDs separated with a "¦". | Common
2595 | General Alarm Status | Long (Multiset) | Alarms | unitless | Contains every alarm code that is currently active | Common
2596 | Alarm Severity Level | UByte (Multiset) | Alarms | unitless | This variable is in multiset format. This variable contains alarm priority level, variable 6164 contains alarm codes, and variable 3878 contains alarm messages. They need to be interpreted together. One alarm code may have one or more priority values and one or more alarm messages associated with it. The possible alarms are depended on device firmware version. Please see device user manuals for details. See variables 177 and 588 and 4131 for string format. | Evita Infinity V500 [MEDIBUS.X] | Alarm And Text Message
2976 | Patient Disconnect Alarm | Boolean | Alarms | unitless |  | Common
3253 | Blood Circuit Pressure Alarm | Boolean | Alarms | unitless | Value of "1" when alarm is active. | Common
3592 | Low Airway Flow Alarm | Long | Alarms | unitless | Consult detailed help for more information on this variable | V60 (PVOI)
3636 | Device Alarm Status | UShort | Alarms | unitless | The alarms code are output as decimal values. Consult detailed help for more information on this variable | Common
3691 | HLM System Control Module Technical Alarm | ULong | Alarms | unitless | This variable is 24 bits long. | Common
3705 | Generic Pump Flow Alarm | UShort | Alarms | unitless | This variable is 16 bits long. | Common
3849 | Alarm History | UShort (Multiset) | Alarms | unitless | 16 last alarms | Common
3858 | Volume Not Achieved Alarm | Boolean | Alarms | unitless |  | Common
3878 | Generic Alarm Message | String (Multiset) | Alarms | unitless | This variable is in multiset format. This variable contains alarm messages, variable 2596 contains alarm priority level, variable 6164 contains alarm codes. They need to be interpreted together. One alarm code may have one or more priority values and one or more alarm messages associated with it. The possible alarms are dependent on device firmware version. Please see device user manuals for details. See variables 177 and 588 and 4131 for string format. The HL7 output can contain an '&' character to indicate a LOW alarm. In order to avoid issues, the '&' character will be replaced by a '?' character in the DDI output. | Evita Infinity V500 [MEDIBUS.X] | Alarm And Text Message
4131 | Alarm Priority | String | Alarms | unitless | This variable is same as 2596 but sent as a string, which is an array containing all alarm priorities and separated by "," delimiter. This variable, along with variables 177 and 588, are marked as obsolete because they use either"¦¦" or "," in the array string as delimiter, which are often not compatible with customers' applications. Variables 6164, 3878, and 2596 should be used instead. | Evita Infinity V500 [MEDIBUS.X] | Alarm And Text Message
4391 | Apnea Ventilation Alarm | Boolean | Alarms | unitless |  | Common
4457 | System Alarm Status | Double | Alarms | unitless | Overall alert state of the system. | Common | Alarms
4458 | Patient Alarm Priority | Double | Alarms | unitless | Overall Patient alarm priority. This is the current highest priority value of currently active Patient Alarms. | Common | Alarms
4459 | Technical Alarm Priority | Double | Alarms | unitless | Overall Technical alarm priority. This is the current highest priority value of currently active Technical Alarms. | Common | Alarms
4460 | Audio Alarm Priority | Double | Alarms | unitless | Overall audible alarm priority. This is the current highest priority value of currently active audible alarms. | Common | Alarms
4461 | Technical Alarm Source ID | Double (Multiset) | Alarms | unitless |  | Common | Alarms
4462 | Technical Alarm Global State | Double (Multiset) | Alarms | unitless | Consult detailed help for more information on this variable | Common | Alarms
4463 | Technical Alarm Priority | Double (Multiset) | Alarms | unitless | Consult detailed help for more information on this variable | Common | Alarms
4464 | Technical Alarm Flag | Double (Multiset) | Alarms | unitless | Consult detailed help for more information on this variable | Common | Alarms
4465 | Technical Alarm Priority | Double (Multiset) | Alarms | unitless |  | Common | Alarms
4466 | Technical Alarm Description | String (Multiset) | Alarms | unitless |  | Common | Alarms
4467 | Patient Alarm Source ID | Double (Multiset) | Alarms | unitless |  | Common | Alarms
4468 | Patient Alarm Global State | Double (Multiset) | Alarms | unitless | Consult detailed help for more information on this variable | Common | Alarms
4469 | Patient Alarm Priority | Double (Multiset) | Alarms | unitless | Consult detailed help for more information on this variable | Common | Alarms
4470 | Patient Alarm Flag | Double (Multiset) | Alarms | unitless | Consult detailed help for more information on this variable | Common | Alarms
4471 | Patient Alarm Priority | Double (Multiset) | Alarms | unitless |  | Common | Alarms
4472 | Patient Alarm Description | String (Multiset) | Alarms | unitless |  | Common | Alarms
4596 | System Pressure High Alarm | Boolean | Alarms | unitless | Possible values: 0 = Active 1 = Not Active | Common
5235 | Low Water Level Alarm | Boolean | Alarms | unitless | Low water level alarm Possible values: 0 = On 1 = Off | Common
5318 | Blood Temperature Alarm | Boolean | Alarms | unitless | Value of "1" when alarm is active. | Common
5684 | Continuous Cardiac Output (CCO) Alarm | Short | Alarms | unitless | High/Low Alarm of PCCO Consult detailed help for more information on this variable | PiCCO2 | Alarm
5712 | Low EtCO2 Alarm | Boolean | Alarms | unitless |  | Monnal T60
5713 | High EtCO2 Alarm | Boolean | Alarms | unitless |  | Monnal T60
5961 | PEEP/CPAP Alarm | Boolean | Alarms | unitless |  | Common
6153 | ET Tube Blocked Alarm | Boolean | Alarms | unitless |  | Common
6164 | General Alarm Code | UShort (Multiset) | Alarms | unitless | This variable is in multiset format. This variable contains alarm codes, variable 3878, alarm messages, variable 2596 contains alarm priority level. They need to be interpreted together. One alarm code may have one or more priority values and one or more alarm messages associated with it. The possible alarms are depended on device firmware version. Please see device user manuals for details. See variables 177 and 588 and 4131 for string format. | Evita Infinity V500 [MEDIBUS.X] | Alarm And Text Message
6296 | O2 Supply Alarm | Boolean | Alarms | unitless |  | Monnal T60
6376 | High Leak Alarm | Boolean | Alarms | unitless |  | Monnal T60
6891 | Expiratory Flow Sensor Failure Alarm | Boolean | Alarms | unitless |  | Common
7033 | Patient Demand High Alarm | Boolean | Alarms | unitless |  | Common
7038 | High Inspired Minute Volume Alarm | Boolean | Alarms | unitless |  | Common
7039 | Low Inspired Minute Volume Alarm | Boolean | Alarms | unitless |  | Common
7941 | SpO2 Alarm | Short | Alarms | unitless | High/Low Alarm of SpO2 Consult detailed help for more information on this variable | PiCCO2 | Alarm
7980 | Dialysate Conductivity Alarm | Boolean | Alarms | unitless | Value of "1" when alarm is active. | Common
8012 | Plateau Airway Pressure Alarm | Boolean | Alarms | unitless |  | Monnal T60
8301 | Nebulizer Error | Boolean | Alarms | unitless |  | Monnal T60
8579 | Plimit Reached Alarm | Boolean | Alarms | unitless |  | Monnal T60
8589 | High Expired Tidal Volume Alarm | Boolean | Alarms | unitless |  | Common
8590 | Vt Not Achieved Alarm | Long | Alarms | unitless | Consult detailed help for more information on this variable | Common
8596 | Low FiO2 Alarm | Boolean | Alarms | unitless |  | Common
8597 | High FiO2 Alarm | Boolean | Alarms | unitless |  | Common
8598 | Low Respiration Rate Alarm | Boolean | Alarms | unitless |  | Common
8599 | High Respiration Rate Alarm | Boolean | Alarms | unitless |  | Common
8600 | CO2 Apnea Alarm | Boolean | Alarms | unitless |  | Monnal T60
8606 | Battery Failure Alarm | Boolean | Alarms | unitless |  | Common
8614 | Ventilator Failure Alarm | Boolean | Alarms | unitless |  | Common
8640 | Technical Alarm Internal Source ID | Double (Multiset) | Alarms | unitless | Consult detailed help for more information on this variable | Common | Alarms
8641 | Technical Alarm Source Type | Double (Multiset) | Alarms | unitless |  | Common | Alarms
8642 | Device Technical Alarm | Double (Multiset) | Alarms | unitless | Consult detailed help for more information on this variable | Common | Alarms
8643 | Patient Alarm Internal Source ID | Double (Multiset) | Alarms | unitless | Consult detailed help for more information on this variable | Common | Alarms
8644 | Patient Alarm Source Type | Double (Multiset) | Alarms | unitless |  | Common | Alarms
8645 | Patient Alarm | Double (Multiset) | Alarms | unitless | Consult detailed help for more information on this variable | Common | Alarms
8653 | Tidal Volume Exceeded Alarm | Long | Alarms | unitless | Consult detailed help for more information on this variable | Common
8678 | General Temperature Alarm | Double | Alarms | unitless | Consult detailed help for more information on this variable | Common
9266 | Generic ECG Alarm | UByte | Alarms | unitless | This variable is 8 bits long. | Common
9267 | Device Pre-alarm Status | UShort | Alarms | unitless | The alarms code are output as decimal values. Consult detailed help for more information on this variable | Common
9537 | SmartBed General Alarm | Byte | Alarms | unitless | The current status of the patient positioning monitor (PPM) alarm system. | Bed Data Server (Common) | BED
9584 | High Expired Minute Volume Alarm | Boolean | Alarms | unitless |  | Common
9836 | CO2 Sensor Failure Alarm | Boolean | Alarms | unitless |  | Monnal T60
9882 | Check Maximum Pressure Alarm | Long | Alarms | unitless | Consult detailed help for more information on this variable | Common
31065 | Low Ventilator Total Respiration Rate Alarm | Long | Alarms | unitless | Consult detailed help for more information on this variable | Common
32211 | Low Expired Tidal Volume Alarm | Long | Alarms | unitless | Consult detailed help for more information on this variable | Common
32221 | Low Expired Minute Volume Alarm | Long | Alarms | unitless | Consult detailed help for more information on this variable | Common
32311 | Low Peak Inspiratory Pressure Alarm | Long | Alarms | unitless | Consult detailed help for more information on this variable | Common
71065 | High Ventilator Total Respiration Rate Alarm | Long | Alarms | unitless | Consult detailed help for more information on this variable | Common
72211 | High Expired Tidal Volume Alarm | Long | Alarms | unitless | Consult detailed help for more information on this variable | Common
72311 | High Peak Inspiratory Pressure Alarm | Long | Alarms | unitless | Consult detailed help for more information on this variable | Common
90714 | Battery Low Alarm | Long | Alarms | unitless | This variable is associated with a ChannelID identifier. | Common
90719 | Main Power Disconnected Alarm | Long | Alarms | unitless | Consult detailed help for more information on this variable | Common
92001 | Ventilator Patient Disconnected Alarm | Long | Alarms | unitless | Consult detailed help for more information on this variable | Common
92011 | Ventilator Inoperative Alarm | Long | Alarms | unitless | Consult detailed help for more information on this variable | Common
92021 | Ventilator Obstruction Alarm | Long | Alarms | unitless | Consult detailed help for more information on this variable | Common
92405 | Pressure Sensor Disconnected Alarm | Long | Alarms | unitless | Consult detailed help for more information on this variable | Common
92612 | O2 Supply Alarm | Long | Alarms | unitless | Consult detailed help for more information on this variable | Common
92631 | O2 Supply High Alarm | Long | Alarms | unitless | Consult detailed help for more information on this variable | Common
92632 | O2 Supply Low Alarm | Long | Alarms | unitless | Consult detailed help for more information on this variable | Common