## Sheet1
ID | Label | Type | Group | Unit | Description | Device | Source
26 | Device Inactivation Status | String | DeviceInfo | unitless | Possible values: "enabled" and "audio-paused". See "Multiset Alarms" on help file. | Common | Alarms
168 | Device ID | String | DeviceInfo | unitless | The globally unique identifier for a Bed entity as defined by Hill-Rom (e.g. “VersaCare-0x0001”). | Bed Data Server (Common) | BED
449 | Recirculation Measurement Active Status | Boolean | DeviceInfo | unitless | Possible values: 0=off, 1=on | Common
459 | Auto Dialysate Flow Enabled Status | Boolean | DeviceInfo | unitless | Possible values: 0=disabled, 1=enabled | Common
461 | Auto AntiCoagulant Enabled Status | Boolean | DeviceInfo | unitless | Possible values: 0=disabled, 1=enabled | Common
462 | Auto Anticoagulant Bolus Enabled Status | Boolean | DeviceInfo | unitless | Possible values: 0=off, 1=on | Common
463 | Auto Anticoagulant Bolus Rate Enabled Status | Boolean | DeviceInfo | unitless | Possible values: 0=disabled, 1=enabled | Common
464 | Single Needle Auto Rate Enabled Status | Boolean | DeviceInfo | unitless | Possible values: 0=disabled, 1=enabled | Common
468 | Replacement Pump Running Status | Boolean | DeviceInfo | unitless | Possible values: 0=off, 1=on | Common
538 | Device Language | String | DeviceInfo | unitless | Consult detailed help for more information on this variable | Common | System Information
584 | Ventilation Mode | String | DeviceInfo | unitless | The ventilation mode is reported in variable 584 and 2756. Variable 2756 contains mode ID numeric value, while variable 584 contains string representation which is the complete ventilation mode information concatenated from basic ventilation mode together with available options such as "Mode VC-CMV /AF". See variable 2756 for full list of supported ventilation mode. | Evita Infinity V500 [MEDIBUS.X]
754 | Device Name | String | DeviceInfo | unitless | This variable contains the device name defined by the manufacture. See variable 3591 for device model identification number. | Common | Device Information
762 | Predilution Enable Status | UShort | DeviceInfo | unitless | Internal status for the mixed HDF control | Common
764 | Device Specification ID | String (Multiset) | DeviceInfo | unitless | Variables 764, 765 and 4201 must be interpreted conjointly. | Common | System Information
765 | Device Specification Type | String (Multiset) | DeviceInfo | unitless | Variables 764, 765 and 4201 must be interpreted conjointly. | Common | System Information
845 | Clearance Measurement Enabled Status | Boolean | DeviceInfo | unitless | Possible values: 0=off, 1=on | Common
851 | Residual Blood Volume Active Status | Boolean | DeviceInfo | unitless | Possible values: 0=off, 1=on | Common
854 | Total Number of Used Disposable | Double | DeviceInfo | - |  | IntelliVue X3 | nFilt
860 | Auto Reinfusion Active Status | Boolean | DeviceInfo | unitless | Possible values: 0=off, 1=on | Common
914 | Low Volume Treatment Active Status | Boolean | DeviceInfo | unitless | Possible values: 0=off, 1=on | Common
980 | Base Device ID | String | DeviceInfo | unitless | The globally unique identifier for the BedConnector entity associated with this Bed entity or an empty string if no association currently exists. | Bed Data Server (Common) | BED
982 | Device IP Address | String | DeviceInfo | unitless |  | Common
1011 | Therapy Running Status | Long | DeviceInfo | unitless | Possible values: | Common
1033 | Air Bubble Detection State | Byte | DeviceInfo | unitless | Consult detailed help for more information on this variable | Common
1378 | Device Configuration | String | DeviceInfo | unitless | For CritiCool: This variable contains information on the fixed step rate at which the patient gradually warms up. For Allon: This variable contains information on if the temperature alarms are on or off.Consult detailed help for more information on this variable | Common
1450 | Infusion Pump Type | UByte | DeviceInfo | unitless | 0: Smart or Go 2: Vision or Go+ | Common
1451 | Pump Infusion Status | UByte | DeviceInfo | unitless | Consult detailed help for more information on this variable | Common
1510 | Concentrate Source | Long | DeviceInfo | unitless | Consult detailed help for more information on this variable | Common
1540 | TcGas Heating Power | Double | DeviceInfo | - |  | IntelliVue X3 | TCUT: HeatPw
1869 | Module State | UShort | DeviceInfo | unitless | This variable is 16 bits long. | Common
1891 | Power Supply Status | String | DeviceInfo | unitless | The current status of the AC power system. | Bed Data Server (Common) | BED
1911 | Device Unique ID | String | DeviceInfo | unitless |  | Common | System Information
1913 | Device Serial Number | String | DeviceInfo | unitless | Format: 8 characters string. This variable is buffered and sent repeatedly by the DDI with each data message. | Common
1927 | Instrument Code | String | DeviceInfo | unitless | Format: 6 characters string | Common
1946 | Remote Alarm Application Mode | String | DeviceInfo | unitless | The current configuration mode of the SafeView alarm system. | Bed Data Server (Common) | BED
1947 | Remote Alarm Application Indicator Status | String | DeviceInfo | unitless | The current status of the SafeView alarm system indicators. | Bed Data Server (Common) | BED
1948 | Patient Weight Scale Activity | String | DeviceInfo | unitless | The last command to be successfully processed by the weigh scale system. | Bed Data Server (Common) | BED
1957 | Remote CIS Application Status | String | DeviceInfo | unitless | The current status for the NaviCare alerts switch. | Bed Data Server (Common) | BED
1958 | Remote CIS Application Indicator Status | String | DeviceInfo | unitless | The current status for the NaviCare alerts indicator. | Bed Data Server (Common) | BED
1963 | Nurse Call Indicator Status | String | DeviceInfo | unitless | The current status for the nurse call indicator. | Bed Data Server (Common) | BED
1964 | Nurse Answer Indicator Status | String | DeviceInfo | unitless | The current status for the nurse answer indicator. | Bed Data Server (Common) | BED
1986 | Bed Height Position | String | DeviceInfo | unitless | The bed’s physical articulated position. | Bed Data Server (Common) | BED
1987 | Bed Lockout Status | String | DeviceInfo | unitless |  | Bed Data Server (Common) | BED
1988 | Bed Head Rails Position | String | DeviceInfo | unitless | The physical positions of the head-end siderails. | Bed Data Server (Common) | BED
1989 | Bed Foot RailsPosition | String | DeviceInfo | unitless | The physical positions of the foot-end siderails. | Bed Data Server (Common) | BED
1990 | Bed Surface Support Mode | String | DeviceInfo | unitless | The current configuration mode of the patient support surface system. | Bed Data Server (Common) | BED
1991 | Bed Emergency CPR Status | String | DeviceInfo | unitless | The current status for the CPR switch. | Bed Data Server (Common) | BED
1992 | Bed Clearance Status | String | DeviceInfo | unitless | The current status for the Bed Cleaned switch. | Bed Data Server (Common) | BED
1993 | Bed Rotation Therapy Status | String | DeviceInfo | unitless | The current status for the rotation therapy system. | Bed Data Server (Common) | BED
1994 | Bed Percussion Therapy Status | String | DeviceInfo | unitless | The current status for the percussion therapy system. | Bed Data Server (Common) | BED
1995 | Bed Vibration Therapy Status | String | DeviceInfo | unitless | The current status for the vibration therapy system. | Bed Data Server (Common) | BED
1996 | Patient Environment Control Activity | String | DeviceInfo | unitless | The last command completed by the patient environment control system. | Bed Data Server (Common) | BED
1997 | Patient Environment Activity Cleared Status | Byte | DeviceInfo | unitless | The last command completed by the patient environment control system. | Bed Data Server (Common) | BED
1998 | Bed Head Angle Alarm Mode | String | DeviceInfo | unitless | The current mode configuration for the head angle alarm system. | Bed Data Server (Common) | BED
1999 | Bed Head Angle Alarm Silence Status | String | DeviceInfo | unitless | The current mode configuration for the audible annunciation portion of the head angle alarm system. | Bed Data Server (Common) | BED
2000 | Bed Clearance Indicator Status | String | DeviceInfo | unitless | The current status for the bed cleaned indicator. | Bed Data Server (Common) | BED
2003 | Therapy Options | Long | DeviceInfo | unitless | Consult detailed help for more information on this variable | Common
2022 | Device System Code | String | DeviceInfo | unitless |  | Multifiltrate PRO | Treatment Status
2049 | Breath Cycle Type | Double | DeviceInfo | - |  | IntelliVue X3 | sCycTy
2200 | Device Hardware Revision | String | DeviceInfo | unitless | This variable is obsolete because it may contains incorrect information. Please use variables 764, 765 and 4201 must be used instead. | Common | Obsolete Variables
2201 | Serial Number | String | DeviceInfo | unitless | This variable is obsolete because it may contains incorrect information. Please use variables 764, 765 and 4201 must be used instead. | Common | Obsolete Variables
2255 | Device Date/Time | String | DeviceInfo | unitless |  | Common
2377 | Fresh Gas O2 Flow Rate | Double | DeviceInfo | L/min | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Evita V300. O2 therapy flow, Flow. For Evita Infinity V500, Babylog VN500: Flow O2 Therapy Flow. | Evita Infinity V500 [MEDIBUS.X]
2378 | Fresh Gas N2O Flow | Double | DeviceInfo | - | Mixer parameter. Setting of fresh gas N2O flow. | IntelliVue X3 | sfgN2O
2379 | Fresh Gas Air Flow | Double | DeviceInfo | - |  | IntelliVue X3 | fgAir
2382 | Ventilator Operation Mode | String | DeviceInfo | unitless | This variable contains ventilation operation mode in string format, which is also send as a numerical value in variable 2758. See detailed description of variable 2758 for the list of supported values. For Evita 4, Evita 2 dura, Evita XL: Firmware version 04.00 or higher For Evita 2: Variable label is not correct. Please use variable 3853 instead. Obsolete variable that may be removed in future release. | Evita Infinity V500 [MEDIBUS.X]
2477 | Battery Module Status | String | DeviceInfo | unitless | The current status of the battery power system. | Bed Data Server (Common) | BED
2486 | NIBP Interval Time | Long | DeviceInfo | min |  | Common
2537 | Citrate Calcium Activation Status | Boolean | DeviceInfo | unitless | Possible values: | Multifiltrate PRO | Dialysis Settings
2573 | Ventilator Control Mode | String | DeviceInfo | unitless | The variable indicates if the ventilator is in Non-Invasive Ventilation. ON means non-invasive ventilation and OFF means invasive ventilation. | Common
2754 | Ventilator General Status | UByte | DeviceInfo | unitless | Ventilator Standby mode status. | Evita Infinity V500 [MEDIBUS.X]
2756 | Ventilation Mode ID | UByte | DeviceInfo | unitless | This variable contains ID of basic ventilation mode ID without options. See also variable 584. | Evita Infinity V500 [MEDIBUS.X]
2758 | Ventilator Operation Mode ID | UByte | DeviceInfo | unitless | For Evita 4, Evita 2 dura, Evita XL: Firmware version 04.00 or higher For Evita 2: Variable label is not correct. Please use variable 2467 instead. Obsolete variable that may be removed in future release. | Evita Infinity V500 [MEDIBUS.X]
2819 | SvO2 Calibration Status | Boolean | DeviceInfo | unitless | ScvO2Cal: Point of time of calibration. "1" once during “Bloodsample drawn”, otherwise "0". | PiCCO2 | ScvO2 Calibration
2821 | System General Status | String | DeviceInfo | unitless | This variable identifies the current operating mode of the device. Possible values: "not specified", "monitoring", "demo", and "service". Moreover, "- standby" or "- config" can be appended | Common | System Information
2845 | Frequency | String | DeviceInfo | Hz | This variable describes the frequency of the main power supply. Possible values: "unspecified", "50", and "60". | Common | System Information
2855 | Device Status | UShort | DeviceInfo | unitless | Consult detailed help for more information on this variable | Multifiltrate PRO | Treatment Status
2879 | Device Software Revision | String | DeviceInfo | unitless | Format: vv.r.c vv = software version r = software revision c = software correction | Common
2880 | Maximum Apnea Time Setting | Double | DeviceInfo | - |  | IntelliVue X3 | sTAmax
2899 | Flow Sensor Type | UByte | DeviceInfo | unitless | Consult detailed help for more information on this variable | Common
2952 | PS Activation State | Boolean | DeviceInfo | unitless | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300, Fabius Family. A value of 1 indicates that Pressure Support (PS) is added to the current ventilation mode. A value of 0 indicates that Pressure Support (PS) is not added to the current ventilation mode. Text message displayed on device screen: "/PS" | Evita Infinity V500 [MEDIBUS.X]
2953 | VS Activation State | Boolean | DeviceInfo | unitless | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. A value of 1 indicates that Volume support (VS) is added to the current ventilation mode. A value of 0 indicates that Volume support (VS) is not added to the current ventilation mode. Text messages displayed on device screen: "/VS" | Evita Infinity V500 [MEDIBUS.X]
2969 | Disposable Lifetime | Long | DeviceInfo | min | The resolution of this variable is 10 min. This variable is sent from the device to the DDI on a regular basis, but the device will keep on sending the same value until a minimal change of 10 min is observed. | Multifiltrate PRO | Dialysis Course
2995 | Device Altitude | Double | DeviceInfo | m |  | Common | System Information
3050 | Therapy Type | String | DeviceInfo | unitless | Consult detailed help for more information on this variable | Multifiltrate PRO | Treatment Status
3052 | Device Screen State | UByte | DeviceInfo | unitless | Possible values : | Common
3496 | TcGas Transducer Temperature | Double | DeviceInfo | - |  | IntelliVue X3 | TCUT: SensrT
3503 | Auto Flow Active Status | Boolean | DeviceInfo | unitless | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300. A value of 1 indicates that AutoFlow (AF) is added to the current ventilation mode. A value of 0 indicates that AutoFlow (AF) is not added to the current ventilation mode. Text message displayed on device screen: "/AF" | Evita Infinity V500 [MEDIBUS.X]
3555 | ATC Activation Status | Boolean | DeviceInfo | unitless | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Evita V300. A value of 1 indicates that automatic tube compensation (ATC) is added to the current ventilation mode. A value of 0 indicates that automatic tube compensation (ATC) is not added to the current ventilation mode. Text message displayed on device screen: "/ATC" For Savina 300 (MEDIBUS.X): added with SW 5.00 | Evita Infinity V500 [MEDIBUS.X]
3590 | Device Type | String | DeviceInfo | unitless |  | Common
3591 | Device Number | String | DeviceInfo | unitless | Consult detailed help for more information on this variable | Common | Device Information
3593 | Device Communications Status | String | DeviceInfo | unitless | Value "stop" sent when a request to stop the communication is received from the device | Common | Device Information
3615 | Battery Charge | UShort | DeviceInfo | unitless | 58112 => battery charged | Common
3635 | Drug Delivery Rate | Double | DeviceInfo | - |  | IntelliVue X3 | sDlvRt
3639 | Manufacturers Name | String | DeviceInfo | unitless | "Philips" | Common | System Information
3640 | Pressure Alarm Delay | Double | DeviceInfo | - |  | IntelliVue X3 | sPADel
3696 | Generic Pump Running Status | ULong | DeviceInfo | unitless | This variable is 32 bits long. | Common
3829 | Device Total Operation Time | String | DeviceInfo | unitless | Format: [hours]:[minutes] | Multifiltrate PRO | Treatment Status
3833 | Device Protocol Mode | String | DeviceInfo | unitless | The connectivity protocol supported by the bed communication controller. | Bed Data Server (Common) | BED
3877 | Device Protocol Revision | String | DeviceInfo | unitless | Release PiCCO2 software version, composed by one letter (type of protocol: V for usual protocol, D for demo version and X for unreleased version) and from one to three digits seperated by dots (version number). For example: D5.1 | PiCCO2 | Device Information
3998 | Auto Release Activation Status | Boolean | DeviceInfo | unitless | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. A value of 1 indicates that AutoRelease is added to the current ventilation mode. A value of 0 indicates that AutoRelease is not added to the current ventilation mode. Text messages displayed on device screen: "/AutoRelease" | Evita Infinity V500 [MEDIBUS.X]
4044 | Medibus Revision | String | DeviceInfo | unitless |  | Common | Device Information
4045 | Device Revision | String | DeviceInfo | unitless | Format: mm.nn | Common | Device Information
4150 | VG Activation Status | Boolean | DeviceInfo | unitless | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Carina, Evita V300. A value of 1 indicates that Volume Guarantee (VG) is added to the current ventilation mode. A value of 0 indicates that Volume Guarantee (VG) is not added to the current ventilation mode. Text message displayed on device screen: "/VG" | Evita Infinity V500 [MEDIBUS.X]
4201 | Product Descritpion | String | DeviceInfo | unitless | 4 characters for product identification | Common
4281 | Nurse Call Status | String | DeviceInfo | unitless | The current status for the nurse call switch. | Bed Data Server (Common) | BED
4365 | Device Communication Mode | UByte | DeviceInfo | unitless | Consult detailed help for more information on this variable | Common
4473 | Device Nomenclature Revision | String | DeviceInfo | unitless | Format: major.minor (such as 1.0) | Common | System Information
4474 | Device Text Catalog Revision | String | DeviceInfo | unitless | Format: "MM.NN mm.nn", where "MM.NN" is text catalog revision and "mm.nn" is language revision. | Common | System Information
4475 | Device Data Format | String | DeviceInfo | unitless | Always 11 (Unicode). | Common | System Information
4476 | Device Operating Area | String | DeviceInfo | unitless | Possible values: "unspecified", "operating room", "intensive care unit", "neonatal intensive care unit", and "cardiology care unit". | Common | System Information
4570 | Dialyzer Name | String | DeviceInfo | unitless |  | Common
4573 | Dialysis Filter Lifetime | Long | DeviceInfo | min | Filter Lifetime | Multifiltrate PRO | Dialysis Course
4661 | Pleth Gain | Double | DeviceInfo | - |  | IntelliVue X3 | PlGain
4830 | Device Simulator Status | Double | DeviceInfo | - | Device data with the values: Passed and Failed. | IntelliVue X3 | Available only in revF or earlier
4966 | Sigh Activation Status | Boolean | DeviceInfo | unitless | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. A value of 1 indicates that the sigh function has been added to the current ventilation mode. A value of 0 indicates that the sigh function has not been added to the current ventilation mode. Text message displayed on device screen: "/Sigh" | Evita Infinity V500 [MEDIBUS.X]
4989 | O2 Monitoring Mode | Double | DeviceInfo | - |  | IntelliVue X3 | VENT: sO2Mon
4990 | O2 Calibration Mode | Double | DeviceInfo | - |  | IntelliVue X3 | VENT: sO2Cal
5005 | Mechanical Ventilation Activated Status | Double | DeviceInfo | - | Controlled Mechanical Ventilation Setting | IntelliVue X3 | VENT: sCMV
5006 | SIMV Mode | Double | DeviceInfo | - | Synchronized Intermittent Mandatory Ventilation Setting | IntelliVue X3 | VENT: sSIMV
5013 | Apnea Alarm Delay | Double | DeviceInfo | - |  | IntelliVue X3 | sAADel
5041 | Gas Measurement Sensor Position | Double | DeviceInfo | - |  | IntelliVue X3 | VENT: sGasPr
5042 | O2 Measurement Sensor Position | Double | DeviceInfo | - |  | IntelliVue X3 | VENT: sO2Pr
5059 | NMT TOF Repetition Time | Double | DeviceInfo | - | Preset Train of Four (Slow TOF) Repetition Time Setting | IntelliVue X3 | NMT: sRepTi
5497 | Pump Infusion Mode | Double | DeviceInfo | - |  | IntelliVue X3 | sPuMod
5642 | SvO2 Calibration Data | Double | DeviceInfo | unitless | ScvO2Cal: ScvO2 Calibration input. | PiCCO2 | ScvO2 Calibration
5858 | Disinfect Mode Active Status | UByte | DeviceInfo | unitless | Consult detailed help for more information on this variable | Common
5859 | Dialysis Mode Active Status | Long | DeviceInfo | unitless | Consult detailed help for more information on this variable | Common
5860 | Blood Sensed Status | Long | DeviceInfo | unitless | 0=no blood sensed 1=blood sensed | Multifiltrate PRO | Dialysis Course
6334 | Temperature Module General Status | Boolean | DeviceInfo | unitless | Possible values: 0=off, 1=on | Common
6388 | Endotracheal Tube Type | Double | DeviceInfo | - |  | IntelliVue X3 | sTube
6389 | Endotracheal Tube Size | Double | DeviceInfo | mm | For Evita 4, Evita 2 dura, Evita XL: Possible units: "0.1mm" in firmware versions 7.00 and 7.01, "mm" otherwise. Available only with EvitaXL firmware version 7.00 or higher. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Evita V300. Inner diameter of the endotracheal/tracheostoma tube. For Evita Infinity V500, Babylog VN500: ATC Tube Diameter. | Evita Infinity V500 [MEDIBUS.X]
6427 | Last Desinfection Time | String | DeviceInfo | unitless | Format: YYYYMMDDhhmm | Common
6428 | PICCO Catheter | Short | DeviceInfo | unitless | Range 0-14 Consult detailed help for more information on this variable | Common | Device Information
6429 | Last Therapy End Time | String | DeviceInfo | unitless | Format: YYYYMMDDhhmm | Common
6430 | Dialysis Pump Stop Status | UShort | DeviceInfo | unitless | Consult detailed help for more information on this variable | Common
6768 | Device Quick Test Status | String | DeviceInfo | unitless | T0 Test Status. Possible values: Start, End, E1, E2, E3. | Multifiltrate PRO | Treatment Status
6772 | Heater Power | Double | DeviceInfo | - |  | IntelliVue X3 | sPower
6967 | Ventilation Mode Additional Information | Double | DeviceInfo | - |  | IntelliVue X3 | sVSupp
7452 | Device Operation Mode | String | DeviceInfo | unitless | Consult detailed help for more information on this variable | Common
7587 | Endotracheal Tube Diameter | Double | DeviceInfo | - |  | IntelliVue X3 | sTubeD
7637 | Artificial Airway Compensation Status | Double | DeviceInfo | % | For Evita 4, Evita 2 dura, Evita XL: Tube Compensation. Available only with EvitaXL firmware version 7.00 or higher. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. Compensation rate during ATC, % comp. For Evita Infinity V500, Babylog VN500: ATC Compensation Compens. | Evita Infinity V500 [MEDIBUS.X]
7765 | Device Power Up Time | Double | DeviceInfo | s | Number of seconds since the Monitor device's last power-on. | Common | System Information
8244 | Bolus Manual Enable Status | Boolean | DeviceInfo | unitless | Possible values: 0=disabled, 1=enabled | Common
8261 | Patient Expiratory Port Type | Long | DeviceInfo | unitless | Consult detailed help for more information on this variable | Common
8646 | Patient Interface Type | Long | DeviceInfo | unitless | Consult detailed help for more information on this variable | Common
8856 | Therapy Operation Status | Long | DeviceInfo | unitless | Consult detailed help for more information on this variable | Common
9154 | Humidifier Activation Status | Double | DeviceInfo | - |  | IntelliVue X3 | sHmCtl
9221 | Error Response From Device | UByte | DeviceInfo | unitless | Error code sent by the pump when an incorrect transmission happens. | Common
9522 | Device Manufacturer | String | DeviceInfo | unitless |  | Common
9526 | Bed Head Angle | Double | DeviceInfo | ° | The articulated angle of the bed’s head section in degrees. -360.0 -- 360.0 | Bed Data Server (Common) | BED
9530 | Bed Brake Status | String | DeviceInfo | unitless | The current status for the brake switch. | Bed Data Server (Common) | BED
9662 | Device Location | String | DeviceInfo | unitless |  | Common
50021 | Device Alarm Inactivation State | String | DeviceInfo | unitless |  | Evita Infinity V500 [MEDIBUS.X]
50201 | Device Type | UByte | DeviceInfo | unitless | This variable contains a fixed value "48 (0x30)" for Infusion Pump. | Common | Recommended Variables
50203 | Device Model | String | DeviceInfo | unitless |  | Common
50212 | Device Software Revision | String | DeviceInfo | unitless |  | Common
53001 | Infusion Pump Type | UByte | DeviceInfo | unitless | Consult detailed help for more information on this variable | Common | Recommended Variables
53012 | Infusion Pump Serial Number | String | DeviceInfo | unitless | 8 characters string | Common | Recommended Variables
53021 | Infusion Pump Rack Index | UByte | DeviceInfo | unitless | The variable has a fixed value "1" because it is a standalone pump. | Common | Recommended Variables
53022 | Infusion Pump Slot Index | UByte | DeviceInfo | unitless | The variable has a fixed value "1" because it is a standalone pump. | Common | Recommended Variables