## Sheet1
ID | Label | Unit | Description | Type | Group | Device | Source
1 | Heart Rate | 1/min | This variable is associated with LowerAlarmLimit and UpperAlarmLimit attributes. | Long | VitalSigns | Common
2 | NIBP Systolic | mmHg | This variable is associated with LowerAlarmLimit and UpperAlarmLimit attributes. | Long | VitalSigns | Common
3 | NIBP Diastolic | mmHg | This variable is associated with LowerAlarmLimit and UpperAlarmLimit attributes. | Long | VitalSigns | Common
4 | NIBP Mean | mmHg | This variable is associated with LowerAlarmLimit and UpperAlarmLimit attributes. | Long | VitalSigns | Common
7 | Patient Delta Temperature | °C |  | Double | VitalSigns | Common
356 | Generic Dialysate Alarm | unitless | Value of "1" when alarm is active. | Boolean | Alarms | Common
357 | Urea Clearance Alarm | unitless | Value of "1" when alarm is active. | Boolean | Alarms | Common
358 | Blood Volume Alarm | unitless | Value of "1" when alarm is active. | Boolean | Alarms | Common
359 | Mixed Dilution Alarm | unitless | Value of "1" when alarm is active. | Boolean | Alarms | Common
393 | Single Needle Ratio Setting | % |  | Long | Settings | Common
444 | Target Treatement Effectiveness (Kt/V) Remaining Time | s |  | Long | VitalSigns | Common
448 | Shunt Blood Flow Rate | mL/min |  | Short | VitalSigns | Common
449 | Recirculation Measurement Active Status | unitless | Possible values: 0=off, 1=on | Boolean | DeviceInfo | Common
450 | Isolated Ultrafiltration Therapy Time | s |  | Long | VitalSigns | Common
451 | Isolated Ultrafiltration Rate | mL/h |  | Long | VitalSigns | Common
452 | Initial Hemoglobin | g/dL |  | Double | VitalSigns | Common
453 | Initial Protein Concentration | g/dL |  | Double | VitalSigns | Common
454 | Single Needle Time | min |  | Long | VitalSigns | Common
455 | Blood Outlet Pressure | mmHg |  | Short | VitalSigns | Common
456 | Dialysate Inlet Pressure | mmHg |  | Short | VitalSigns | Common
457 | Dialysate Outlet Pressure | mmHg |  | Short | VitalSigns | Common
458 | Maximum Transmembrane Pressure | mmHg |  | Short | VitalSigns | Common
459 | Auto Dialysate Flow Enabled Status | unitless | Possible values: 0=disabled, 1=enabled | Boolean | DeviceInfo | Common
460 | Auto Dialysate Flow Factor Setting | unitless |  | Double | Settings | Common
461 | Auto AntiCoagulant Enabled Status | unitless | Possible values: 0=disabled, 1=enabled | Boolean | DeviceInfo | Common
462 | Auto Anticoagulant Bolus Enabled Status | unitless | Possible values: 0=off, 1=on | Boolean | DeviceInfo | Common
463 | Auto Anticoagulant Bolus Rate Enabled Status | unitless | Possible values: 0=disabled, 1=enabled | Boolean | DeviceInfo | Common
464 | Single Needle Auto Rate Enabled Status | unitless | Possible values: 0=disabled, 1=enabled | Boolean | DeviceInfo | Common
468 | Replacement Pump Running Status | unitless | Possible values: 0=off, 1=on | Boolean | DeviceInfo | Common
508 | Clearance Measurement Interval Setting | s |  | UShort | Settings | Common
540 | Low Venous Line Pressure Alarm | unitless | Value of "1" when alarm is active. | Boolean | Alarms | Common
543 | High Venous Line Pressure Alarm | unitless | Value of "1" when alarm is active. | Boolean | Alarms | Common
546 | High Arterial Line Pressure Alarm | unitless | Value of "1" when alarm is active. | Boolean | Alarms | Common
547 | Low Arterial Line Pressure Alarm | unitless | Value of "1" when alarm is active. | Boolean | Alarms | Common
558 | Heparin Pump Alarm | unitless | Value of "1" when alarm is active. | Boolean | Alarms | Common
569 | Patient Specific Filtration Ratio Setting | % |  | UShort | Settings | Common
576 | Average Dialysate Temperature | °C |  | Double | VitalSigns | Common
610 | Cuff Inflation Pressure | mmHg |  | UShort | OtherMeasuredData | Common
676 | Bicarbonate Concentration (HCO3) | mmol/L |  | Double | VitalSigns | Common
754 | Device Name | unitless |  | String | DeviceInfo | Common
757 | Patient Weight | kg |  | Double | PatientInfo | Common
762 | Predilution Enable Status | unitless | Internal status for the mixed HDF control | UShort | DeviceInfo | Common
Consult detailed help for more information on this variable
845 | Clearance Measurement Enabled Status | unitless | Possible values: 0=off, 1=on | Boolean | DeviceInfo | Common
846 | Maximum Ultrafiltration Rate Setting | mL/h |  | Long | Settings | Common
847 | Ultrafiltration Goal Positive Deviation Setting | mL |  | Long | Settings | Common
848 | Ultrafiltration Goal Negative Deviation Setting | mL |  | Long | Settings | Common
851 | Residual Blood Volume Active Status | unitless | Possible values: 0=off, 1=on | Boolean | DeviceInfo | Common
852 | Temperature Change Rate Setting | °C/h |  | Double | Settings | Common
859 | Reinfusion Type Setting | unitless | Consult detailed help for more information on this variable | UShort | Settings | Common
860 | Auto Reinfusion Active Status | unitless | Possible values: 0=off, 1=on | Boolean | DeviceInfo | Common
914 | Low Volume Treatment Active Status | unitless | Possible values: 0=off, 1=on | Boolean | DeviceInfo | Common
955 | Battery Running Alarm | unitless | Value of "1" when alarm is active. | Boolean | Alarms | Common
965 | Cumulated Therapy Time | s |  | Long | OtherMeasuredData | Common
966 | Current Therapy Time | s |  | Long | OtherMeasuredData | Common
967 | Current Therapy ID | unitless |  | String | PatientInfo | Common
982 | Device IP Address | unitless |  | String | DeviceInfo | Common
1011 | Therapy Running Status | unitless | Consult detailed help for more information on this variable | String | DeviceInfo | Common
1043 | Manually Entered Hematocrit | % |  | UShort | Settings | Common
1090 | Anticoagulant Cumulated Bolus Infused Volume | mL |  | Double | VitalSigns | Common
1100 | Average Anticoagulant Rate | mL/h |  | Double | VitalSigns | Common
1101 | Single Needle Stroke Volume | mL |  | Long | Settings | Common
1102 | Average Urea Clearance | mL/min |  | Long | VitalSigns | Common
1103 | Online Bolus Volume | mL |  | Long | VitalSigns | Common
1104 | Online Bolus Rate Setting | mL/min |  | Long | Settings | Common
1220 | Patient Birth Date | unitless | Format: YYYYMMDD | String | PatientInfo | Common
1246 | Hematocrit | % |  | Double | VitalSigns | Common
1463 | Hemoglobin (Hb) | g/dL |  | Double | VitalSigns | Common
1510 | Concentrate Source | unitless | Consult detailed help for more information on this variable | Long | DeviceInfo | Common
1525 | Maximum Fluid Removal Rate Setting | mL/h |  | Long | Settings | Common
1591 | Average Blood Flow Rate | mL/min |  | Long | VitalSigns | Common
1602 | Reinfusion Volume | mL |  | Long | VitalSigns | Common
1803 | Device Alarm Global Status | unitless | Format: Bit 2: 0 --> Alarm | String | Alarms | Common
1822 | Blood Pump Alarm | unitless | Value of "1" when alarm is active. | Boolean | Alarms | Common
1913 | Device Serial Number | unitless |  | String | DeviceInfo | Common
1929 | Patient ID | unitless |  | String | PatientInfo | Common
2003 | Therapy Options | unitless | Consult detailed help for more information on this variable | Long | DeviceInfo | Common
2486 | NIBP Interval Time | min |  | Long | DeviceInfo | Common
2499 | Relative Blood Volume Change Limit | % |  | Long | Settings | Common
2590 | Generic Device Alarm | unitless | Format: list of maximum 16 alarm IDs separated with a "¦". | String | Alarms | Common
Consult detailed help for more information on this variable
2847 | Urea Distribution Volume Setting | L |  | Double | Settings | Common
2848 | Auto Ultrafiltration Control Activation Status | unitless | Possible values: 0=off, 1=on | Boolean | Settings | Common
2879 | Device Software Revision | unitless |  | String | DeviceInfo | Common
2901 | Patient First Name | unitless |  | String | PatientInfo | Common
3003 | Blood Flow Rate Setting | mL/min |  | UShort | Settings | Common
3004 | Replacement Flow Rate Setting | mL/min |  | Long | Settings | Common
3005 | Dialysate Flow Rate Setting | mL/min |  | ULong | Settings | Common
3006 | Fluid Removal Volume Setting | mL |  | Long | Settings | Common
3029 | Transmembrane Pressure | mmHg |  | Long | VitalSigns | Common
3070 | Patient Profile | unitless | Na profile number | UByte | PatientInfo | Common
3071 | Therapy Profile | unitless |  | UByte | PatientInfo | Common
3100 | Dialysate Sodium Start Concentration Setting | mmol/L |  | Double | Settings | Common
3168 | Fluid Removal Actual Rate | mL/h |  | Long | VitalSigns | Common
3189 | Average Replacement Flow Rate | mL/min |  | Long | VitalSigns | Common
3190 | Minimum Relative Treated Blood Volume | % |  | Double | VitalSigns | Common
3194 | Auto Replacement Control Activation Status | unitless | Possible values: 0=disabled, 1=enabled | Boolean | Settings | Common
3237 | Relative Blood Volume Change | % |  | Double | VitalSigns | Common
3239 | Replacement Volume Setting | L |  | Double | Settings | Common
3248 | Generic Venous Line Pressure | mmHg |  | Long | VitalSigns | Common
3252 | Total Replacement Cumulated Volume | L | Expected total substitution volume (pre + post) at the end of the treatment. | Double | VitalSigns | Common
3253 | Blood Circuit Pressure Alarm | unitless | Value of "1" when alarm is active. | Boolean | Alarms | Common
3263 | Recirculation Percentage | % |  | Double | VitalSigns | Common
3269 | Filtration Ratio | % |  | UShort | VitalSigns | Common
3417 | Ultrafiltration Remaining Time | min | Remaining ISO UF time. | Long | VitalSigns | Common
3430 | Total Replacement Flow Rate | mL/min |  | Long | VitalSigns | Common
3437 | Replacement Post-delivered Percentage | % |  | UShort | VitalSigns | Common
3575 | Arterial Temperature | °C |  | Double | VitalSigns | Common
3576 | Venous Temperature | °C |  | Double | VitalSigns | Common
3590 | Device Type | unitless |  | String | DeviceInfo | Common
3626 | Therapy Ending Time | unitless | Format: YYYYMMDDhhmm | String | OtherMeasuredData | Common
3628 | Therapy Starting Time | unitless | Format: YYYYMMDDhhmm | String | OtherMeasuredData | Common
4229 | Fluid Removal Cumulated Volume | mL |  | ULong | VitalSigns | Common
4238 | Single Needle Max Pressure Setting | mmHg |  | Long | Settings | Common
4239 | Single Needle Min Pressure Setting | mmHg |  | Long | Settings | Common
4240 | Plasma Sodium (Na) | mmol/L | For 6008 only: This variable will contain the Online Clearance Monitoring plasma sodium concentration when sodium management is being used. | Double | VitalSigns | Common
4347 | Therapy Remaining Time | min | Remaining UF time including ISO UF time. | ULong | OtherMeasuredData | Common
4365 | Device Communication Mode | unitless | Consult detailed help for more information on this variable | UByte | DeviceInfo | Common
4570 | Dialyzer Name | unitless |  | String | DeviceInfo | Common
4579 | Dialysate Name | unitless |  | String | PatientInfo | Common
4592 | Dialysate Average Flow Rate | mL/min |  | ULong | VitalSigns | Common
4601 | Dialysate Sodium (Na+) Concentration Setting | mmol/L |  | Double | Settings | Common
4605 | Dialysate Bicarbonate (HCO3) Concentration Setting | mmol/L |  | Double | Settings | Common
4608 | Urea Clearance | mL/min |  | Short | VitalSigns | Common
4609 | Fluid Removal Average Rate | mL/h |  | Long | VitalSigns | Common
4610 | Anticoagulant Stopping Interval Setting | s |  | Long | Settings | Common
4614 | Target Treatement Effectiveness (Kt/V) Setting | unitless |  | Double | Settings | Common
4615 | Initial Hematocrit | % |  | Double | VitalSigns | Common
4685 | Serum Protein | g/dL |  | Double | VitalSigns | Common
4705 | Total Protein | g/dL |  | Double | VitalSigns | Common
5318 | Blood Temperature Alarm | unitless | Value of "1" when alarm is active. | Boolean | Alarms | Common
5845 | Total Replacement Cumulated Volume | mL | Cumulated total substitution volume (pre + post). | Long | VitalSigns | Common
5846 | Single Needle Blood Pump Rate | mL/min |  | Long | VitalSigns | Common
5849 | Therapy Duration Setting | s |  | Long | Settings | Common
5850 | Anticoagulant Cumulated Volume | mL |  | Double | VitalSigns | Common
5853 | Delivered Treatment Effectiveness (Kt/V) | unitless |  | Double | VitalSigns | Common
5856 | Dialysate Flow Rate | mL/min |  | ULong | VitalSigns | Common
5858 | Disinfect Mode Active Status | unitless | Consult detailed help for more information on this variable | UByte | DeviceInfo | Common
5859 | Dialysis Mode Active Status | unitless | Consult detailed help for more information on this variable | Long | DeviceInfo | Common
5861 | Effective Blood Volume Processed (Kt) | L |  | Double | VitalSigns | Common
5863 | Dialysate Temperature | °C |  | Double | VitalSigns | Common
5868 | Fluid Removal Rate Setting | mL/h |  | Long | Settings | Common
5869 | Dialysate Temperature Setting | °C |  | Double | Settings | Common
5884 | Replacement Cumulated Bolus Volume | mL |  | Long | VitalSigns | Common
5899 | Actual Anticoagulant Rate | mL/h |  | Double | VitalSigns | Common
5905 | Therapy Processed Percentage | % |  | Byte | VitalSigns | Common
6328 | Blood Plasma Inlet Pressure | mmHg |  | Short | VitalSigns | Common
6334 | Temperature Module General Status | unitless | Possible values: 0=off, 1=on | Boolean | DeviceInfo | Common
6427 | Last Desinfection Time | unitless | Format: YYYYMMDDhhmm | String | DeviceInfo | Common
6429 | Last Therapy End Time | unitless | Format: YYYYMMDDhhmm | String | DeviceInfo | Common
6430 | Dialysis Pump Stop Status | unitless | Consult detailed help for more information on this variable | UShort | DeviceInfo | Common
7351 | Isolated Ultrafiltration Volume | mL |  | Long | VitalSigns | Common
7352 | Single Needle Stroke Volume | mL | Average single needle stroke volume | Long | VitalSigns | Common
7353 | Initial Ultrafiltration Factor Setting | unitless |  | UShort | Settings | Common
7438 | Generic Arterial Line Pressure | mmHg |  | Long | VitalSigns | Common
7979 | Dialysate Conductivity | mS/cm |  | Double | VitalSigns | Common
7980 | Dialysate Conductivity Alarm | unitless | Value of "1" when alarm is active. | Boolean | Alarms | Common
8244 | Bolus Manual Enable Status | unitless | Possible values: 0=disabled, 1=enabled | Boolean | DeviceInfo | Common
8262 | Plasma Filtration Fraction Setting | % |  | UShort | Settings | Common
8340 | Patient Family Name | unitless |  | String | PatientInfo | Common
8803 | Replacement Pre-delivered Rate Setting | mL/min |  | UShort | Settings | Common
8804 | Replacement Post-delivered Rate Setting | mL/min |  | UShort | Settings | Common
8807 | Anticoagulant Bolus Volume Setting | mL |  | Double | Settings | Common
8809 | Anticoagulant Infusion Rate Setting | mL/h |  | Double | Settings | Common
8813 | Processed Cumulated Blood Volume | L |  | Double | VitalSigns | Common
8815 | Replacement Post-delivered Cumulated Volume | mL |  | Long | VitalSigns | Common
8856 | Therapy Operation Status | unitless | Consult detailed help for more information on this variable | Short | DeviceInfo | Common
9522 | Device Manufacturer | unitless |  | String | DeviceInfo | Common
9643 | Blood Flow Rate | mL/min |  | Long | VitalSigns | Common
9662 | Device Location | unitless |  | String | DeviceInfo | Common