## Sheet1
ID | Label | Unit | Description | Type | Group | Device | Source
20 | End Tidal CO2 (EtCO2) | - | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Evita V300, Fabius Family. This variable is associated with UpperAlarmLimit attribute, which is available for Perseus A500, Atlan, Evita Infinity V500, Babylog VN500, Savina 300, Primus, Primus IE, Apollo, Zeus IE, Oxylog 3000 plus, Evita V300. This variable is associated with LowerAlarmLimit attribute, which is available for Perseus A500, Atlan, Evita Infinity V500, Babylog VN500, Savina 300, Primus, Primus IE, Apollo, Zeus IE, Oxylog 3000 plus, Evita V300. Possible units: mmHg, %, kPa. This variable is in Multiset. See variables 2355 and 6551 for concentration and partial pressure respectively. For Evita Infinity V500, Babylog VN500: This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. end tidal CO2 etCO2 Note: etCO2 values and alarm limits are transmitted according to the configuration on the device screen (mmHg or Vol%). If kPa is configured be aware that the Medibus alarm limits will be transmitted with mmHg. | Double (Multiset) | VitalSigns | Evita Infinity V800
60 | Tidal Volume | mL | For Evita 4, Evita 2 dura, Evita XL: Expiratory tidal volume. This variable is associated with UpperAlarmLimit attribute, which is also sent in separate variable 2763. Values below 1 L will be transmitted in units of mL, otherwise in units of L. Device firmware version 04.00 and higher: Limit is not transmitted, if alarm is switched off. For Savina: This variable is associated with UpperAlarmLimit attribute.Device SW version 02.00 and higher: Limit is not transmitted, if alarm is switched off. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300, Fabius Family. This variable is associated with UpperAlarmLimit attribute, which is available for Perseus A500, Evita Infinity V500, Savina 300, Evita V300. This variable is associated with LowerAlarmLimit attribute, which is available for Evita Infinity V500, Evita V300. For Evita Infinity V500, Babylog VN500: This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. For Zeus, Zeus IE, Primus, Primus IE, Apollo, Pallas: Exp. Tidal Volume. | Double | VitalSigns | Evita Infinity V800
61 | Dynamic Compliance | L/bar | For Evita 4, Evita 2 dura, Evita XL: Values below 100 will be transmitted with one decimal place, otherwise without a decimal place. For Savina: Available since device SW version 02.00. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Evita V300. | Double | VitalSigns | Evita Infinity V800
114 | CO2 Concentration Waveform | - | For Evita 4, Evita 2 dura, Evita XL: Possible units: mmHg, kPa, and % For Evita 2: Possible units: mmHg, kPa, and % Firmware version 2.00 or higher. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Evita V300. Possible units: mmHg, %, kPa. Unit "mmHg" is available for Perseus A500, Atlan, Evita Infinity V500, Babylog VN500, Savina 300, Primus, Primus IE, Apollo, Zeus IE, Oxylog 3000 plus, Evita V300. Unit "kPa" and "%" are available for Perseus A500, Atlan, Savina 300, Primus, Primus IE, Apollo, Zeus IE, Oxylog 3000 plus. See variables 8866 and 7872 for concentration and partial pressure respectively. For Evita Infinity V500, Babylog VN500: Possible units: mmHg, %. Expiratory CO2 concentration. For Zeus, Zeus IE, Primus, Primus IE, Apollo, Pallas, Vamos, Vamos plus: Possible units: mmHg, %, kPa. | Double | Waveforms | Evita Infinity V800
118 | Airway Flow Wave | L/min | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300, Fabius Family. For Evita Infinity V500, Babylog VN500, Carina, Zeus, Zeus IE, Primus, Primus IE, Apollo, Pallas: Flow (insp./exp.) | Double | Waveforms | Evita Infinity V800
582 | Resistance | mbar/L/s | For Evita 4, Evita 2 dura, Evita XL: Values below 100 will be transmitted with one decimal place, otherwise without a decimal place. For Savina: Available since device SW version 02.00. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Zeus IE, Evita V300. | Double | VitalSigns | Evita Infinity V800
584 | Ventilation Mode | unitless | The ventilation mode is reported in variable 584 and 2756. Variable 2756 contains mode ID numeric value, while variable 584 contains string representation which is the complete ventilation mode information concatenated from basic ventilation mode together with available options such as "Mode VC-CMV /AF". See variable 2756 for full list of supported ventilation mode. | String | DeviceInfo | Evita Infinity V800
635 | Inspired O2 (FiO2) | % | For Savina: This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. Available since device SW version 03.00. Limits are transmitted in LPO-Mode (Option) only. For Fabius GS, Fabius GS premium, Fabius Tiro, Fabius plus, Fabius plus XL, Fabius OS, Fabius MRI: This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Evita V300, Fabius Family. This variable is associated with UpperAlarmLimit attribute, which is available for Perseus A500, Atlan, Primus, Primus IE, Apollo, Zeus IE, Fabius Family. This variable is associated with LowerAlarmLimit attribute, which is available for Perseus A500, Atlan, Primus, Primus IE, Apollo, Zeus IE, Fabius Family. Inspiratory oxygen fraction, FiO2. For Zeus, Zeus IE: This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. For Primus, Primus IE, Apollo, Pallas: This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. | Double | VitalSigns | Evita Infinity V800
754 | Device Name | unitless | This variable contains the device name defined by the manufacture. See variable 3591 for device model identification number. | String | DeviceInfo | Common | Device Information
775 | Inspiratory Component of I:E Ratio | unitless | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: Obsolete variable that may be removed in future release. Partial information of I:E ratio. For Evita Infinity V500, Savina: Partial information. Please use variable 1190 instead. Obsolete variable that may be removed in future release. | Double | VitalSigns | Evita Infinity V800
776 | Expiratory Component of I:E Ratio | unitless | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: Obsolete variable that may be removed in future release. Partial information of I:E ratio. For Evita Infinity V500, Savina: Partial information. Please use variable 1190 instead. Obsolete variable that may be removed in future release. | Double | VitalSigns | Evita Infinity V800
1189 | Positive End Expiratory Pressure (PEEP) | mbar | For Fabius GS, Fabius GS premium, Fabius Tiro, Fabius plus, Fabius plus XL, Fabius OS, Fabius MRI: 1mbar = 1cmH2O. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300, Fabius Family. | Double | VitalSigns | Evita Infinity V800
1190 | I:E Ratio | unitless | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Evita V300. Format: 1:x if E > I, x:1 if I > E, or 1:1 if I = E. The useless leading or ending “0” or “.” is not transmitted, and the maximum digits after the floating point is 2. For Evita Infinity V500, Babylog VN500: Format: 1:x if E > I, x:1 if I > E, or 1:1 if I = E. The useless leading or ending “0” or “.” is not transmitted, and the maximum digits after the floating point is 2. | String | VitalSigns | Evita Infinity V800
1307 | Minute Volume | L/min | For Evita 4, Evita 2 dura, Evita XL: This variable is associated with UpperAlarmLimit attribute, which is also sent in separate variable 1400. This variable is associated with LowerAlarmLimit attribute, which is also sent in separate variable 1399. For device firmware version 04.00 and higher: Lower alarm limit is not transmitted, if alarm is switched off. Values may be transmitted with one or two decimal place, depended on firmware version. Note that the value sent by the device is less precise than what is displayed on device screen. For example, a displayed value of 5.12 is sent by the device and thus by DDI as 5.1. For Evita 2, Evita, Fabius GS, Fabius GS premium, Fabius Tiro, Fabius plus, Fabius plus XL, Fabius OS, Fabius MRI, Evita Infinity V500, Babylog VN500, Carina: This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. This variable is available only with a precision to the tenth. For example: for a minute volume value of 7.43, the device displays 7.43 L/min but the DDI will output it as 7.4 L/min. For Savina: This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300, Fabius Family. This variable is associated with UpperAlarmLimit attribute, which is available for Perseus A500, Atlan, Savina 300, Primus, Primus IE, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Fabius Family. This variable is associated with LowerAlarmLimit attribute, which is available for Perseus A500, Atlan, Savina 300, Primus, Primus IE, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Fabius Family. This variable is available only with a precision to the tenth. For example: for a minute volume value of 7.43, the device displays 7.43 L/min but the DDI will output it as 7.4 L/min. For Zeus, Zeus IE, Primus, Primus IE, Apollo, Pallas: This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. Respiratory Minute Volume This variable is available only with a precision to the tenth. For example: for a minute volume value of 7.43, the device displays 7.43 L/min but the DDI will output it as 7.4 L/min. For Evita Infinity V500 and Babylog VN500: This variable is only output if the device version is less than 2.30, otherwise variable 2326 is exported. | Double | VitalSigns | Evita Infinity V800
1314 | Spontaneous Respiration Rate | 1/min | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Zeus IE, Oxylog 3000 plus, Evita V300. This variable is associated with UpperAlarmLimit attribute, which is available for Oxylog 3000 plus. | Double | VitalSigns | Evita Infinity V800
1318 | Static Compliance | mL/mbar | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500 Evita Infinity V600, Evita Infinity V800. For Evita Infinity V500: Cstat (Low Flow) Cstat. | Double | VitalSigns | Evita Infinity V800
1320 | Mandatory Respiration Rate Setting | 1/min | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300, Fabius Family. Respiratory rate, RR. For Evita Infinity V500, Babylog VN500: Respiratory Rate RR. | Double | Settings | Evita Infinity V800
1324 | Sigh Interval Setting | s | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. Time interval between sighs, Interval sigh. | Double | Settings | Evita Infinity V800
1326 | Sigh Number Setting | unitless | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. Number of sigh cycles, Cycles sigh. | Double | Settings | Evita Infinity V800
1328 | Apnea Tidal Volume Setting | - | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Oxylog 3000 plus, Carina, Evita V300. Tidal volume for apnea ventilation, Vtapn. Possible units: L, mL. Unit "L" is available for Evita Infinity V500, Babylog VN500, Savina 300, Oxylog 3000 plus, Carina, Evita V300. Unit "mL" is available for Evita Infinity V500, Babylog VN500, Evita V300. For Evita Infinity V500, Babylog VN500: Possible units: L, mL. VT Apnea VTapn. | Double (Multiset) | Settings | Evita Infinity V800
1329 | Apnea Respiration Rate Setting | 1/min | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Oxylog 3000 plus, Carina, Evita V300. Respiratory rate for apnea ventilation, Rrapn. | Double | Settings | Evita Infinity V800
1332 | Pressure Support Level Above PEEP Setting | mbar | For Evita 4, Evita 2 dura, Evita XL: This variable is calculated from variables 1392 and 2104. Available only with Evita XL. For Medibus.X compatible device, Perseus A500, Atlan, Evita V300, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300, Fabius Family. Pressure amplitude above PEEP in Pressure Support, DPsupp. For Evita Infinity V500, Babylog VN500: ΔPsupp. For Carina: Assisted spontaneous breathing. | Double | Settings | Evita Infinity V800
1381 | Inspiratory Flow Setting | L/min | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300, Fabius Family. | Double | Settings | Evita Infinity V800
1383 | Inspiratory Flow Trigger Setting | L/min | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Evita V300, Fabius Family. Flow trigger threshold, Flow trigger. For Primus, Primus IE, Apollo, Pallas: Flow trigger. | Double | Settings | Evita Infinity V800
1386 | Intermittent PEEP Setting | mbar | For Fabius GS, Fabius GS premium, Fabius Tiro, Fabius plus, Fabius plus XL, Fabius OS, Fabius MRI: 1mbar = 1cmH2O. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Evita V300. Additional intermittent PEEP for sighs, PEEPsigh. For Evita Infinity V500, Babylog VN500: Interm. PEEP ΔintPEEP. For Zeus, Zeus IE: PEEP for all vent mode different from Man/Spont. | Double | Settings | Evita Infinity V800
1387 | APRV Low Pressure Setting | mbar | For Evita: BIPAP low pressure For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Evita V300. Lower pressure level during APRV, Plow. For Evita Infinity V500, Babylog VN500: APRV Low Pressure Plow. | Double | Settings | Evita Infinity V800
1388 | APRV Low Time Setting | s | For Evita: BIPAP low time For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Evita V300. Time of lower pressure level during APRV, Tlow. For Evita Infinity V500, Babylog VN500: APRV Low Time Tlow. | Double | Settings | Evita Infinity V800
1389 | APRV High Pressure Setting | mbar | For Evita: BIPAP high pressure For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Evita V300. Upper pressure level during APRV, Phigh. For Evita Infinity V500, Babylog VN500: APRV High Pressure Phigh. | Double | Settings | Evita Infinity V800
1390 | APRV High Time Setting | s | For Evita: BIPAP high time For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Evita V300. Time of upper pressure level during APRV, Thigh. For Evita Infinity V500, Babylog VN500: APRV High Time Thigh. | Double | Settings | Evita Infinity V800
1391 | Apnea Time Setting | s | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Oxylog 3000 plus, Carina, Evita V300. For Evita Infinity V500, Babylog VN500: Apnea Time Tapn. For Carina: Apnea alarm time. | Double | Settings | Evita Infinity V800
1393 | Maximum Inspired Pressure Setting | mbar | For Fabius GS, Fabius GS premium, Fabius Tiro, Fabius plus, Fabius plus XL, Fabius OS, Fabius MRI: 1mbar = 1cmH2O. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300, Fabius Family. Pressure limitation, Pmax. For Evita Infinity V500, Babylog VN500: Max Pressure Pmax. For Carina: Pressure limitation. For Primus, Primus IE, Apollo, Pallas: Max. insp. Airway pressure. | Double | Settings | Evita Infinity V800
1415 | Mean Airway Pressure | mbar | For Fabius GS, Fabius GS premium, Fabius Tiro, Fabius plus, Fabius plus XL, Fabius OS, Fabius MRI: This variable is associated with LowerAlarmLimit attribute. 1mbar = 1cmH2O. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300, Fabius Family. For Zeus, Zeus IE, Primus, Primus IE, Apollo, Pallas: Mean Breathing Pressure. | Double | VitalSigns | Evita Infinity V800
1422 | Rapid Shallow Breathing Index | - | For Evita 4, Evita 2 dura, Evita XL: Firmware version 04.00 or higher For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Atlan, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Evita V300. Possible units: 1/L * min, 1/mL*min. Unit "1/L * min" is available for Evita Infinity V500, Babylog VN500, Savina 300, Evita V300. Unit "1/mL*min" is available for Evita Infinity V500, Babylog VN500, Evita V300 (added with SW 2.30 for Babylog VN500, added with SW 2.30 for Evita Infinity V500). | Double (Multiset) | VitalSigns | Evita Infinity V800
1426 | Intrinsic PEEP | mbar | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. | Double | VitalSigns | Evita Infinity V800
1569 | Minimum Airway Pressure | mbar | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. | Double | VitalSigns | Evita Infinity V800
1570 | P0.1 Occlusion Pressure | mbar | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. Occlusion pressure in P0.1 maneuver. For Evita Infinity V500, Babylog VN500: Only available for Evita Infinity V500 | Double | VitalSigns | Evita Infinity V800
1571 | Lung Volume Trapped (TrpVol) | mL | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. Volume trapped in the lung by intrinsic PEEP, Vtrap. For Evita Infinity V500: Trapped Volume Vtrap. | Double | VitalSigns | Evita Infinity V800
1575 | Respired Volume Wave | mL | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Evita V300. Respiratory volume since start of inspiration, Volume. For Evita Infinity V500, Babylog VN500: Volume since inspiration began. | Double | Waveforms | Evita Infinity V800
1637 | Ventilation Pressure Variability Setting | % | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800. For Evita Infinity V500, Babylog VN500: Supported through Medibus.X protocol. | Double | Settings | Evita Infinity V800
1638 | APRV Low Time Limitation | s | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. Maximum time of low pressure level during APRV/AutoRelease, Tlow max. For Evita Infinity V500, Babylog VN500: Tlow max. | Double | Settings | Evita Infinity V800
1640 | Spontaneous Mean Tidal Volume | mL | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500 Evita Infinity V600, Evita Infinity V800. Spontaneous mean tidal volume, VTspon mean. For Evita Infinity V500, Babylog VN500: VTspon mean (Not available on Babylog VN500). Only available for Evita Infinity V500 | Double | VitalSigns | Evita Infinity V800
1641 | Spontaneous Inspired Mean Tidal Volume | mL | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500 Evita Infinity V600, Evita Infinity V800. For Evita Infinity V500, Babylog VN500: VTispon mean (Not available on Babylog VN500). Only available for Evita Infinity V500 | Double | VitalSigns | Evita Infinity V800
1642 | Spontaneous Expired Mean Tidal Volume | mL | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800. For Evita Infinity V500, Babylog VN500: VTespon mean (Not available on Babylog VN500). Only available for Evita Infinity V500 | Double | VitalSigns | Evita Infinity V800
1650 | SmartCare FiO2 Limitation | % | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. FiO2 in SmartCare - upper limit, SC-FiO2high For Evita Infinity V500, Babylog VN500: SC-Upper limit inspiratory O2 FiO2 max. | Double | Settings | Evita Infinity V800
1651 | SmartCare PEEP Limitation | mbar | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. PEEP in SmartCare - upper limit, SC-PEEPmax. For Evita Infinity V500, Babylog VN500: SC-Upper limit PEEP, PEEP max. | Double | Settings | Evita Infinity V800
1652 | SmartCare Spontaneous Respiration Rate Lower Limitation | 1/min | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. Spontaneous respiratory rate in SmartCare - lower limit, SC-RRspon low. For Evita Infinity V500, Babylog VN500: SC-Lower limit respiratory rate in the guideline, RRspon low. | Double | Settings | Evita Infinity V800
1653 | SmartCare Spontaneous Respiration Rate Upper Limitation | 1/min | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. Spontaneous respiratory rate in SmartCare - upper limit, SC-RRspon high. For Evita Infinity V500, Babylog VN500: SC-Upper limit respiratory rate in the guideline, RRspon high. | Double | Settings | Evita Infinity V800
1654 | SmartCare Tidal Volume Limitation | mL/kg | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. Tidal volume in SmartCare - lower limit, SC-VTlow. For Evita Infinity V500, Babylog VN500: SC-Lower limit tidal volume, VT low. | Double | Settings | Evita Infinity V800
1655 | SmartCare EtCO2 Limitation | mmHg | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. End-tidal CO2 in SmartCare - upper limit, SC-etCO2high. For Evita Infinity V500, Babylog VN500: SC-Upper limit endtidal CO2 in the guideline etCO2 high. | Double | Settings | Evita Infinity V800
1656 | SmartCare Patient Height | cm | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. Patient height in SmartCare, SC-Ht. For Evita Infinity V500: SC-Body height. | Double | Settings | Evita Infinity V800
1761 | Airway Resistance | mbar/L/s | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. | Double | VitalSigns | Evita Infinity V800
2044 | P01 Repetition Interval Setting | min | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. P0.1 repetition interval. | Double | Settings | Evita Infinity V800
2092 | Inspired Minute Volume | L/min | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Carina, Evita V300. This variable is associated with UpperAlarmLimit attribute, which is available for Carina. This variable is associated with LowerAlarmLimit attribute, which is available for Carina. | Double | VitalSigns | Evita Infinity V800
2102 | Inspiratory Time Setting | s | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300, Fabius Family. | Double | Settings | Evita Infinity V800
2104 | Positive End Expiratory Pressure (PEEP) Setting | mbar | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300, Fabius Family. For Evita Infinity V500, Babylog VN500: PEEP / CPAP PEEP. For Zeus, Zeus IE: PEEP (CPAP). PEEP for mode Man/Spont. | Double | Settings | Evita Infinity V800
2110 | Inspired O2 (FiO2) Setting | % | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Zeus IE, Oxylog 3000 plus, Carina, Evita V300. For Primus, Primus IE, Apollo, Pallas: Not in device Apollo and Pallas | Double | Settings | Evita Infinity V800
2217 | Manual Breath Ventilation Pressure Setting | mbar | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300(added with SW 2.30 for Babylog VN500, added with SW 2.30 for Evita Infinity V500). Manual inspiratory pressure for Neonates during NIV - Non-Invasive Ventilation, PmanInsp. | Double | Settings | Evita Infinity V800
2245 | Manual Breath Maximum Time Limit | s | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300(added with SW 2.30 for Babylog VN500, added with SW 2.30 for Evita Infinity V500). Maximum time of pressure for manual inspiration pressure for Neonates during NIV - Non-Invasive Ventilation, TmanInsp. | Double | Settings | Evita Infinity V800
2250 | CO2 Tidal Production | mL | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. CO2 production volume per breath, VTCO2. For Evita Infinity V500, Babylog VN500: Supported through Medibus.X protocol. | Double | VitalSigns | Evita Infinity V800
2251 | CO2 Minute Production | mL/min | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Evita V300. CO2 production, V'CO2. | Double | VitalSigns | Evita Infinity V800
2255 | Device Date/Time | unitless | Format: YYYYMMDDhhmmss, where ss is always present but might be 00. | String | DeviceInfo | Evita Infinity V800 | Device Information
2324 | Inspired Tidal Volume | mL | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Carina, Evita V300. This variable is associated with UpperAlarmLimit attribute, which is available for Evita Infinity V500, Zeus IE, Evita V300. This variable is associated with LowerAlarmLimit attribute, which is available for Evita Infinity V500, Evita V300. For Primus, Primus IE, Apollo, Pallas: Insp. Tidal Volume. | Double | VitalSigns | Evita Infinity V800
2325 | Expired Tidal Volume | mL | For Evita XL: Neo expired tidal volume. Available only with firmware version 7.01 or higher. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Evita V300. | Double | VitalSigns | Evita Infinity V800
2326 | Expired Minute Volume | L/min | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. This variable is associated with UpperAlarmLimit attribute, which is available for Evita Infinity V500, Babylog VN500, Evita V300. This variable is associated with LowerAlarmLimit attribute, which is available for Evita Infinity V500, Babylog VN500, Evita V300. For Evita Infinity V500, Babylog VN500: This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. | Double | VitalSigns | Evita Infinity V800
2355 | End Tidal CO2 Percentage | % | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Evita V300, Fabius Family. This variable is associated with UpperAlarmLimit attribute, which is available for Perseus A500, Atlan, Evita Infinity V500, Babylog VN500, Savina 300, Primus, Primus IE, Apollo, Zeus IE, Oxylog 3000 plus, Evita V300. This variable is associated with LowerAlarmLimit attribute, which is available for Perseus A500, Atlan, Evita Infinity V500, Babylog VN500, Savina 300, Primus, Primus IE, Apollo, Zeus IE, Oxylog 3000 plus, Evita V300. end tidal CO2 etCO2 | Double | VitalSigns | Evita Infinity V800
2377 | Fresh Gas O2 Flow Rate | L/min | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Evita V300. O2 therapy flow, Flow. For Evita Infinity V500, Babylog VN500: Flow O2 Therapy Flow. | Double | DeviceInfo | Evita Infinity V800
2382 | Ventilator Operation Mode | unitless | This variable contains ventilation operation mode in string format, which is also send as a numerical value in variable 2758. See detailed description of variable 2758 for the list of supported values. For Evita 4, Evita 2 dura, Evita XL: Firmware version 04.00 or higher For Evita 2: Variable label is not correct. Please use variable 3853 instead. Obsolete variable that may be removed in future release. | String | DeviceInfo | Evita Infinity V800
2416 | Spontaneous Inspired Tidal Volume | mL | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. For Zeus, Zeus IE: This variable is associated with UpperAlarmLimit attribute. Insp. Tidal volume. | Double | VitalSigns | Evita Infinity V800
2421 | Airway Dead Space | mL | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. Serial dead space volume, Vds. For Evita Infinity V500, Babylog VN500: Serial dead space Vds. | Double | VitalSigns | Evita Infinity V800
2423 | Dynamic Compliance Over Last 20% Breath | unitless | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. Ratio of compliance during the last 20 % of inspiration over dynamic compliance, C20/Cdyn. For Evita Infinity V500, Babylog VN500: C20/Cdyn. | Double | VitalSigns | Evita Infinity V800
2432 | Expired Spontaneous Tidal Volume | mL | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. | Double | VitalSigns | Evita Infinity V800
2433 | Mechanical Expired Minute Volume | L/min | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Primus, Primus Infinity Empowered, Apollo, Evita V300. Mandatory expiratory minute volume, MVemand | Double | VitalSigns | Evita Infinity V800
2434 | Spontaneous Expired Minute Volume | L/min | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Primus, Primus Infinity Empowered, Apollo, Evita V300. | Double | VitalSigns | Evita Infinity V800
2467 | Patient Type ID | unitless | Consult detailed help for more information on this variable | UByte | PatientInfo | Evita Infinity V800
2542 | Airway Pressure Wave | mbar | For Fabius GS, Fabius GS premium, Fabius Tiro, Fabius plus, Fabius plus XL, Fabius OS, Fabius MRI: 1mbar = 1cmH2O. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300, Fabius Family. For Evita Infinity V500, Babylog VN500: Airway pressure Paw. | Double | Waveforms | Evita Infinity V800
2564 | Peak Airway Pressure | mbar | For Evita 4, Evita 2 dura, Evita XL: This variable is associated with UpperAlarmLimit attribute, which is also sent in separate variable 7040. For Evita 2, Evita, Savina: This variable is associated with UpperAlarmLimit attribute. For Fabius GS, Fabius GS premium, Fabius Tiro, Fabius plus, Fabius plus XL, Fabius OS, Fabius MRI: 1mbar = 1cmH2O. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300, Fabius Family. This variable is associated with UpperAlarmLimit attribute, which is available for Perseus A500, Atlan, Evita Infinity V500, Babylog VN500, Savina 300, Zeus IE, Oxylog 3000 plus, Carina, Evita V300, Fabius Family. This variable is associated with LowerAlarmLimit attribute, which is available for Zeus IE, Fabius Family. Peak inspiratory pressure, PIP. For Evita Infinity V500, Babylog VN500, Carina: This variable is associated with UpperAlarmLimit attribute. For Zeus, Zeus IE: This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. Peak Breathing Press. For Primus, Primus IE, Apollo, Pallas: Peak Breathing Pressure. | Double | VitalSigns | Evita Infinity V800
2565 | Plateau Airway Pressure | mbar | For Fabius GS, Fabius GS premium, Fabius Tiro, Fabius plus, Fabius plus XL, Fabius OS, Fabius MRI: 1mbar = 1cmH2O. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Evita V300, Fabius Family. For Evita Infinity V500, Babylog VN500: Only available for Evita Infinity V500 | Double | VitalSigns | Evita Infinity V800
2570 | Spontaneous Tidal Volume | - | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Zeus IE, Evita V300. Possible units: L, mL. For Evita Infinity V500, Babylog VN500: Possible units: L, mL. | Double (Multiset) | VitalSigns | Evita Infinity V800
2583 | I:E Ratio Setting | unitless | String format: 1:x if E > I, x:1 if I > E, or 1:1 if I = E. The useless leading or ending “0” or “.” is not transmitted, and the maximum digits after the floating point is 2. For Carina, Primus, Primus IE, Apollo, Pallas, Perseus A500: This variable is calculated by the DDI from respiration rate and inspiratory time settings. For Medibus.X compatible device: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300, Fabius Family. For Primus, Primus IE, Apollo, Perseus A500, Carina and Oxylog 3000 plus, it is calculated by the DDI from respiration rate and inspiratory time settings. For Atlan, Evita Infinity V500, Babylog VN500, Savina 300, Zeus IE, Evita V300, and Fabius Family, it is calculated by the DDI from inspiratory and expiratory components. For Evita, Evita 2, Evita 4, Evita 2 dura, and Evita XL, Fabius GS, Fabius GS premium, Fabius Tiro, Fabius plus, Fabius plus XL, Fabius OS, Fabius MRI, Evita Infinity V500, Babylog VN500, Zeus IE, Evita V300: This variable is calculated by the DDI from inspiratory and expiratory components of I:E ratio sent by the devices. | String | Settings | Evita Infinity V800
2596 | Alarm Severity Level | unitless | This variable is in multiset format. This variable contains alarm priority level, variable 6164 contains alarm codes, and variable 3878 contains alarm messages. They need to be interpreted together. One alarm code may have one or more priority values and one or more alarm messages associated with it. The possible alarms are depended on device firmware version. Please see device user manuals for details. See variables 177 and 588 and 4131 for string format. | UByte (Multiset) | Alarms | Evita Infinity V800 | Alarm And Text Message
2754 | Ventilator General Status | unitless | Ventilator Standby mode status. | UByte | DeviceInfo | Evita Infinity V800
Consult detailed help for more information on this variable
2756 | Ventilation Mode ID | unitless | This variable contains ID of basic ventilation mode ID without options. See also variable 584. | UByte | DeviceInfo | Evita Infinity V800
Consult detailed help for more information on this variable
2758 | Ventilator Operation Mode ID | unitless | For Evita 4, Evita 2 dura, Evita XL: Firmware version 04.00 or higher For Evita 2: Variable label is not correct. Please use variable 2467 instead. Obsolete variable that may be removed in future release. | UByte | DeviceInfo | Evita Infinity V800
Consult detailed help for more information on this variable
2771 | Tidal Volume Setting | L | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300, Fabius Family. Inspiratory tidal volume, Vti. For Evita Infinity V500, Babylog VN500: See also variable 1321 and 5044. For Zeus, Zeus IE: This variable contains same information as variable 1321, except that the value of different units (L or mL) are sent in separate variables. See also variable 5044. | Double | Settings | Evita Infinity V800
2784 | ASB Ramp Time Setting | s | For Evita 2: Firmware version 2.00 or higher. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300. Pressure rise time, Slope. For Carina: ASB ramp. For Primus, Primus IE, Apollo, Pallas: Slope time (Ramp). | Double | Settings | Evita Infinity V800
2952 | PS Activation State | unitless | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300, Fabius Family. A value of 1 indicates that Pressure Support (PS) is added to the current ventilation mode. A value of 0 indicates that Pressure Support (PS) is not added to the current ventilation mode. Text message displayed on device screen: "/PS" | Boolean | DeviceInfo | Evita Infinity V800
2953 | VS Activation State | unitless | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. A value of 1 indicates that Volume support (VS) is added to the current ventilation mode. A value of 0 indicates that Volume support (VS) is not added to the current ventilation mode. Text messages displayed on device screen: "/VS" | Boolean | DeviceInfo | Evita Infinity V800
2954 | APRV High Pressure (Phigh) | mbar | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. Upper pressure level during APRV, Phigh. | Double | VitalSigns | Evita Infinity V800
2959 | APRV Low Pressure (Plow) | mbar | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. Lower pressure level during APRV, Plow. | Double | VitalSigns | Evita Infinity V800
2987 | Inspired Pressure Setting | mbar | For Evita 4, Evita 2 dura, Evita XL: Available only with EvitaXL firmware version 7.00 or higher. For Fabius GS, Fabius GS premium, Fabius Tiro, Fabius plus, Fabius plus XL, Fabius OS, Fabius MRI: 1mbar = 1cmH2O. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300, Fabius Family. | Double | Settings | Evita Infinity V800
3326 | Device Average Gas Flow | L/min | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. Average device flow, FlowDev. | Double | OtherMeasuredData | Evita Infinity V800
3503 | Auto Flow Active Status | unitless | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300. A value of 1 indicates that AutoFlow (AF) is added to the current ventilation mode. A value of 0 indicates that AutoFlow (AF) is not added to the current ventilation mode. Text message displayed on device screen: "/AF" | Boolean | DeviceInfo | Evita Infinity V800
3555 | ATC Activation Status | unitless | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Evita V300. A value of 1 indicates that automatic tube compensation (ATC) is added to the current ventilation mode. A value of 0 indicates that automatic tube compensation (ATC) is not added to the current ventilation mode. Text message displayed on device screen: "/ATC" For Savina 300 (MEDIBUS.X): added with SW 5.00 | Boolean | DeviceInfo | Evita Infinity V800
3591 | Device Number | unitless | Consult detailed help for more information on this variable | String | DeviceInfo | Common | Device Information
3593 | Device Communications Status | unitless | Value "stop" sent when a request to stop the communication is received from the device | String | DeviceInfo | Common | Device Information
3853 | Patient Type | unitless | This variable contains patient type in string format, which is also send as a numerical value in variable 2467. See detailed description of variable 2467 for the list of supported values. | String | PatientInfo | Evita Infinity V800
3878 | Generic Alarm Message | unitless | This variable is in multiset format. This variable contains alarm messages, variable 2596 contains alarm priority level, variable 6164 contains alarm codes. They need to be interpreted together. One alarm code may have one or more priority values and one or more alarm messages associated with it. The possible alarms are dependent on device firmware version. Please see device user manuals for details. See variables 177 and 588 and 4131 for string format. The HL7 output can contain an '&' character to indicate a LOW alarm. In order to avoid issues, the '&' character will be replaced by a '?' character in the DDI output. | String (Multiset) | Alarms | Evita Infinity V800 | Alarm And Text Message
3966 | Inspired Mandatory Tidal Volume | mL | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. | Double | VitalSigns | Evita Infinity V800
3968 | Spontaneous Fraction of Minute Volume | % | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. Spontaneous breathing portion of minute volume, % Mvspon. For Evita Infinity V500, Babylog VN500: Spontaneous Fraction Min. Vol. %MVspon (MVspon : MVtotal). | Double | VitalSigns | Evita Infinity V800
3971 | Negative Inspiratory Force | mbar | For Evita 4, Evita 2 dura, Evita XL: The value is not transmitted, if it is displayed on Evita 4 / Evita 2 dura as "[less than] –45 mbar". Firmware version 04.00 or higher For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. | Double | VitalSigns | Evita Infinity V800
3998 | Auto Release Activation Status | unitless | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. A value of 1 indicates that AutoRelease is added to the current ventilation mode. A value of 0 indicates that AutoRelease is not added to the current ventilation mode. Text messages displayed on device screen: "/AutoRelease" | Boolean | DeviceInfo | Evita Infinity V800
4003 | Generic Respiration Rate | 1/min | For Evita 4, Evita 2 dura, Evita XL: This variable is associated with UpperAlarmLimit attribute, which is also sent in separate variable 1339. Total Respiratory Rate (RRtotal). For Evita 2, Evita, Savina, Carina: This variable is associated with UpperAlarmLimit attribute. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Carina, Evita V300, Fabius Family. This variable is associated with UpperAlarmLimit attribute, which is available for Evita Infinity V500, Babylog VN500, Savina 300, Carina, Evita V300. Respiratory rate, RR. For Evita Infinity V500, Babylog VN500: This variable is associated with UpperAlarmLimit attribute. Respiratory Rate (RRtotal) RR. | Double | VitalSigns | Evita Infinity V800
4010 | Minute Volume Alarm Delay | s | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. Delay for minute volume high alarm, MVhigh delay. This variable is associated with a ChannelID identifier. | Double | Settings | Evita Infinity V800
4011 | Flow Assist Level Setting | 0.1 mbar/L/s | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. Flow assist during PPS ventilation, Flow Assist. For Evita Infinity V500, Babylog VN500: Flow Assist. | Double | Settings | Evita Infinity V800
4012 | Volume Assist Level Setting | 0.1 mbar/L | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. Volume assist during PPS ventilation, Vol. Assist. For Evita Infinity V500, Babylog VN500: Vol. Assist. | Double | Settings | Evita Infinity V800
4015 | Disconnection Alarm Delay Time | s | For Evita 4, Evita 2 dura, Evita XL: Firmware version 4.00 or higher. For Savina: Tdeconnect. Available since Device SW version 02.00. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Carina, Evita V300. Disconnection alarm time, Tdiscon. For Evita Infinity V500, Babylog VN500: Tdisconnect. For Carina: Disconnection detection timeout. | Double | Settings | Evita Infinity V800
4044 | Medibus Revision | unitless |  | String | DeviceInfo | Common | Device Information
4045 | Device Revision | unitless | Format: mm.nn | String | DeviceInfo | Common | Device Information
4150 | VG Activation Status | unitless | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Carina, Evita V300. A value of 1 indicates that Volume Guarantee (VG) is added to the current ventilation mode. A value of 0 indicates that Volume Guarantee (VG) is not added to the current ventilation mode. Text message displayed on device screen: "/VG" | Boolean | DeviceInfo | Evita Infinity V800
4732 | Inspiratory Flow Waveform | L/min | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. | Double | Waveforms | Evita Infinity V800
4889 | Spontaneous Inspiration Time | s | For Evita 4, Evita 2 dura, Evita XL: Available only with EvitaXL firmware version 7.02 or higher. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. | Double | VitalSigns | Evita Infinity V800
4891 | Lung Time Constant | s | For Evita 4, Evita 2 dura, Evita XL: Available only with EvitaXL firmware version 7.00 or higher. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. | Double | VitalSigns | Evita Infinity V800
4892 | Airway Correlation Coefficient | unitless | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. Correlation factor, r2. For Evita Infinity V500, Babylog VN500: r². | Double | VitalSigns | Evita Infinity V800
4966 | Sigh Activation Status | unitless | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. A value of 1 indicates that the sigh function has been added to the current ventilation mode. A value of 0 indicates that the sigh function has not been added to the current ventilation mode. Text message displayed on device screen: "/Sigh" | Boolean | DeviceInfo | Evita Infinity V800
5044 | Inspired Tidal Volume Setting | mL | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Zeus IE, Evita V300. Inspiratory tidal volume, Vti. For Evita Infinity V500, Babylog VN500: See also variable 1321 and 2771. For Zeus, Zeus IE: This variable contains same information as variable 1321, except that the value of different units (L or mL) are sent in separate variables. See also variable 2771. | Double | Settings | Evita Infinity V800
5087 | Dead Space to Tidal Volume Ratio | % | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. Ratio of serial dead space volume to expiratory tidal volume, Vds/VTe. For Evita Infinity V500, Babylog VN500: Dead space ventilation Vds/VTe. | Double | VitalSigns | Evita Infinity V800
5340 | Expired Mandatory Tidal Volume | mL | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. | Double | VitalSigns | Evita Infinity V800
5374 | Expiratory Time Constant (RCexp) | s | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. | Double | VitalSigns | Evita Infinity V800
5728 | EtCO2 Phase III Increment | - | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. Possible units: mmHg/L, kPa/L, Vol%/L. Increase of measured CO2 value in phase III of the capnogram, CO2 slope. | Double (Multiset) | VitalSigns | Evita Infinity V800
5815 | Spontaneous I:E Ratio | unitless | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. Format: 1:x if E > I, x:1 if I > E, or 1:1 if I = E. The useless leading or ending “0” or “.” is not transmitted, and the maximum digits after the floating point is 2. For Evita Infinity V500, Babylog VN500: Format: 1:x if E > I, x:1 if I > E, or 1:1 if I = E. The useless leading or ending “0” or “.” is not transmitted, and the maximum digits after the floating point is 2. | String | VitalSigns | Evita Infinity V800
6164 | General Alarm Code | unitless | This variable is in multiset format. This variable contains alarm codes, variable 3878, alarm messages, variable 2596 contains alarm priority level. They need to be interpreted together. One alarm code may have one or more priority values and one or more alarm messages associated with it. The possible alarms are depended on device firmware version. Please see device user manuals for details. See variables 177 and 588 and 4131 for string format. | UShort (Multiset) | Alarms | Evita Infinity V800 | Alarm And Text Message
Consult detailed help for more information on this variable
6389 | Endotracheal Tube Size | mm | For Evita 4, Evita 2 dura, Evita XL: Possible units: "0.1mm" in firmware versions 7.00 and 7.01, "mm" otherwise. Available only with EvitaXL firmware version 7.00 or higher. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Evita V300. Inner diameter of the endotracheal/tracheostoma tube. For Evita Infinity V500, Babylog VN500: ATC Tube Diameter. | Double | DeviceInfo | Evita Infinity V800
6551 | End Tidal CO2 Pressure | - | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Evita V300, Fabius Family. This variable is associated with UpperAlarmLimit attribute, which is available for Perseus A500, Atlan, Evita Infinity V500, Babylog VN500, Savina 300, Primus, Primus IE, Apollo, Zeus IE, Oxylog 3000 plus, Evita V300. This variable is associated with LowerAlarmLimit attribute, which is available for Perseus A500, Atlan, Evita Infinity V500, Babylog VN500, Savina 300, Primus, Primus IE, Apollo, Zeus IE, Oxylog 3000 plus, Evita V300. Possible units: mmHg, kPa. end tidal CO2 etCO2 | Double | VitalSigns | Evita Infinity V800
6701 | Expiratory Time Setting | s | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Zeus IE, Evita V300. | Double | Settings | Evita Infinity V800
6702 | Inspiratory Component of I:E Ratio Setting | unitless | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Zeus IE, Evita V300, Fabius Family, Atlan For Evita Infinity V500: Partial information. Please use variable 2583 instead. Obsolete variable that may be removed in future release. | Double | Settings | Evita Infinity V800
6703 | Expiratory Component of I:E Ratio Setting | unitless | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800,Savina 300, Zeus IE, Evita V300, Fabius Family, Atlan For Evita Infinity V500: Partial information. Please use variable 2583 instead. Obsolete variable that may be removed in future release. | Double | Settings | Evita Infinity V800
6735 | Lung Elastance | mbar/L | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. | Double | VitalSigns | Evita Infinity V800
7076 | Generic Ventilation Mode | unitless | Consult detailed help for more information on this variable | String | Settings | Evita Infinity V800
7077 | Ventilation Mode Primary | unitless | Consult detailed help for more information on this variable | String | Settings | Evita Infinity V800
7078 | Ventilation Mode Options | unitless | Consult detailed help for more information on this variable | String | Settings | Evita Infinity V800
7318 | Variable PS Activation Status | unitless | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800. A value of 1 indicates that Variable PS (VPS) has been added to the current ventilation mode. A value of 0 indicates that Variable PS (VPS) has not been added to the current ventilation mode. Text message displayed on device screen: "VariablePS" | Boolean | Settings | Evita Infinity V800
7364 | Patient Operation Mode | unitless | Possible values: 1 for Invasive, 2 for Noninvasive | UByte | PatientInfo | Evita Infinity V800
7637 | Artificial Airway Compensation Status | % | For Evita 4, Evita 2 dura, Evita XL: Tube Compensation. Available only with EvitaXL firmware version 7.00 or higher. For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. Compensation rate during ATC, % comp. For Evita Infinity V500, Babylog VN500: ATC Compensation Compens. | Double | DeviceInfo | Evita Infinity V800
7640 | Mandatory Respiration Rate | 1/min | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Zeus IE, Evita V300. | Double | VitalSigns | Evita Infinity V800
7642 | Mandatory Tidal Volume | - | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Babylog VN500, Babylog VN600, Babylog VN800, Zeus IE, Evita V300(added with SW 2.30 for Babylog VN500, added with SW 2.30 for Evita Infinity V500). Possible units: L, mL. Unit "L" available for Evita Infinity V500, Babylog VN500, Evita V300 (added with SW 2.30 for Babylog VN500, added with SW 2.30 for Evita Infinity V500). Unit "mL" available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Zeus IE, Evita V300. | Double (Multiset) | VitalSigns | Evita Infinity V800
7872 | CO2 Partial Pressure Waveform | - | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Primus, Primus Infinity Empowered, Apollo, Zeus IE, Oxylog 3000 plus, Evita V300. Possible units: mmHg, kPa. Unit "mmHg" is available for Perseus A500, Atlan, Evita Infinity V500, Babylog VN500, Savina 300, Primus, Primus IE, Apollo, Zeus IE, Oxylog 3000 plus, Evita V300. Unit "kPa" is available for Perseus A500, Atlan, Savina 300, Primus, Primus IE, Apollo, Zeus IE, Oxylog 3000 plus. | Double | Waveforms | Evita Infinity V800
7879 | Time Duration at Low Pressure (Tlow) | s | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. Effective expiratory time during APRV/AutoRelease, Tlow. For Evita Infinity V500, Babylog VN500: Tlow (resulting Tlow from expiratory flow cycling). | Double | VitalSigns | Evita Infinity V800
7904 | Spontaneous Tidal Volume | - |  | Double (Multiset) | VitalSigns | Evita Infinity V800
8366 | Generic Message Code | unitless | This variable is in multiset format. This variable, along with variable 8696, contains text message code and alarm text respectively. They need to be interpreted together. | UShort (Multiset) | Messages | Evita Infinity V800 | Alarm And Text Message
8696 | Generic Message Text | unitless | This variable is in multiset format. This variable, along with variable 8366, contains text message code and alarm text respectively. They need to be interpreted together. | String (Multiset) | Messages | Evita Infinity V800 | Alarm And Text Message
9146 | End Inspiratory Pressure | mbar | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. End-inspiratory pressure, EIP. | Double | VitalSigns | Evita Infinity V800
9163 | Tracheal Pressure Waveform | mbar | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Evita V300. For Evita Infinity V500, Babylog VN500: Ptrach. | Double | Waveforms | Evita Infinity V800
9288 | Tidal Volume Indexed By Body Weight | mL/kg | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. Tidal volume per kg body weight, VT/Wt. The leakage is determined by the difference between the delivered flow on the inspiratory side and the measured flow on the expiratory side, which is displayed as the leakage minute volume MVleak and relative as % leak (MVleak to MV). For Evita Infinity V500, Babylog VN500: VT/kg BW. | Double | VitalSigns | Evita Infinity V800
9466 | Mandatory Triggered Respiration Rate | 1/min | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300(added with SW 2.42 for Evita Infinity V500, Evita V300 and Babylog VN500). Respiratory rate of triggered, mandatory breaths, RRtrig For Evita 4, Evita 2 dura, Evita XL: Available only with Evita XL firmware version 7.00 or higher. | Double | VitalSigns | Evita Infinity V800
9475 | Volume of Leakage (MVLeak) | L/min | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Carina, Evita V300. Leakage minute volume, Mvleak. The leakage is determined by the difference between the delivered flow on the inspiratory side and the measured flow on the expiratory side, which is displayed as the leakage minute volume MVleak and relative as % leak (MVleak to MV). For Evita Infinity V500, Babylog VN500: Mvleak. The leakage is determined by the difference between the delivered flow on the inspiratory side and the measured flow on the expiratory side, which is displayed as the leakage minute volume MVleak and relative as % leak (MVleak to MV). For Carina: Leakage minute volume. For Evita XL: Leakage minute volume. Available since device SW version 7.02. | Double | OtherMeasuredData | Evita Infinity V800
9476 | Percent of Leakage | % | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Savina 300, Evita V300. Leakage minute volume in % of inspiratory minute volume, % leak Leakage rel. % leak - Infinity V500 determines the difference between the delivered flow on the inspiratory side and the measured flow on the expiratory side. This difference provides a measure of the amount of leakage and is displayed by Infinity V500 as the leakage minute volume MVleak and relative as % leak (MVleak to MV). For Evita Infinity V500, Babylog VN500: Leakage rel. % leak. | Double | OtherMeasuredData | Evita Infinity V800
9477 | Inspiratory Component of Spontaneous I:E Ratio | unitless | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: Obsolete variable that may be removed in future release. Partial information of I:E ratio. For Evita Infinity V500: Partial information. Please use variable 5815 instead. Obsolete variable that may be removed in future release. | Double | VitalSigns | Evita Infinity V800
9478 | Expiratory Component of Spontaneous I:E Ratio | unitless | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: Obsolete variable that may be removed in future release. Partial information of I:E ratio. For Evita Infinity V500: Partial information. Please use variable 5815 instead. Obsolete variable that may be removed in future release. | Double | VitalSigns | Evita Infinity V800
9479 | SmartCare Target Delta Support Pressure | mbar | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. Dpressure support goal in SmartCare, SC-DPsupp goal. For Evita Infinity V500: SC-Target ΔPsupp ΔPsupp goal. | Double | VitalSigns | Evita Infinity V800
9480 | SmartCare Rated Delta Support Pressure | mbar | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. Rated Dpressure support by internal controller in SmartCare, SC-DPsupp rated. For Evita Infinity V500: SC-Rated ΔPsupp SC-Δpsupp. | Double | VitalSigns | Evita Infinity V800
9481 | SmartCare Duration of Patient Session | - | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. Duration of session (hours and minutes) in SmartCare, SC-duration. Variable is in Multiset format (need combine the first value in hours with the second value in minutes to obtain the exact duration). Please use variable 4346 instead. For Evita Infinity V500: Multiset variable (need combine the first value in hours with the second value in minutes to obtain the exact duration). | Double (Multiset) | VitalSigns | Evita Infinity V800
9482 | SmartCare Spontaneous Respiration Rate | 1/min | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. Spontaneous respiratory rate in SmartCare, SC-Rrspon. For Evita Infinity V500: SC-Rrspon. | Double | VitalSigns | Evita Infinity V800
9483 | SmartCare Tidal Volume | mL | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. Tidal volume in SmartCare, SC-VT. For Evita Infinity V500: SC-VT. | Double | VitalSigns | Evita Infinity V800
9484 | SmartCare EtCO2 | mmHg | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. End-tidal CO2 concentration in SmartCare, SC-etCO2. For Evita Infinity V500: SC-etCO2. | Double | VitalSigns | Evita Infinity V800
9485 | Maximum Inspiratory Time Setting | s | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Carina, Evita V300. Maximal inspiratory time for flow cycled strokes, Timax. For Carina: Maximal inspiration time for CPAP, ASB. | Double | Settings | Evita Infinity V800
9486 | Inspiratory Flow Termination Level Setting | % | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Perseus A500, Atlan, Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Savina 300, Oxylog 3000 plus, Evita V300. Inspiratory termination criterion based on peak inspiratory flow, Insp. term. For Evita Infinity V500: Inspiratory termination. | Double | Settings | Evita Infinity V800
9487 | Expiratory Flow Termination Setting | % | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Babylog VN500, Babylog VN600, Babylog VN800, Evita V300. Expiratory termination criterion during APRV/AutoRelease, Exp. term. For Evita Infinity V500, Babylog VN500: Expiratory termination. | Double | Settings | Evita Infinity V800
9488 | SmartCare Body Weight | kg | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. Ideal body weight in SmartCare, SC-Wt. For Evita Infinity V500: SC-Body weight. | Double | Settings | Evita Infinity V800
9489 | SmartCare Start of Night Rest | h | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. Start of night rest in SmartCare, SC-startrest. For Evita Infinity V500: SC-Start of night rest. | Double | Settings | Evita Infinity V800
9490 | SmartCare End of Night Rest | h | For Medibus.X compatible device, Perseus A500, Evita V300, Evita Infinity V600, Evita Infinity V800, Babylog VN600, Babylog VN800, Babyleo TN500, BiliLux, Isolette 8000 plus, Atlan: This variable is available for Evita Infinity V500, Evita Infinity V600, Evita Infinity V800, Evita V300. End of night rest in SmartCare, SC-endrest. For Evita Infinity V500: SC-End of night rest. | Double | Settings | Evita Infinity V800
50021 | Device Alarm Inactivation State | unitless |  | String | DeviceInfo | Evita Infinity V800