## Sheet1
ID | Label | Unit | Description | Type | Group | Device | Source
754 | Device Name | unitless | This variable is obsolete for MultiFiltrate and might be removed in a future version. | String | DeviceInfo | Common | Treatment Status/Data Management
965 | Cumulated Therapy Time | min | Treatment time, running, total | Long | OtherMeasuredData | Multifiltrate PRO | Dialysis Course
1060 | Anticoagulant Cumulated Continuous Infused Volume | 0.1 mL | Anticoagulation, continuous, total | Long | VitalSigns | Multifiltrate PRO | Dialysis Course
1090 | Anticoagulant Cumulated Bolus Infused Volume | 0.1 mL | Anticoagulation bolus, total | Long | VitalSigns | Multifiltrate PRO | Dialysis Course
1728 | Ultrafiltration/Blood Flow Ratio | % |  | Long | VitalSigns | Multifiltrate PRO | Dialysis Course
1842 | Arterial Pressure Upper Alarm Limit | mmHg | This variable is also sent as alarm limit of variableID 7438. | Long | Settings | Multifiltrate PRO | Dialysis Settings
1843 | Arterial Pressure Lower Alarm Limit | mmHg | This variable is also sent as alarm limit of variableID 7438. | Long | Settings | Multifiltrate PRO | Dialysis Settings
1913 | Device Serial Number | unitless |  | String | DeviceInfo | Common | Treatment Status
1929 | Patient ID | unitless |  | String | PatientInfo | Common | Administrative Data
2022 | Device System Code | unitless |  | String | DeviceInfo | Multifiltrate PRO | Treatment Status
2537 | Citrate Calcium Activation Status | unitless | Possible values: | Boolean | DeviceInfo | Multifiltrate PRO | Dialysis Settings
1: CiCa requested
0: CiCa not requested.
2539 | Therapy Pause Start Time | unitless | Format: YYYYMMDDhhmmss | String | OtherMeasuredData | Multifiltrate PRO | Treatment Status
2855 | Device Status | unitless | Consult detailed help for more information on this variable | UShort | DeviceInfo | Multifiltrate PRO | Treatment Status
2879 | Device Software Revision | unitless |  | String | DeviceInfo | Common | Treatment Status
2961 | Therapy Pause End Time | unitless | Format: YYYYMMDDhhmmss | String | OtherMeasuredData | Multifiltrate PRO | Treatment Status
2969 | Disposable Lifetime | min | The resolution of this variable is 10 min. This variable is sent from the device to the DDI on a regular basis, but the device will keep on sending the same value until a minimal change of 10 min is observed. | Long | DeviceInfo | Multifiltrate PRO | Dialysis Course
3003 | Blood Flow Rate Setting | mL/min |  | Long | Settings | Common | Dialysis Settings
3004 | Replacement Flow Rate Setting | mL/h | Substitution rate setting | Long | Settings | Multifiltrate PRO | Dialysis Settings
3005 | Dialysate Flow Rate Setting | mL/h |  | Long | Settings | Common | Dialysis Settings
3006 | Fluid Removal Volume Setting | mL | UF goal setting | Long | Settings | Common | Dialysis Settings
3016 | Transmembrane Pressure Upper Alarm Limit | mmHg | This variable is also sent as alarm limit of variableID 3029. | Long | Settings | Multifiltrate PRO | Dialysis Settings
3026 | Filter Pod Pressure | mmHg | Filtrate pressure | Long | OtherMeasuredData | Multifiltrate PRO | Dialysis Course
For Multifiltrate PRO: The resolution of this variable is 10 mmHg. This variable is sent from the device to the DDI on a regular basis, but the device will keep on sending the same value until a minimal change of 10 mmHg is observed.
3029 | Transmembrane Pressure | mmHg | This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. | Long | VitalSigns | Multifiltrate PRO | Dialysis Course
For Multifiltrate PRO: The resolution of this variable is 10 mmHg. This variable is sent from the device to the DDI on a regular basis, but the device will keep on sending the same value until a minimal change of 10 mmHg is observed.
3050 | Therapy Type | unitless | Consult detailed help for more information on this variable | String | DeviceInfo | Multifiltrate PRO | Treatment Status
3248 | Generic Venous Line Pressure | mmHg | This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. | Long | VitalSigns | Multifiltrate PRO | Dialysis Course
For Multifiltrate PRO: The resolution of this variable is 10 mmHg. This variable is sent from the device to the DDI on a regular basis, but the device will keep on sending the same value until a minimal change of 10 mmHg is observed.
3626 | Therapy Ending Time | unitless |  | String | OtherMeasuredData | Multifiltrate PRO | Treatment Status
3628 | Therapy Starting Time | unitless | Date/time in the format of "YYYYMMDDhhmmss" | String | OtherMeasuredData | Common | Treatment Status
3829 | Device Total Operation Time | unitless | Format: [hours]:[minutes] | String | DeviceInfo | Multifiltrate PRO | Treatment Status
4226 | Replacement Cumulated Volume | mL | Substituate-Total, total | Long | VitalSigns | Multifiltrate PRO | Dialysis Course
4229 | Fluid Removal Cumulated Volume | mL | Effective total withdrawal, total | Long | VitalSigns | Common | Dialysis Course
4573 | Dialysis Filter Lifetime | min | Filter Lifetime | Long | DeviceInfo | Multifiltrate PRO | Dialysis Course
For Multifiltrate PRO: The resolution of this variable is 10 min. This variable is sent from the device to the DDI on a regular basis, but the device will keep on sending the same value until a minimal change of 10 min is observed.
5860 | Blood Sensed Status | unitless | 0=no blood sensed 1=blood sensed | Long | DeviceInfo | Multifiltrate PRO | Dialysis Course
5868 | Fluid Removal Rate Setting | mL/h | UF rate setting | Long | Settings | Common | Dialysis Settings
5869 | Dialysate Temperature Setting | 0.1 °C |  | Long | Settings | Common | Dialysis Settings
5870 | Plasma Volume Setting | mL |  | Long | Settings | Multifiltrate PRO | Dialysis Settings
5871 | Citrate Dose Setting | 0.01 mmol/L |  | Long | Settings | Multifiltrate PRO | Dialysis Settings
5872 | Calcium Dose Setting | 0.01 mmol/L |  | Long | Settings | Multifiltrate PRO | Dialysis Settings
5875 | Citrate Flow Rate | mL/h |  | Long | VitalSigns | Multifiltrate PRO | Dialysis Course
5876 | Calcium Flow Rate | mL/h |  | Long | VitalSigns | Multifiltrate PRO | Dialysis Course
5878 | Citrate Cumulated Volume | 0.1 mL |  | Long | VitalSigns | Multifiltrate PRO | Dialysis Course
5880 | Calcium Cumulated Volume | 0.1 mL |  | Long | VitalSigns | Multifiltrate PRO | Dialysis Course
5882 | Filtration Fluid Cumulated Balance | mL | DF-Total, total balance | Long | VitalSigns | Multifiltrate PRO | Dialysis Course
5884 | Replacement Cumulated Bolus Volume | mL | Bolus-Substituate-Total, total | Long | VitalSigns | Multifiltrate PRO | Dialysis Course
6063 | Central Venous Pressure Lower Alarm Limit | mmHg | This variable is also sent as alarm limit of variableID 3248. | Long | Settings | Multifiltrate PRO | Dialysis Settings
6064 | Central Venous Pressure Upper Alarm Limit | mmHg | This variable is also sent as alarm limit of variableID 3248. | Long | Settings | Multifiltrate PRO | Dialysis Settings
6164 | General Alarm Code | unitless | Consult detailed help for more information on this variable | Long | Alarms | Multifiltrate PRO | Dialysis Course
6768 | Device Quick Test Status | unitless | T0 Test Status. Possible values: Start, End, E1, E2, E3. | String | DeviceInfo | Multifiltrate PRO | Treatment Status
7438 | Generic Arterial Line Pressure | mmHg | This variable is associated with UpperAlarmLimit and LowerAlarmLimit attributes. | Long | VitalSigns | Multifiltrate PRO | Dialysis Course
For Multifiltrate PRO: The resolution of this variable is 10 mmHg. This variable is sent from the device to the DDI on a regular basis, but the device will keep on sending the same value until a minimal change of 10 mmHg is observed.
8809 | Anticoagulant Infusion Rate Setting | 0.1 mL/h | Heparin continuous rate | Long | Settings | Common | Dialysis Settings
8817 | Patient Weight Cumulated Value | g | Absolute balancing error | Long | VitalSigns | Multifiltrate PRO | Dialysis Course
8840 | Patient Sensor Value 1 | g | Weight of scale 1 | Long | OtherMeasuredData | Multifiltrate PRO | Dialysis Course
For Multifiltrate PRO: The resolution of this variable is 10 g. This variable is sent from the device to the DDI on a regular basis, but the device will keep on sending the same value until a minimal change of 10 g is observed. The value of "0" will always be sent regardless of the delta change.
8841 | Patient Sensor Value 2 | g | Weight of scale 2 | Long | OtherMeasuredData | Multifiltrate PRO | Dialysis Course
For Multifiltrate PRO: The resolution of this variable is 10 g. This variable is sent from the device to the DDI on a regular basis, but the device will keep on sending the same value until a minimal change of 10 g is observed. The value of "0" will always be sent regardless of the delta change.
8842 | Patient Sensor Value 3 | g | Weight of scale 3 | Long | OtherMeasuredData | Multifiltrate PRO | Dialysis Course
For Multifiltrate PRO: The resolution of this variable is 10 g. This variable is sent from the device to the DDI on a regular basis, but the device will keep on sending the same value until a minimal change of 10 g is observed. The value of "0" will always be sent regardless of the delta change.
8843 | Patient Sensor Value 4 | g | Weight of scale 4 | Long | OtherMeasuredData | Multifiltrate PRO | Dialysis Course
For Multifiltrate PRO: The resolution of this variable is 10 g. This variable is sent from the device to the DDI on a regular basis, but the device will keep on sending the same value until a minimal change of 10 g is observed. The value of "0" will always be sent regardless of the delta change.
8853 | Patient Fluid Balance | mL | Balance, total | Long | VitalSigns | Multifiltrate PRO | Dialysis Course
8856 | Therapy Operation Status | unitless | Consult detailed help for more information on this variable | String | DeviceInfo | Common | Treatment Status
8967 | Case ID | unitless |  | String | PatientInfo | Common | Administrative Data
9643 | Blood Flow Rate | mL/min | Actual blood flow This variable contains value "0" when blood pump is stopped. | Long | VitalSigns | Multifiltrate PRO | Dialysis Course
For Multifiltrate PRO: The resolution of this variable is 10 mL/min. This variable is sent from the device to the DDI on a regular basis, but the device will keep on sending the same value until a minimal change of 10 mL/min is observed. The value of "0" will always be sent regardless of the delta change.
9644 | Prefilter Pod Pressure | mmHg | Pre-Hemofilter pressure | Long | OtherMeasuredData | Multifiltrate PRO | Dialysis Course
For Multifiltrate PRO: This variable is associated with LowerAlarmLimit and UpperAlarmLimit attributes.
The resolution of this variable is 10 mmHg. This variable is sent from the device to the DDI on a regular basis, but the device will keep on sending the same value until a minimal change of 10 mmHg is observed.
9648 | Transmembrane Pressure Lower Alarm Limit | mmHg | This variable is also sent as alarm limit of variableID 3029. | Long | Settings | Multifiltrate PRO | Dialysis Settings