## Sheet1
ID | Label | Unit | Description | Type | Group | Device | Source
1 | Heart Rate | 1/min |  | Short | VitalSigns | Common | Continuous Parameter
14 | SpO2 | % |  | Double | VitalSigns | PiCCO2 | Continuous Parameter
414 | Cardiac Output | L/min |  | Double | VitalSigns | Common | Discontinuous Mean Parameter
519 | Central Venous Pressure (CVP) | mmHg | This variable transmits either a continuous measurement or a manual input value. | Short | VitalSigns | Common | Continuous Parameter
585 | Patient Category | unitless | Consult detailed help for more information on this variable | Short | PatientInfo | PiCCO2 | Patient Setting
649 | Cardiac Index | L/min/m2 |  | Double | VitalSigns | PiCCO2 | Discontinuous Mean Parameter
650 | Stroke Volume | mL | Beat to beat stroke volume. | Double | VitalSigns | Common | Continuous Parameter
654 | Systemic Vascular Resistance | dyn·s/cm5 |  | Short | VitalSigns | Common | Continuous Parameter
756 | Patient Height | cm |  | Short | PatientInfo | Common | Patient Setting
757 | Patient Weight | kg |  | Double | PatientInfo | Common | Patient Setting
768 | Generic Error Code | unitless | TD measurement error, "0" is transmitted if no error occurs. | Short | Alarms | Common | Error
1246 | Hematocrit | % |  | Double | VitalSigns | PiCCO2 | ScvO2 Calibration
1434 | Patient Body Surface Area | m2 |  | Double | PatientInfo | PiCCO2 | Patient Calculation
1435 | Systemic Vascular Resistance Index | dyn·s·m2/cm5 |  | Short | VitalSigns | PiCCO2 | Continuous Parameter
1468 | O2 Consumption (VO2) | mL/min |  | Short | VitalSigns | PiCCO2 | Continuous Parameter
1472 | O2 Consumption Index | mL/min/m2 |  | Short | VitalSigns | PiCCO2 | Continuous Parameter
1710 | Cardiac Power Index (CCO) | Not Available |  | Double | VitalSigns | PiCCO2 | Continuous Parameter
1711 | Cardiac Power Output (CCO) | W |  | Double | VitalSigns | PiCCO2 | Continuous Parameter
1720 | ICG Measurement Counter | unitless | Counter for signalizing update of ICG measurement. | Short | OtherMeasuredData | PiCCO2 | Discontinuous ICG
1721 | Plasma Disappearance Rate | Not Available |  | Double | VitalSigns | PiCCO2 | Discontinuous ICG
1722 | ICG Retention Rate | % |  | Double | VitalSigns | PiCCO2 | Discontinuous ICG
1730 | Cardiac Index ((Thermodilution) | L/min/m2 |  | Double | VitalSigns | PiCCO2 | Discontinuous TD
1731 | Cardiac Output (Thermodilution) | L/min |  | Double | VitalSigns | PiCCO2 | Discontinuous TD
1732 | Global Ejection Fraction (Thermodilution) | % |  | Short | VitalSigns | PiCCO2 | Discontinuous TD
1733 | Global End-Diastolic Volum (Thermodilution) | mL |  | Short | VitalSigns | PiCCO2 | Discontinuous TD
1734 | Global End-Diastolic Volum Index (Thermodilution) | mL/m2 |  | Short | VitalSigns | PiCCO2 | Discontinuous TD
1735 | Pulmonary Vascular Permeability Index (Thermodilution) | unitless |  | Double | VitalSigns | PiCCO2 | Discontinuous TD
1821 | Generic ICG Error | unitless | ICG measurement error, "0" is transmitted if no error occurs. | Short | Alarms | PiCCO2 | Error
1929 | Patient ID | unitless | 9 digits maximum, only contain numbers from 0 to 9. | String | PatientInfo | Common | Patient Setting
1930 | Patient Name | unitless | Allowed characters include A..Z, 0..9, ' ', '.', and '-'. The device sends a single character each 400 ms, so it takes time to transmit the full patient name. | String | PatientInfo | PiCCO2 | Patient Setting
1943 | Predicted Patient Body Surface Area | m2 |  | Double | PatientInfo | PiCCO2 | Patient Calculation
1944 | Predicted Patient Weight | kg |  | Double | PatientInfo | PiCCO2 | Patient Calculation
2255 | Device Date/Time | unitless | This variable contains the current time of the device in the form of "MM/dd/yyyy - hh:mm:ss". It is also sent in the "DeviceTime" attribute of other variables. | String | DeviceInfo | PiCCO2 | Device Information
2259 | Arterial BP Systolic | mmHg |  | Double | VitalSigns | Common | Continuous Parameter
2260 | Arterial BP Diastolic | mmHg |  | Double | VitalSigns | Common | Continuous Parameter
2261 | Arterial BP Mean | mmHg |  | Double | VitalSigns | Common | Continuous Parameter
2341 | Patient Blood Temperature | °C | A offset of 30°C is already taken into account by the DDI and consequently, the user receives the blood temperature real value, without having to interpret it. | Double | VitalSigns | Common | Discontinuous TD
2342 | Injectate Temperature | °C | A offset of 15°C is already taken into account by the DDI and consequently, the user receives the injectate temperature real value, without having to interpret it. | Double | OtherMeasuredData | Common | Discontinuous TD
2452 | O2 Delivery (DO2) | mL/min |  | Short | VitalSigns | PiCCO2 | Continuous Parameter
2462 | SvO2 Alarm | unitless | High/Low Alarm of ScvO2 Consult detailed help for more information on this variable | Short | Alarms | PiCCO2 | Alarm
2819 | SvO2 Calibration Status | unitless | ScvO2Cal: Point of time of calibration. "1" once during “Bloodsample drawn”, otherwise "0". | Boolean | DeviceInfo | PiCCO2 | ScvO2 Calibration
2879 | Device Software Revision | unitless | Release PiCCO software version, composed by one letter (type of protocol: V for usual protocol, D for demo version and X for unreleased version) and from one to three digits seperated by dots (version number). For example: D5.1 | String | DeviceInfo | Common | Device Information
2915 | Ejection Fraction | % | Global Ejection Fraction | Short | VitalSigns | PiCCO2 | Discontinuous Mean Parameter
3184 | Alternate Patient ID | unitless | Alternate display of Patient ID. Allowed characters include A..Z, 0..9, ' ', '.', and '-'. The device sends a single character each 400 ms, so it takes time to transmit the full patient ID. | String | PatientInfo | PiCCO2 | Patient Setting
3789 | Arterial O2 Saturation (SaO2) | % | Arterial oxygenation saturation (input value for display of parameters DO2 and VO2). | Double | VitalSigns | PiCCO2 | ScvO2 Calibration
3877 | Device Protocol Revision | unitless | Release PiCCO2 software version, composed by one letter (type of protocol: V for usual protocol, D for demo version and X for unreleased version) and from one to three digits seperated by dots (version number). For example: D5.1 | String | DeviceInfo | PiCCO2 | Device Information
4532 | Stroke Volume Variation | % |  | Short | VitalSigns | Common | Continuous Parameter
4575 | O2 Delivery Index (DO2I) | mL/min/m2 |  | Short | VitalSigns | PiCCO2 | Continuous Parameter
4865 | Cardiac Function Index (Thermodilution) | 1/min |  | Double | VitalSigns | PiCCO2 | Discontinuous TD
4866 | Intrathoracic Blood Volume (Thermodilution) | mL |  | Short | VitalSigns | PiCCO2 | Discontinuous TD
4867 | Intrathoracic Blood Volume Index (Thermodilution) | mL/m2 |  | Short | VitalSigns | PiCCO2 | Discontinuous TD
4868 | Extravascular Lung Water (Thermodilution) | mL |  | Short | VitalSigns | PiCCO2 | Discontinuous TD
4869 | Extravascular Lung Water Index (Thermodilution) | mL/kg |  | Short | VitalSigns | PiCCO2 | Discontinuous TD
5072 | Intrathoracic Blood Volume (ITBV) | mL |  | Short | VitalSigns | Common | Discontinuous Mean Parameter
5073 | Intrathoracic Blood Volume Index (ITBI) | mL/m2 |  | Short | VitalSigns | PiCCO2 | Discontinuous Mean Parameter
5074 | Extravascular Lung Water | mL |  | Short | VitalSigns | Common | Discontinuous Mean Parameter
5075 | Extravascular Lung Water Index | mL/kg |  | Short | VitalSigns | PiCCO2 | Discontinuous Mean Parameter
5076 | Global End Diastolic Volume | mL |  | Short | VitalSigns | Common | Discontinuous Mean Parameter
5077 | Global End Diastolic Volume Index | mL/m2 |  | Short | VitalSigns | PiCCO2 | Discontinuous Mean Parameter
5078 | Cardiac Function Index | 1/min |  | Double | VitalSigns | PiCCO2 | Discontinuous Mean Parameter
5642 | SvO2 Calibration Data | unitless | ScvO2Cal: ScvO2 Calibration input. | Double | DeviceInfo | PiCCO2 | ScvO2 Calibration
5684 | Continuous Cardiac Output (CCO) Alarm | unitless | High/Low Alarm of PCCO Consult detailed help for more information on this variable | Short | Alarms | PiCCO2 | Alarm
6403 | Pulse Contour Cardiac Output | L/min |  | Double | VitalSigns | Common | Continuous Parameter
6404 | Thermodilution Measurement Status | unitless | Consult detailed help for more information on this variable | Short | OtherMeasuredData | Common | Discontinuous TD
6407 | Index Of Left Ventricular Contractility (dPmx) | mmHg/s |  | Short | VitalSigns | Common | Continuous Parameter
6410 | Mean Transit Time (MTt) | s |  | Double | VitalSigns | Common | Discontinuous TD
6411 | Down Slope Time | s |  | Double | VitalSigns | Common | Discontinuous TD
6422 | Temperature Difference (Thermodilution) | °C |  | Double | VitalSigns | PiCCO2 | Discontinuous TD
6424 | Injectate Temperature (Thermodilution) | °C |  | Double | OtherMeasuredData | PiCCO2 | Discontinuous TD
6428 | PICCO Catheter | unitless | Range 0-14 Consult detailed help for more information on this variable | Short | DeviceInfo | Common | Device Information
6434 | Pulse Pressure Variation | % |  | Short | VitalSigns | Common | Continuous Parameter
6493 | Pulmonary Vascular Permeability Index | unitless |  | Double | VitalSigns | PiCCO2 | Discontinuous Mean Parameter
6533 | Patient Gender | unitless | Consult detailed help for more information on this variable | Short | PatientInfo | PiCCO2 | Patient Setting
6689 | Pulse Contour Cardiac Index | L/min/m2 |  | Double | VitalSigns | PiCCO2 | Continuous Parameter
6690 | Thermodilution Measurement Counter | unitless | Counter for signalizing when all discontinuous thermodilution parameters have been updated. | Short | OtherMeasuredData | PiCCO2 | Discontinuous TD
6692 | Right to Left Heart Shunt Fraction | % |  | Short | VitalSigns | PiCCO2 | Discontinuous TD
6694 | Arterial Blood Pressure Wave | mmHg |  | Double | Waveforms | Common | Continuous Parameter
6695 | Injectate Volume Setting | mL |  | Double | Settings | Common | Discontinuous TD
6696 | Stroke Volume Min | mL |  | Double | VitalSigns | Common | Continuous Parameter
6697 | Stroke Volume Max | mL |  | Double | VitalSigns | Common | Continuous Parameter
6698 | Stroke Volume Index | mL/m2 |  | Double | VitalSigns | PiCCO2 | Continuous Parameter
6699 | Stroke Volume Mean | mL |  | Double | VitalSigns | Common | Continuous Parameter
7823 | Central Venous O2 Saturation (ScvO2) | % |  | Double | VitalSigns | PiCCO2 | Continuous Parameter
7941 | SpO2 Alarm | unitless | High/Low Alarm of SpO2 Consult detailed help for more information on this variable | Short | Alarms | PiCCO2 | Alarm